## 1 Diagnostics Categories
| GDF TB Diagnostics Product categories | Please click any category to view the full product specifications
1. DETECTION AND RESISTANCE OF MTB
1.1 MOLECULAR TESTING: POC / NEAR-POC
1.1.1 Truenat MTB equipment
1.1.2 Truenat MTB reagents
1.1.3 Truenat consumables
1.1.4 Truenat services
1.1.5 TB-LAMP equipment
1.1.6 TB-LAMP reagents
1.1.7 TB-LAMP consumables
1.1.8 TB-LAMP services
1.1.11 GeneXpert Equipment (10 color)
1.1.12 GeneXpert Satellite (10 color)
1.1.13 GeneXpert reagents (6 and 10 color)
1.1.14 GeneXpert consumables (6 and 10 color)
1.1.15 GeneXpert services (same for 6 and 10 color)
1.2 RAPID ANTIGEN DETECTION
1.2.1 Rapid antigen detection tests (LAM)
1.3 SMEAR MICROSCOPY
1.3.1 Smear microscopy kits
1.3.2 Smear microscopy equipment
1.3.3 Smear microscopy reagents
1.3.4 Smear microscopy consumables
1.4 MOLECULR TESTING: CENTRALIZED
1.4.1 LPA equipment
1.4.2 LPA reagents
1.4.3 LPA consumables
1.4.6 BD MAX equipment
1.4.7 BD MAX reagents
1.4.8 BD MAX Consumables
1.4.9 Fluorotype equipment
1.4.10 Fluorotype reagents
1.4.11 Fluorotype consumables
1.4.12 Fluorotype services
1.5 SPECIMEN CULTURE AND DST
1.5.1 Solid culture equipment
1.5.2 Solid culture reagents
1.5.3 Solid culture consumables
1.5.4 Liquid culture equipment
1.5.5 Liquid culture reagents
1.5.6 Liquid culture consumables
1.5.7 Solid and liquid culture DST: FLD pure substances
1.5.8 Solid and liquid culture DST: SLD pure substances
1.6 MTB IDENTIFICATION
1.6.1 MTB Identification Rapid test
2. DETECTION OF LATENT TB
2.1 IGRA
2.1.1 IGRA Equipment
2.1.2 IGRA Testing
3. DETECTION OF COVID-19
3.1 COVID-19
3.1.1 COVID-19 testing
4. SAMPLE AND MEDIA PREPARATION
4.1 SPECIMEN TRANPORTATION AND PROCESSING
4.1.1 Specimen transport consumables
4.1.2 Specimen processing equipment
4.1.3 Specimen processing reagents
4.1.4 Specimen processing consumables
4.2 MEDIA PREPARATION
4.2.1 Media preparation equipment
4.2.2 Media preparation plasticware
4.2.3 Media preparation glassware
4.2.4 Media preparation consumables
5. INFECTION PREVENTION PRODUCTS
5.1 PERSONAL PROTECTIVE EQUIPMENT (PPE)
5.1.1 PPE consumables
5.2 BIOSAFETY PRODUCTS
5.2.1 Biosafety consumables
5.2.2 Disinfectants
5.3 WASTE MANAGEMENT
5.3.1 Waste management reagents
5.3.2 Waste management consumables
6. LABORATORY EQUIPMENT
6.1 UPS AND SOLAR PANELS
6.1.1 UPSs
6.1.2 Solar panels
6.2 AUTOCLAVES
6.2.1 Horizontal Autoclaves
6.2.2 Vertical Autoclaves
6.2.3 Autoclave water supply
6.3 BIOSAFETY CABINETS
6.3.1 Biosafety cabinet Class II - 3 filters
6.3.2 Biosafety cabinet Class II Type A2, 4 feet
6.3.3 Biosafety cabinet Class II Type A2, 5 feet
6.3.4 Biosafety cabinet Class II Type A2, 6 feet
6.4 LABORATORY FURNITURE
6.4.1 Laboratory furniture

## 2 Diagnostics Ordering list
|  | GDF TB Diagnostics Ordering List May, 2024 |  |  |  | Quantity to order | Product price Concessional prices for the eligible countries are found here: https://dxc-marketplace.finddx.org/pages/accessible-price-agreements. For the Non concessional prices of the non-eligible countries, please contact the GDF Country Supply Officer in charge. |  |  | Product transport type and packaging | [5 empty] | Product information |  | Products links
GDF Code | Product Name | Pack size | Minimum remaining shelf life when requesting shelf life approval | Product technical specifications | Supplier |  | Indicative FCA Price | Currency | Total Cost | Transport type | Weight Kg | Depth cm | Width cm | Height cm | Volume m3 | Manufacturer product code | HS Code | Product picture, Product information link
1. DETECTION AND RESISTANCE OF MTB
1.1 MOLECULAR TESTING: POC / NEAR-POC
1.1.1 Truenat MTB equipment
106689 | Truelab Uno Dx Workstation (623010001) | 1 set | N/A | Truelab Real Time Quantitative micro PCR System to run 10 -12 samples / 8 hours. The workstation has Optics: Fluorescence, Three Wavelength with a speed of 40 cycles of PCR/35 minutes. Interface: Wi-Fi, 3G, Bluetooth with Auto-calibration with a memory of 20,000 test results; Operating temperature: 15-40oC, RH: 10 -80%; Display Capacitive 5” Touch Screen LCD Display. The initial period of the warranty is 12 months from the date of acceptance of the goods following their successful installation at site, or 14 months form date of invoice, whichever is earlier. This item includes: 1 Truelab Uno Dx Real Time Quantitative micro PCR Analyzer (603021001), 1 Trueprep AUTO v2 Universal Cartridges based sample prep device (603042001); 1 Truelab micro PCR Printer (603050001), 1 Truepet SPA fixed volume precision micropipette-6 µl (604070006), 1 Field carry case (603120001) and accessories relevant to the equipment Truelab Software updates are offered for FREE- subject to Internet /WIFI connectivity of the installed devices at their respective sites. The services come separately. Therefore, it needs to be ordered with the Truelab and Trueprep Installation and training and the Truelab UNO DX Workstation Extended warranty 1 year (or the numbe rof years needed) | Molbio Diagnostics |  | 10000 | USD | 0 | General Cargo UN 3481 UN 3481, Section II, PI -967 as per ICAO/IATA Lithium-ion batteries contained in equipment | 11 | 44 | 83 | 30 | 0.11 | 623010001 | 9018 | https://www.molbiodiagnostics.com/product_details.php?id=44
106690 | Truelab Duo Workstation (643010001) | 1 set | N/A | Truelab Real Time Quantitative micro PCR System to run 22 -25 samples / 8 hours. The workstation has Optics: Fluorescence, Three Wavelength with a speed of 40 cycles of PCR/35 minutes. Interface: Wi-Fi, 3G, Bluetooth with Auto-calibration with a memory of 20,000 test results; Operating temperature: 15-40oC, RH: 10 -80%; Display Capacitive 5” Touch Screen LCD Display. The initial period of the warranty is 12 months from the date of acceptance of the goods following their successful installation at site, or 14 months form date of invoice, whichever is earlier. This item includes: 1 Truelab duo Real Time Quantitative micro PCR Analyzer (603022001), 1 Trueprep AUTO v2 Universal Cartridges based sample prep device (603042001); 1 Truelab micro PCR Printer (603050001), 1 Truepet fixed micropipette-6 µl (604070006) and accessories relevant to the equipment Truelab Software updates are offered for FREE- subject to Internet /WIFI connectivity of the installed devices at their respective sites. The services come separately. Therefore, it needs to be ordered with the Truelab and Trueprep Installation and training and the Truelab DUO Workstation Extended warranty 1 year (or the numbe rof years needed) | Molbio Diagnostics |  | 14000 | USD | 0 | General Cargo UN 3481 UN 3481, Section II, PI -967 as per ICAO/IATA Lithium-ion batteries contained in equipment | 14 | 44 | 83 | 30 | 0.11 | 643010001 | 9018 | https://www.molbiodiagnostics.com/product_details.php?id=45
106691 | Truelab Quattro Workstation (653010001) | 1 set | N/A | Truelab Real Time Quantitative micro PCR System to run 42 -45 samples / 8 hours. The workstation has Optics: Fluorescence, Three Wavelength with a speed of 40 cycles of PCR/35 minutes. Interface: Wi-Fi, 3G, Bluetooth with Auto-calibration with a memory of 20,000 test results; Operating temperature: 15-40oC, RH: 10 -80%; Display Capacitive 5” Touch Screen LCD Display. The initial period of the warranty is 12 months from the date of acceptance of the goods following their successful installation at site, or 14 months form date of invoice, whichever is earlier. This item includes: 1 Truelab Quattro Real Time Quantitative micro PCR Analyzer (603023001), 2 Trueprep AUTO v2 Universal Cartridges based sample prep device (603042001); 1 Truelab micro PCR Printer (603050001), 1 Truepet fixed micropipette-6 µl (604070006) and accessories relevant to the equipment Truelab Software updates are offered for FREE- subject to Internet /WIFI connectivity of the installed devices at their respective sites. The services come separately. Therefore, it needs to be ordered with the Truelab and Trueprep Installation and training and the Truelab Quattro Workstation Extended warranty 1 year (or the numbe rof years needed) | Molbio Diagnostics |  | 18000 | USD | 0 | General Cargo UN 3481 UN 3481, Section II, PI -967 as per ICAO/IATA Lithium-ion batteries contained in equipment | 22 | 56 | 104 | 34 | 0.198 | 653010001 | 9018 | https://www.molbiodiagnostics.com/product_details.php?id=46
1.1.2 Truenat MTB reagents
106692 | Truenat MTB, 5 tests (601030005) | 5 tests | 19 months | Chip-based Real Time PCR test for Mycobacterium tuberculosis. Assay based on Taqman chemistry with minimal sample requirement @ 6µL. Test Duration: 35 minutes. This item includes: A. Individually sealed pouches with: Truenat MTB micro PCR chip; Microtube with freeze dried PCR reagents; DNase & RNase free pipette tip and Desiccant pouch. B. Trueprep AUTO v2 Universal Cartridge Based sample prep kit. C. Trueprep AUTO MTB Sample Pre-treatment Pack (Sample pre-treatment pack for MTB samples). D. Package Insert. · For an order of 100 tests (20 X 5 tests) or more, 20% of Truenat MTB RIF (any pack size: 5, 20 or 50 tests) will be supplied without any extra costs (at 0 USD). Given that the free 20% tests will not always end up in even quantity, a round up the quantities upwards will be ordered. · For an order of 500 test (100 X 5 tests), one (1)Truenat Positive Control Kit - Panel I, will be supplied without any extra costs (at 0 USD). | Molbio Diagnostics |  | 39.5 | USD | 0 | CRT cargo (+15-+25°C) | 10 | 89 | 79 | 37 | 0.07 | 601030005 | 3822 | https://www.molbiodiagnostics.com/product_details.php?id=1
106959 | Truenat MTB, 25 tests (601030025) | 25 tests | 19 months | Chip-based Real Time PCR test for Mycobacterium tuberculosis. Assay based on Taqman chemistry with minimal sample requirement @ 6µL. Test Duration: 35 minutes. This item includes: A. Individually sealed pouches with: Truenat MTB micro PCR chip; Microtube with freeze dried PCR reagents; DNase & RNase free pipette tip and Desiccant pouch. B. Trueprep AUTO v2 Universal Cartridge Based sample prep kit. C. Trueprep AUTO MTB Sample Pre-treatment Pack (Sample pre-treatment pack for MTB samples). D. Package Insert. · For an order of 100 tests (5 X 20 tests) or more, 20% of Truenat MTB RIF (any pack size: 5, 20 or 50 tests) will be supplied without any extra costs (at 0 USD). Given that the free 20% tests will not always end up in even quantity, a round up the quantities upwards will be ordered. · For an order of 500 test (25 X 20 tests), one (1)Truenat Positive Control Kit - Panel I, will be supplied without any extra costs (at 0 USD). | Molbio Diagnostics |  | 197.5 | USD | 0 | CRT cargo (+15-+25°C) | 10 | 143 | 89 | 72 | 0.23 | 601030025 | 3822 | https://www.molbiodiagnostics.com/product_details.php?id=1
106694 | Truenat MTB, 50 tests (601030050) | 50 tests | 19 months | Chip-based Real Time PCR test for Mycobacterium tuberculosis. Assay based on Taqman chemistry with minimal sample requirement @ 6µL. Test Duration: 35 minutes. This item includes: A. Individually sealed pouches with: Truenat MTB micro PCR chip; Microtube with freeze dried PCR reagents; DNase & RNase free pipette tip and Desiccant pouch. B. Trueprep AUTO v2 Universal Cartridge Based sample prep kit. C. Trueprep AUTO MTB Sample Pre-treatment Pack (Sample pre-treatment pack for MTB samples). D. Package Insert. · For an order of 100 tests (2 X 50 tests) or more, 20% of Truenat MTB RIF (any pack size: 5, 20 or 50 tests) will be supplied without any extra costs (at 0 USD). Given that the free 20% tests will not always end up in even quantity, a round up the quantities upwards will be ordered. · For an order of 500 test (10 X 50 tests), one (1)Truenat Positive Control Kit - Panel I, will be supplied without any extra costs (at 0 USD). | Molbio Diagnostics |  | 395 | USD | 0 | CRT cargo (+15-+25°C) | 10 | 180 | 143 | 80 | 0.51 | 601030050 | 3822 | https://www.molbiodiagnostics.com/product_details.php?id=1
106695 | Truenat MTB Plus, 5 tests (601130005) | 5 tests | 19 months | Chip-based Real Time PCR test for Mycobacterium tuberculosis. High sensitive semi Quantitative detection and diagnosis of Mycobacterium tuberculosis (MTB) with minimal sample requirement @ 6µL. Duration: 35 minutes. This item includes: A. Individually sealed pouches with: Truenat MTB Plus micro PCR chip; Microtube with freeze dried PCR reagents; DNase & RNase free pipette tip and Desiccant pouch. B. Trueprep AUTO v2 Universal Cartridge Based sample prep kit. C. Trueprep AUTO MTB Sample Pre-treatment Pack (Sample pre-treatment pack for MTB samples). D. Package Insert. · For an order of 100 tests (20 X 5 tests) or more, 20% of Truenat MTB RIF (any pack size: 5, 20 or 50 tests) will be supplied without any extra costs (at 0 USD). Given that the free 20% tests will not always end up in even quantity, a round up the quantities upwards will be ordered. · For an order of 500 test (100 X 5 tests), one (1)Truenat Positive Control Kit - Panel I, will be supplied without any extra costs (at 0 USD). | Molbio Diagnostics |  | 39.5 | USD | 0 | CRT cargo (+15-+25°C) | 10 | 89 | 79 | 37 | 0.07 | 601130005 | 3822 | https://www.molbiodiagnostics.com/product_details.php?id=49
106960 | Truenat MTB Plus, 25 tests (601130025) | 25 tests | 19 months | Chip-based Real Time PCR test for Mycobacterium tuberculosis. High sensitive semi Quantitative detection and diagnosis of Mycobacterium tuberculosis (MTB) with minimal sample requirement @ 6µL. Duration: 35 minutes. This item includes: A. Individually sealed pouches with: Truenat MTB Plus micro PCR chip; Microtube with freeze dried PCR reagents; DNase & RNase free pipette tip and Desiccant pouch. B Trueprep AUTO v2 Universal Cartridge Based sample prep kit. C. Trueprep AUTO MTB Sample Pre-treatment Pack (Sample pre-treatment pack for MTB samples). D. Package Insert. · For an order of 100 tests (5 X 20 tests) or more, 20% of Truenat MTB RIF (any pack size: 5, 20 or 50 tests) will be supplied without any extra costs (at 0 USD). Given that the free 20% tests will not always end up in even quantity, a round up the quantities upwards will be ordered. · For an order of 500 test (25 X 20 tests), one (1)Truenat Positive Control Kit - Panel I, will be supplied without any extra costs (at 0 USD). | Molbio Diagnostics |  | 197.5 | USD | 0 | CRT cargo (+15-+25°C) | 10 | 143 | 89 | 72 | 0.23 | 601130025 | 3822 | https://www.molbiodiagnostics.com/product_details.php?id=49
106697 | Truenat MTB Plus, 50 tests (601130050) | 50 tests | 19 months | Chip-based Real Time PCR test for Mycobacterium tuberculosis. High sensitive semi Quantitative detection and diagnosis of Mycobacterium tuberculosis (MTB) with minimal sample requirement @ 6µL. Duration: 35 minutes. This item includes: A. Individually sealed pouches with: Truenat MTB Plus micro PCR chip; Microtube with freeze dried PCR reagents; DNase & RNase free pipette tip and Desiccant pouch. B. Trueprep AUTO v2 Universal Cartridge Based sample prep kit. C. Trueprep AUTO MTB Sample Pre-treatment Pack (Sample pre-treatment pack for MTB samples). D. Package Insert. · For an order of 100 tests (2 X 50 tests) or more, 20% of Truenat MTB RIF (any pack size: 5, 20 or 50 tests) will be supplied without any extra costs (at 0 USD). Given that the free 20% tests will not always end up in even quantity, a round up the quantities upwards will be ordered. · For an order of 500 test (10 X 50 tests), one (1)Truenat Positive Control Kit - Panel I, will be supplied without any extra costs (at 0 USD). | Molbio Diagnostics |  | 395 | USD | 0 | CRT cargo (+15-+25°C) | 10 | 180 | 143 | 80 | 0.51 | 601130050 | 3822 | https://www.molbiodiagnostics.com/product_details.php?id=49
106698 | Truenat MTB RIF, 5 tests (601210005) | 5 tests | 19 months | Detection of Rifampicin resistance in Mycobacterium tuberculosis in Truenat MTB/MTB Plus positive human specimen and aids in the diagnosis of MDR-TB with minimal sample requirement @ 6µL. Duration: 1 hour. This item includes: Individually sealed pouches with: Truenat MTB-RIF Dx micro PCR chip; Microtube with freeze dried PCR reagents; DNase & RNase free pipette tip; Desiccant pouch and the package Insert. 20% of Truenat MTB RIF tests are free of cost (0 USD) when the quantity ordered of Truenat MTB tests or MTB Plus tests is above 100 tests. | Molbio Diagnostics |  | 40 | USD | 0 | CRT cargo (+15-+25°C) | 10 | 89 | 79 | 37 | 0.07 | 601210005 | 3822 | https://www.molbiodiagnostics.com/product_details.php?id=25
106699 | Truenat MTB RIF, 20 tests (601210020) | 20 tests | 19 months | Detection of Rifampicin resistance in Mycobacterium tuberculosis in Truenat MTB/MTB Plus positive human specimen and aids in the diagnosis of MDR-TB with minimal sample requirement @ 6µL. Duration: 1 hour. This item includes: Individually sealed pouches with: Truenat MTB-RIF Dx micro PCR chip; Microtube with freeze dried PCR reagents; DNase & RNase free pipette tip; Desiccant pouch and the package Insert. 20% of Truenat MTB RIF tests are free of cost (0 USD) when the quantity ordered of Truenat MTB tests or MTB Plus tests is above 100 tests. | Molbio Diagnostics |  | 160 | USD | 0 | CRT cargo (+15-+25°C) | 10 | 143 | 89 | 72 | 0.23 | 601210020 | 3822 | https://www.molbiodiagnostics.com/product_details.php?id=25
106700 | Truenat MTB RIF, 50 tests (601210050) | 50 tests | 19 months | Detection of Rifampicin resistance in Mycobacterium tuberculosis in Truenat MTB/MTB Plus positive human specimen and aids in the diagnosis of MDR-TB with minimal sample requirement @ 6µL. Duration: 1 hour. This item includes: Individually sealed pouches with: Truenat MTB-RIF Dx micro PCR chip; Microtube with freeze dried PCR reagents; DNase & RNase free pipette tip; Desiccant pouch and the package Insert. 20% of Truenat MTB RIF tests are free of cost (0 USD) when the quantity ordered of Truenat MTB tests or MTB Plus tests is above 100 tests. | Molbio Diagnostics |  | 400 | USD | 0 | CRT cargo (+15-+25°C) | 10 | 180 | 143 | 80 | 0.51 | 601210050 | 3822 | https://www.molbiodiagnostics.com/product_details.php?id=25
1.1.3 Truenat consumables
106886 | Truenat Positive Control Kit - Panel I (801010008) | 1 unit | 10 months | Control Kit for Truenat Chip-based micro PCR Tests: MTB, MTB Plus, MTB-RIF Dx, Beta CoV, SARS CoV-2, COVID-19 and H1N1. item includes: A. Individually sealed pouch, containing 1. A strip of 8 micro tubes each containing dried down Positive Control : 1 Strip. 2. A screw Cap vial containing 1 ml of Negative Control : 1 x 1 ml. B. Package Insert One (1)Truenat Positive Control Kit - Panel I is provided free of charge (0 USD), for every 500 Truenat MTB or Truenat MTB plus tests purchased | Molbio Diagnostics |  | 8 | USD | 0 | CRT cargo (+15-+25°C) | 0.1 | 7.5 | 1.6 | 14.2 | 0.00017 | 801010008 | 38220090 | https://www.molbiodiagnostics.com/product_details.php?id=59
106834 | Suitcase for the Truelab Duo Workstation (603170001) | 1 unit | N/A | Suitcase for the Truelab Duo Workstation - field case. The suitcase for the Truenat Duo workstation allows for the system to be easily transported for active case finding activities so the workstations can be easily and securely transported when needed. | Molbio Diagnostics |  | 350 | USD | 0 | General Cargo | 17 | 66 | 49 | 39 | 0.141 | 603170001 | 9027/ 9090 | https://www.molbiodiagnostics.com/
106955 | Suitcase for the Truelab Quattro Workstation (603180001) | 1 unit | N/A | Suitcase for the Truelab Quattro Workstation - field case. The suitcase for the Truenat Quattro workstation allows for the system to be easily transported for active case finding activities so the workstations can be easily and securely transported when needed. | Molbio Diagnostics |  | 350 | USD | 0 | General Cargo | 17 | 66 | 49 | 39 | 0.141 | 603180001 | 9027/ 9090 | https://www.molbiodiagnostics.com/
1.1.4 Truenat services
106701 | Truelab and Trueprep Installation and training | 1 unit | N/A | Installation include: 1. Location check for ideal work area; Check for quality of power outlet especially earthing 2. Unboxing and check equipment for damages; Installation of the equipment, inputting user information and checking for correct settings 3. Pairing of printer and attaching of antennas, checking communication settings including GPRS settings if opted for; Updating profiles 4. Performing soft calibration and Running positive and negative controls/ samples 5. Confirming proper functioning of the devices; Completing installation report In-country travel and labor of company's local agent Training include: 1.Patient’s sample handling and Good laboratory practices/Decontamination Protocol. 2.Operation and Maintenance of Trueprep AUTO and Truelab Analyzer in detail. 3.Preparation of different types of samples (viz. Sputum, Blood, Serum Plasma etc.) using Trueprep Reagents and its extraction in Trueprep AUTO. 4.Loading of elute on Truenat chip and result analysis; Print the results using Truelab Micro PCR Printer via Bluetooth Connection; 5. Proper Disposal of used material after completion of test | Molbio Diagnostics |  | 500 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | Truelab and Trueprep Installation and training | N/A | Not available
106956 | Truelab UNO DX Workstation Extended warranty 1 year (2400000723) | 1 unit | N/A | Extended warranty / Comprehensive Maintenance Contract (CMC) for the Truelab UNO DX Workstation for 1 year includes: · Remote assistance / visit of service engineer and as necessary, · Repair and replacement of parts. If a machine is not repairable at site, a standby machine will be provided. If the machine needs to be replaced, free replacement will be provided at Molbio's cost. · In-country travel and labor of company's local agent · Calibration chips and material which is used as service items | Molbio Diagnostics |  | 800 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | 2400000723 | N/A | Not available
106957 | Truelab DUO Workstation Extended warranty 1 year (2400000724) | 1 unit | N/A | Extended warranty / Comprehensive Maintenance Contract (CMC) for the Truelab DUO Workstation for 1 year includes: · Remote assistance / visit of service engineer and as necessary, · Repair and replacement of parts. If a machine is not repairable at site, a standby machine will be provided. If the machine needs to be replaced, free replacement will be provided at Molbio's cost. · In-country travel and labor of company's local agent · Calibration chips and material which is used as service items | Molbio Diagnostics |  | 1120 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | 2400000724 | N/A | Not available
106958 | Truelab Quattro Workstation Extended warranty 1 year (2400000725 | 1 unit | N/A | Extended warranty / Comprehensive Maintenance Contract (CMC) for the Truelab QUATTRO Workstation for 1 year includes: · Remote assistance / visit of service engineer and as necessary, · Repair and replacement of parts. If a machine is not repairable at site, a standby machine will be provided. If the machine needs to be replaced, free replacement will be provided at Molbio's cost. · In-country travel and labor of company's local agent · Calibration chips and material which is used as service items | Molbio Diagnostics |  | 1440 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | 2400000725 | N/A | Not available
1.1.5 TB-LAMP equipment
106632 | HumaLoop T (961000) | 1 unit | N/A | With HumaLoop T the sensitive and accurate molecular detection of Mycobacterium Tuberculosis bacterium complex can be facilitated on a consolidated platform in order to replace smear microscopy. Preprogrammed incubator with built-in heating block, reaction block, timer and fluorescence detection unit for lysis, amplification, enzyme inactivation and visual result interpretation with1 year warranty. The operator can run a maximum of 16 reactions in one run of HumaLoop T. Of these 16 tests, 14 are patient samples and 2 are controls. Each cycle takes 1.5 hours, therefore 5 cycles may be run per day in a well-functioning laboratory (up to 70 patient samples per day); The recommended UPS to use wit this item is: UPS 700 VA with its battery pack. | Human |  | 3467.54 | USD | 0 | General Cargo | 10.5 | 37.5 | 39 | 35 | 0.0512 | 961000 | 90275000 | https://www.human.de/en/products/Molecular%20DX/TB-LAMP/Instruments/HumaLoop%20T
106633 | HumaTurb C+ HumaTurb A (963200) | 1 unit | N/A | Programmable system with preinstalled settings, amplification and control units for isothermal amplification, enzyme inactivation and real-time turbidity detection; · High throughput system, for reference laboratories (up to 96 tests per run), as an alternative to HumaLoop T. This item includes: 1 Unit C + 1 Unit A | Human |  | 18855.81 | USD | 0 | General Cargo | 5.4 | 52 | 35 | 23.5 | 0.0445 | 963200 | 90275000 | https://www.human.de/en/products/Molecular%20DX/TB-LAMP/Instruments/HumaTurb%20System
106636 | HumaHeat (964000) | 1 unit | N/A | HumaHeat allows an easy lysis of Loopamp™ heating tubes with the corresponding heating block. Incubation temperatures and times can be set up individually. Heater designed for the sample lysis with Loopamp PURE DNA Extraction Kit before performing the LAMP reaction with HumaTurb system. | Human |  | 1138.46 | USD | 0 | General Cargo | 2.6 | 31 | 19.5 | 23.5 | 0.0142 | 964000 | 84198998 | https://www.human.de/en/products/Molecular%20DX/TB-LAMP/Instruments/HumaTurb%20System#unfold-c4029
106663 | HuMax ITA (980000) | 1 unit | N/A | HuMax ITA benchtop micro centrifuge for incubation, mixing and centrifugation of Loopamp reaction tubes. | Human |  | 858.73 | USD | 0 | General Cargo | 2.5 | 28 | 20.5 | 17.5 | 0.01 | 980000 | 84211920 | https://www.human.de/en/products/Molecular%20DX/TB-LAMP/Instruments/HumaTurb%20System#unfold-c4029
1.1.6 TB-LAMP reagents
106634 | Loopamp MTBC Detection Kit, 2x48 tests (972000) | 1 kit | 12 months | Qualitative in vitro diagnostic test for the detection of MTB complex /96 tests9. The amplification under isothermal conditions enables within a short time the fluorescence or turbidity based detection of MTB complex; A detection kit of 96 tests and an Extraction kit of 90 tests will provide the appropriate number of reagents for 6 cycles comprising 84 patient samples plus 6 negative controls and 6 positive controls. This item includes: 2 boxes X 48 Tests | Human |  | 219 | USD | 0 | General Cargo | 0.19 | 19.5 | 17.5 | 7.5 | 0.0026 | 972000 | 38221900 | https://www.human.de/detection-assay
106635 | Loopamp PURE DNA Extraction Kit, 90 tests (970000) | 1 kit | 21 months | DNA extraction and isolation method for sputum samples through lysis/heating, removal of inhibitors and elution of DNA directly into the amplification tube; A detection kit of 96 tests and an Extraction kit of 90 tests will provide the appropriate number of reagents for 6 cycles comprising 84 patient samples plus 6 negative controls and 6 positive controls. This item includes: a kit that allows for 90 extractions. | Human |  | 275 | USD | 0 | General Cargo | 1.9 | 23 | 21.5 | 32 | 0.0158 | 970000 | 38221900 | https://www.human.de/dna-extraction
1.1.7 TB-LAMP consumables
106643 | Pipette-60 set for TB LAMP (971000) | 1 set | N/A | Pipette and filter tips for the Transfer of sputum samples for TB LAMP or about 4x96 tests. · Pipette – 60 set: Aspiration performance 60 µl +/_ 4.0%; 35 mm in diameter x 192 mm I length, Approx. 26 gram · Filtered tips for 4 x 96 tips/box, 7.5 mm in diameter x 75 mm in length, approx. 1.0 gram/tip. | Human |  | 42 | USD | 0 | General Cargo | 0.89 | 13 | 19 | 20.5 | 0.005064 | 971000 | 39269097 | https://www.human.de/dna-extraction
1.1.8 TB-LAMP services
106637 | TB-LAMP Installation | 1 unit | N/A | The installation of the HumaLoop T, HumaHeat, and HumaTurb A and C. On-site inspection according to site preparation check list, and installation of system(s). · Testing incubator of HumaHeat and turbidimetric functionality of HumaTurb C + A; · Testing incubators and fluorescence module of HumaLoop according to Operation Qualification (OQ); · All equipment and test kits needed for OQ; · Issue of Installation and Operation Qualification certificates, and start of basic warranty; Travel and accommodation expenses, where applicable, are not included and will be charged separately by HUMAN to the recipient. | Human |  | 1535.75 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | 210 | N/A | https://www.human.de/service-support
106638 | TB-LAMP Training | 1 unit | N/A | End user training for TB- LAMP is for 2 days and covers: · Introduction to Loopamp technology; Familiarization with products and procedures; · Preparation of materials, workspace and training documentation; · Performing TB assay; Results interpretation & troubleshooting; · Examination of trainees. Each trainee receives a certificate after passing the exam; · Test kits to be used for training purposes are not included in the pricing and should be provided by the trainee or can be ordered via the assigned partner of HUMAN; Travel and accommodation expenses, where applicable, are not included and will be charged separately by HUMAN to the recipient. | Human |  | 296.48 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | 210 | N/A | https://www.human.de/service-support
106639 | TB-LAMP Maintenance for 1 year | 1 unit | N/A | 1 year checking on-site conditions: · Extensive cleaning of working area, instrument surface, heating blocks and fluorescence reading unit according to maintenance check list; Inspection and validation of UPS and power connection; · Temperature check and verifying the incubation and reaction block are reaching the target temperature according to maintenance check list; · Performance of maintenance test kit to assure the instrument is working properly according to Operation Qualification (OQ); · Needed equipment and test kit to perform the OQ; issue of OQ certificate; Travel and accommodation expenses, where applicable, are not included and will be charged separately by HUMAN to the recipient. | Human |  | 1464.59 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | 150 | N/A | https://www.human.de/service-support
106640 | TB-LAMP Maintenance with warranty extension for 1 year | 1 unit | N/A | Warranty extension for 1 year (on-site) if the maintenance is done before the first warranty period ends and a copy of the filled maintenance check list is returned to HUMAN: · Extensive cleaning of working area, instrument surface, heating blocks and fluorescence reading unit according to maintenance check list; Inspection and validation of UPS and power connection; · Temperature check and verifying the incubation and reaction block are reaching the target temperature according to maintenance check list; · Performance of maintenance test kit to assure the instrument is working properly according to Operation Qualification (OQ); · Needed equipment and test kit to perform the OQ; Issue of OQ certificate; Travel and accommodation expenses, where applicable, are not included and will be charged separately by HUMAN to the recipient. | Human |  | 1986.38 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | 150 | N/A | https://www.human.de/service-support
1.1.11 GeneXpert Equipment (10 color)
GXII-1-D-10C | GeneXpert II, 1-Module, Desktop, 10-color | 1 unit | N/A | GeneXpert 1-module instrument with desktop computer. 10-Color assay targets as many as 8 genes for detection of M. tuberculosis and resistance to as many as 6 drugs: isoniazid, fluoroquinolones, amikacin, kanamycin, capreomycin and ethionamide. Multiplexing technology for an expanding test menu including Xpert MTB/XDR. This new 10-color system could also be connected to an existing GeneXpert System and replace an old computer. If that’s the case, make sure to add a separate DAISYKIT part to your new system order. This item includes: cables, Scanner, Desktop, Monitor, GeneXpert Software; WIN10 Computer, Barcode Scanner, latest GeneXpert Software and all accessories. The initial period of the warranty is 24 months from the invoice date of the system. To benefit from the second year of the initial warranty, an Xpert Check kit must be procured separately to calibrate the system at the end of the first year of the initial warranty. The initial warranty includes: Remote assistance; Onsite interventions are not included and will be quoted separately by the assigned authorized service provider and Preventative Maintenance is handled by the customer. Remote installation is included and onsite installation is available upon request. Installation & Training and Super-User training programme is Free of Charge for countries with Cepheid’s representation or Cepheid’s Authorized Service Providers (ASP). Requests for Super User trainings can be sent to: training@cepheid.com or to any Cepheid representative in country or region. | Cepheid HBDC |  | 9420 | USD | 0 | General Cargo | 24 | 69 | 50 | 68 | 0.2346 | GXII-1-D-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXII-2-D-10C | GeneXpert II, 2-Modules, Desktop, 10-color | 1 unit | N/A | GeneXpert 2-modules instrument with desktop computer. 10-Color assay targets as many as 8 genes for detection of M. tuberculosis and resistance to as many as 6 drugs: isoniazid, fluoroquinolones, amikacin, kanamycin, capreomycin and ethionamide. Multiplexing technology for an expanding test menu including Xpert MTB/XDR. This new 10-color system could also be connected to an existing GeneXpert System and replace an old computer. If that’s the case, make sure to add a separate DAISYKIT part to your new system order. This item includes: cables, Scanner, Desktop, Monitor, GeneXpert Software; WIN10 Computer, Barcode Scanner, latest GeneXpert Software and all accessories. The initial period of the warranty is 24 months from the invoice date of the system. To benefit from the second year of the initial warranty, an Xpert Check kit must be procured separately to calibrate the system at the end of the first year of the initial warranty. The initial warranty includes: Remote assistance; Onsite interventions are not included and will be quoted separately by the assigned authorized service provider and Preventative Maintenance is handled by the customer. Remote installation is included and onsite installation is available upon request. Installation & Training and Super-User training programme is Free of Charge for countries with Cepheid’s representation or Cepheid’s Authorized Service Providers (ASP). Requests for Super User trainings can be sent to: training@cepheid.com or to any Cepheid representative in country or region. | Cepheid HBDC |  | 12530 | USD | 0 | General Cargo | 25.5 | 69 | 50 | 66 | 0.2277 | GXII-2-D-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXII-2-D-10C (Trade in IV-2) | GeneXpert II, 2-Modules, D10C, trade in for R1 GeneXpert IV-2 | 1 unit | N/A | Trade in old R1 GeneXpert IV-2 for a new R2 II-2-Desktop-10C. First Generation of GeneXpert systems (R1 model) are eligible to the Cepheid’s Global Access Trade-In Program that offers a new GeneXpert system (R2 model) equipped with 10-color technology at a preferred price Cepheid recycles the GeneXpert instrument (R1 model) with 6-color technology. The original computer with all patient data and results is not collected for data privacy reasons. The technical specifications and the items included with this product are exactly the same as : GXII-2-D-10C GeneXpert II, 2-Modules, Desktop, 10-color plus the Trade in program specifications detailed above. For a successful trade in, we kindly request you to share the serial numbers of the R1 GeneXpert system(s) with your GDF focal point (Country Supply Officer) as they will be used to place the Trade-in order. | Cepheid HBDC |  | 10502 | USD | 0 | General Cargo | 25.5 | 69 | 50 | 66 | 0.2277 | GXII-2-D-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXIV-2-D-10C | GeneXpert IV, 2-Modules, Desktop, 10-color | 1 unit | N/A | GeneXpert 2-modules instrument with desktop computer. 10-Color assay targets as many as 8 genes for detection of M. tuberculosis and resistance to as many as 6 drugs: isoniazid, fluoroquinolones, amikacin, kanamycin, capreomycin and ethionamide. Multiplexing technology for an expanding test menu including Xpert MTB/XDR. This new 10-color system could also be connected to an existing GeneXpert System and replace an old computer. If that’s the case, make sure to add a separate DAISYKIT part to your new system order. This item includes: cables, Scanner, Desktop, Monitor, GeneXpert Software; WIN10 Computer, Barcode Scanner, latest GeneXpert Software and all accessories. The initial period of the warranty is 24 months from the invoice date of the system. To benefit from the second year of the initial warranty, an Xpert Check kit must be procured separately to calibrate the system at the end of the first year of the initial warranty. The initial warranty includes: Remote assistance; Onsite interventions are not included and will be quoted separately by the assigned authorized service provider and Preventative Maintenance is handled by the customer. Remote installation is included and onsite installation is available upon request. Installation & Training and Super-User training programme is Free of Charge for countries with Cepheid’s representation or Cepheid’s Authorized Service Providers (ASP). Requests for Super User trainings can be sent to: training@cepheid.com or to any Cepheid representative in country or region. | Cepheid HBDC |  | 12780 | USD | 0 | General Cargo | 25.5 | 69 | 50 | 66 | 0.2277 | GXIV-2-D-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXIV-2-D-10C (Trade-in IV-2) | GeneXpert IV, 2-Modules, D10C, trade-in for R1 GeneXpert IV-2 | 1 unit | N/A | Trade in old R1 GeneXpert IV-2 for a new R2 IV-2-Desktop-10C First Generation of GeneXpert systems (R1 model) are eligible to the Cepheid’s Global Access Trade-In Program that offers a new GeneXpert system (R2 model) equipped with 10-color technology at a preferred price Cepheid recycles the GeneXpert instrument (R1 model) with 6-color technology. The original computer with all patient data and results is not collected for data privacy reasons. The technical specifications and the items included with this product are exactly the same as : GXIV-2-D-10C GeneXpert IV, 2-Modules, Desktop, 10-color plus the Trade in program specifications detailed above. For a successful trade in, we kindly request you to share the serial numbers of the R1 GeneXpert system(s) with your GDF focal point (Country Supply Officer) as they will be used to place the Trade-in order. | Cepheid HBDC |  | 10502 | USD | 0 | General Cargo | 25.5 | 69 | 50 | 66 | 0.2277 | GXIV-2-D-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXIV-4-D-10C | GeneXpert IV, 4-Modules, Desktop, 10-color | 1 unit | N/A | GeneXpert 4-modules instrument with desktop computer. 10-Color assay targets as many as 8 genes for detection of M. tuberculosis and resistance to as many as 6 drugs: isoniazid, fluoroquinolones, amikacin, kanamycin, capreomycin and ethionamide. Multiplexing technology for an expanding test menu including Xpert MTB/XDR. This new 10-color system could also be connected to an existing GeneXpert System and replace an old computer. If that’s the case, make sure to add a separate DAISYKIT part to your new system order. This item includes: cables, Scanner, Desktop, Monitor, GeneXpert Software; WIN10 Computer, Barcode Scanner, latest GeneXpert Software and all accessories. The initial period of the warranty is 24 months from the invoice date of the system. To benefit from the second year of the initial warranty, an Xpert Check kit must be procured separately to calibrate the system at the end of the first year of the initial warranty. The initial warranty includes: Remote assistance; Onsite interventions are not included and will be quoted separately by the assigned authorized service provider and Preventative Maintenance is handled by the customer. Remote installation is included and onsite installation is available upon request. Installation & Training and Super-User training programme is Free of Charge for countries with Cepheid’s representation or Cepheid’s Authorized Service Providers (ASP). Requests for Super User trainings can be sent to: training@cepheid.com or to any Cepheid representative in country or region. | Cepheid HBDC |  | 19000 | USD | 0 | General Cargo | 30 | 69 | 50 | 66 | 0.2277 | GXIV-2-D-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXIV-4-D-10C (Trade-in IV-2) | GeneXpert IV, 4-Modules, D10C, trade-in for R1 GeneXpert IV-2 | 1 unit | N/A | Trade in old R1 GeneXpert IV-2 for a new R2 IV-4-Desktop-10C First Generation of GeneXpert systems (R1 model) are eligible to the Cepheid’s Global Access Trade-In Program that offers a new GeneXpert system (R2 model) equipped with 10-color technology at a preferred price Cepheid recycles the GeneXpert instrument (R1 model) with 6-color technology. The original computer with all patient data and results is not collected for data privacy reasons. The technical specifications and the items included with this product are exactly the same as : GXIV-4-D-10C GeneXpert IV, 4-Modules, Desktop, 10-color plus the Trade in program specifications detailed above. For a successful trade in, we kindly request you to share the serial numbers of the R1 GeneXpert system(s) with your GDF focal point (Country Supply Officer) as they will be used to place the Trade-in order. | Cepheid HBDC |  | 15150 | USD | 0 | General Cargo | 30 | 69 | 50 | 66 | 0.2277 | GXIV-2-D-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXIV-4-D-10C (Trade-in IV-4) | GeneXpert IV, 4-Modules, D10C, trade-in for R1 GeneXpert IV-4 | 1 unit | N/A | Trade in old R1 GeneXpert IV-4 for a new R2 IV-4-Desktop-10C. First Generation of GeneXpert systems (R1 model) are eligible to the Cepheid’s Global Access Trade-In Program that offers a new GeneXpert system (R2 model) equipped with 10-color technology at a preferred price Cepheid recycles the GeneXpert instrument (R1 model) with 6-color technology. The original computer with all patient data and results is not collected for data privacy reasons. The technical specifications and the items included with this product are exactly the same as : GXIV-4-D-10C GeneXpert IV, 4-Modules, Desktop, 10-color plus the Trade in program specifications detailed above. For a successful trade in, we kindly request you to share the serial numbers of the R1 GeneXpert system(s) with your GDF focal point (Country Supply Officer) as they will be used to place the Trade-in order. | Cepheid HBDC |  | 15100 | USD | 0 | General Cargo | 30 | 69 | 50 | 66 | 0.2277 | GXIV-2-D-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXXVI-8-D-10C | GeneXpert XVI, 8-Modules, Desktop, 10-color | 1 unit | N/A | GeneXpert 8-modules instrument with desktop computer. 10-Color assay targets as many as 8 genes for detection of M. tuberculosis and resistance to as many as 6 drugs: isoniazid, fluoroquinolones, amikacin, kanamycin, capreomycin and ethionamide. Multiplexing technology for an expanding test menu including Xpert MTB/XDR. This new 10-color system could also be connected to an existing GeneXpert System and replace an old computer. If that’s the case, make sure to add a separate DAISYKIT part to your new system order. This item includes: cables, Scanner, Desktop, Monitor, GeneXpert Software; WIN10 Computer, Barcode Scanner, latest GeneXpert Software and all accessories. The initial period of the warranty is 24 months from the invoice date of the system. To benefit from the second year of the initial warranty, an Xpert Check kit must be procured separately to calibrate the system at the end of the first year of the initial warranty. The initial warranty includes: Remote assistance; Onsite interventions are not included and will be quoted separately by the assigned authorized service provider and Preventative Maintenance is handled by the customer. Remote installation is included and onsite installation is available upon request. Installation & Training and Super-User training programme is Free of Charge for countries with Cepheid’s representation or Cepheid’s Authorized Service Providers (ASP). Requests for Super User trainings can be sent to: training@cepheid.com or to any Cepheid representative in country or region. | Cepheid HBDC |  | 43000 | USD | 0 | General Cargo | 96 | 81 | 71 | 139 | 0.79939 | GXXVI-8-D-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXXVI-8-D-10C (Trade-in IV-2) | GeneXpert XVI, 8-Modules, D10C, trade-in for R1 GeneXpert IV-2 | 1 unit | N/A | Trade in old R1 GeneXpert IV-2 for a new R2 XVI-8-Desktop-10C. First Generation of GeneXpert systems (R1 model) are eligible to the Cepheid’s Global Access Trade-In Program that offers a new GeneXpert system (R2 model) equipped with 10-color technology at a preferred price Cepheid recycles the GeneXpert instrument (R1 model) with 6-color technology. The original computer with all patient data and results is not collected for data privacy reasons. The technical specifications and the items included with this product are exactly the same as : GXXVI-8-D-10C GeneXpert XVI, 8-Modules, Desktop, 10-color plus the Trade in program specifications detailed above. For a successful trade in, we kindly request you to share the serial numbers of the R1 GeneXpert system(s) with your GDF focal point (Country Supply Officer) as they will be used to place the Trade-in order. | Cepheid HBDC |  | 33400 | USD | 0 | General Cargo | 96 | 81 | 71 | 139 | 0.79939 | GXXVI-8-D-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXXVI-8-D-10C (Trade-in IV-4) | GeneXpert XVI, 8-Modules, D10C, trade-in for R1 GeneXpert IV-4 | 1 unit | N/A | Trade in old R1 GeneXpert IV-4 for a new R2 XVI-8-Desktop-10C. First Generation of GeneXpert systems (R1 model) are eligible to the Cepheid’s Global Access Trade-In Program that offers a new GeneXpert system (R2 model) equipped with 10-color technology at a preferred price Cepheid recycles the GeneXpert instrument (R1 model) with 6-color technology. The original computer with all patient data and results is not collected for data privacy reasons. The technical specifications and the items included with this product are exactly the same as : GXXVI-8-D-10C GeneXpert XVI, 8-Modules, Desktop, 10-color plus the Trade in program specifications detailed above. For a successful trade in, we kindly request you to share the serial numbers of the R1 GeneXpert system(s) with your GDF focal point (Country Supply Officer) as they will be used to place the Trade-in order. | Cepheid HBDC |  | 34550 | USD | 0 | General Cargo | 96 | 81 | 71 | 139 | 0.79939 | GXXVI-8-D-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXXVI-8-D-10C (Trade-in XVI-8) | GeneXpert XVI, 8-Modules, D10C, trade-in for R1 GeneXpert XVI-8 | 1 unit | N/A | Trade in old R1 GeneXpert XVI-8 for a new R2 XVI-8-Dektop-10C. First Generation of GeneXpert systems (R1 model) are eligible to the Cepheid’s Global Access Trade-In Program that offers a new GeneXpert system (R2 model) equipped with 10-color technology at a preferred price Cepheid recycles the GeneXpert instrument (R1 model) with 6-color technology. The original computer with all patient data and results is not collected for data privacy reasons. The technical specifications and the items included with this product are exactly the same as : GXXVI-8-D-10C GeneXpert XVI, 8-Modules, Desktop, 10-color plus the Trade in program specifications detailed above. For a successful trade in, we kindly request you to share the serial numbers of the R1 GeneXpert system(s) with your GDF focal point (Country Supply Officer) as they will be used to place the Trade-in order. | Cepheid HBDC |  | 34700 | USD | 0 | General Cargo | 96 | 81 | 71 | 139 | 0.79939 | GXXVI-8-D-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXXVI-16-D-10C | GeneXpert XVI, 16-Modules, Desktop, 10-color | 1 unit | N/A | GeneXpert 16-modules instrument with desktop computer. 10-Color assay targets as many as 8 genes for detection of M. tuberculosis and resistance to as many as 6 drugs: isoniazid, fluoroquinolones, amikacin, kanamycin, capreomycin and ethionamide. Multiplexing technology for an expanding test menu including Xpert MTB/XDR. This new 10-color system could also be connected to an existing GeneXpert System and replace an old computer. If that’s the case, make sure to add a separate DAISYKIT part to your new system order. This item includes: cables, Scanner, Desktop, Monitor, GeneXpert Software; WIN10 Computer, Barcode Scanner, latest GeneXpert Software and all accessories. The initial period of the warranty is 24 months from the invoice date of the system. To benefit from the second year of the initial warranty, an Xpert Check kit must be procured separately to calibrate the system at the end of the first year of the initial warranty. The initial warranty includes: Remote assistance; Onsite interventions are not included and will be quoted separately by the assigned authorized service provider and Preventative Maintenance is handled by the customer. Remote installation is included and onsite installation is available upon request. Installation & Training and Super-User training programme is Free of Charge for countries with Cepheid’s representation or Cepheid’s Authorized Service Providers (ASP). Requests for Super User trainings can be sent to: training@cepheid.com or to any Cepheid representative in country or region. | Cepheid HBDC |  | 71850 | USD | 0 | General Cargo | 108 | 81 | 71 | 137 | 0.78789 | GXXVI-16-D-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXXVI-16-D-10C (rade-in IV-2) | GeneXpert XVI, 16-Modules, D10C, trade-in for R1 GeneXpert IV-2 | 1 unit | N/A | Trade in old R1 GeneXpert IV-2 for a new XVI-16-Desktop-10C. First Generation of GeneXpert systems (R1 model) are eligible to the Cepheid’s Global Access Trade-In Program that offers a new GeneXpert system (R2 model) equipped with 10-color technology at a preferred price Cepheid recycles the GeneXpert instrument (R1 model) with 6-color technology. The original computer with all patient data and results is not collected for data privacy reasons. The technical specifications and the items included with this product are exactly the same as : GXXVI-16-D-10C GeneXpert XVI, 16-Modules, Desktop, 10-color plus the Trade in program specifications detailed above. For a successful trade in, we kindly request you to share the serial numbers of the R1 GeneXpert system(s) with your GDF focal point (Country Supply Officer) as they will be used to place the Trade-in order. | Cepheid HBDC |  | 56480 | USD | 0 | General Cargo | 108 | 81 | 71 | 137 | 0.78789 | GXXVI-16-D-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXXVI-16-D-10C (Trade-in IV-4) | GeneXpert XVI, 16-Modules, D10C, trade-in for R1 GeneXpert IV-4 | 1 unit | N/A | Trade in old R1 GeneXpert IV-4 for a new R2 XVI-16-Desktop-10C. First Generation of GeneXpert systems (R1 model) are eligible to the Cepheid’s Global Access Trade-In Program that offers a new GeneXpert system (R2 model) equipped with 10-color technology at a preferred price Cepheid recycles the GeneXpert instrument (R1 model) with 6-color technology. The original computer with all patient data and results is not collected for data privacy reasons. The technical specifications and the items included with this product are exactly the same as : GXXVI-16-D-10C GeneXpert XVI, 16-Modules, Desktop, 10-color plus the Trade in program specifications detailed above. For a successful trade in, we kindly request you to share the serial numbers of the R1 GeneXpert system(s) with your GDF focal point (Country Supply Officer) as ithey will be used to place the Trade-in order. | Cepheid HBDC |  | 55480 | USD | 0 | General Cargo | 108 | 81 | 71 | 137 | 0.78789 | GXXVI-16-D-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXXVI-16-D-10C (Trade-in XVI-8) | GeneXpert XVI, 16-Modules, D10C, trade-in for R1 GeneXpert XVI-8 | 1 unit | N/A | Trade in old R1 GeneXpert XVI-8 for a new R2 XVI-16-Desktop-10C. First Generation of GeneXpert systems (R1 model) are eligible to the Cepheid’s Global Access Trade-In Program that offers a new GeneXpert system (R2 model) equipped with 10-color technology at a preferred price Cepheid recycles the GeneXpert instrument (R1 model) with 6-color technology. The original computer with all patient data and results is not collected for data privacy reasons. The technical specifications and the items included with this product are exactly the same as : GXXVI-16-D-10C GeneXpert XVI, 16-Modules, Desktop, 10-color plus the Trade in program specifications detailed above. For a successful trade in, we kindly request you to share the serial numbers of the R1 GeneXpert system(s) with your GDF focal point (Country Supply Officer) as they will be used to place the Trade-in order. | Cepheid HBDC |  | 57072.5 | USD | 0 | General Cargo | 108 | 81 | 71 | 137 | 0.78789 | GXXVI-16-D-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXXVI-16-D-10C (Trade-in XVI-16) | GeneXpert XVI, 16-Modules D10C, trade-in for R1 GeneXpert XVI-16 | 1 unit | N/A | Trade in old R1 GeneXpert XVI-16 for a new R2 XVI-16-Desktop-10C. First Generation of GeneXpert systems (R1 model) are eligible to the Cepheid’s Global Access Trade-In Program that offers a new GeneXpert system (R2 model) equipped with 10-color technology at a preferred price Cepheid recycles the GeneXpert instrument (R1 model) with 6-color technology. The original computer with all patient data and results is not collected for data privacy reasons. The technical specifications and the items included with this product are exactly the same as : GXXVI-16-D-10C GeneXpert XVI, 16-Modules, Desktop, 10-color plus the Trade in program specifications detailed above. For a successful trade in, we kindly request you to share the serial numbers of the R1 GeneXpert system(s) with your GDF focal point (Country Supply Officer) as they will be used to place the Trade-in order. | Cepheid HBDC |  | 56665 | USD | 0 | General Cargo | 108 | 81 | 71 | 137 | 0.78789 | GXXVI-16-D-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXII-1-L-10C | GeneXpert II, 1-Module, Laptop, 10-color | 1 unit | N/A | GeneXpert 1-module instrument with laptop computer. 10-Color assay targets as many as 8 genes for detection of M. tuberculosis and resistance to as many as 6 drugs: isoniazid, fluoroquinolones, amikacin, kanamycin, capreomycin and ethionamide. Multiplexing technology for an expanding test menu including Xpert MTB/XDR. This new 10-color system could also be connected to an existing GeneXpert System and replace an old computer. If that’s the case, make sure to add a separate DAISYKIT part to your new system order. This item includes: cables, Scanner, Laptop, Monitor, GeneXpert Software; WIN10 Computer, Barcode Scanner, latest GeneXpert Software and all accessories. The initial period of the warranty is 24 months from the invoice date of the system. To benefit from the second year of the initial warranty, an Xpert Check kit must be procured separately to calibrate the system at the end of the first year of the initial warranty. The initial warranty includes: Remote assistance; Onsite interventions are not included and will be quoted separately by the assigned authorized service provider and Preventative Maintenance is handled by the customer. Remote installation is included and onsite installation is available upon request. Installation & Training and Super-User training programme is Free of Charge for countries with Cepheid’s representation or Cepheid’s Authorized Service Providers (ASP). Requests for Super User trainings can be sent to: training@cepheid.com or to any Cepheid representative in country or region. | Cepheid HBDC |  | 9920 | USD | 0 | General cargo UN 3481 | 13 | 59 | 50 | 45 | 0.13275 | GXII-1-L-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXII-2-L-10C | GeneXpert II, 2-Modules, Laptop, 10-color | 1 unit | N/A | GeneXpert 2-modules instrument with laptop computer. 10-Color assay targets as many as 8 genes for detection of M. tuberculosis and resistance to as many as 6 drugs: isoniazid, fluoroquinolones, amikacin, kanamycin, capreomycin and ethionamide. Multiplexing technology for an expanding test menu including Xpert MTB/XDR. This new 10-color system could also be connected to an existing GeneXpert System and replace an old computer. If that’s the case, make sure to add a separate DAISYKIT part to your new system order. This item includes: cables, Scanner, Laptop, Monitor, GeneXpert Software; WIN10 Computer, Barcode Scanner, latest GeneXpert Software and all accessories. The initial period of the warranty is 24 months from the invoice date of the system. To benefit from the second year of the initial warranty, an Xpert Check kit must be procured separately to calibrate the system at the end of the first year of the initial warranty. The initial warranty includes: Remote assistance; Onsite interventions are not included and will be quoted separately by the assigned authorized service provider and Preventative Maintenance is handled by the customer. Remote installation is included and onsite installation is available upon request. Installation & Training and Super-User training programme is Free of Charge for countries with Cepheid’s representation or Cepheid’s Authorized Service Providers (ASP). Requests for Super User trainings can be sent to: training@cepheid.com or to any Cepheid representative in country or region. | Cepheid HBDC |  | 13030 | USD | 0 | General cargo UN 3481 | 14 | 59 | 50 | 43 | 0.12685 | GXII-2-L-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXII-2-L-10C (Trade-in IV-2) | GeneXpert II, 2-Modules, L10C trade-in for R1 GeneXpert IV-2 | 1 unit | N/A | Trade in old R1 GeneXpert IV-2 for a new R2 II-2-Laptop-10C. First Generation of GeneXpert systems (R1 model) are eligible to the Cepheid’s Global Access Trade-In Program that offers a new GeneXpert system (R2 model) equipped with 10-color technology at a preferred price Cepheid recycles the GeneXpert instrument (R1 model) with 6-color technology. The original computer with all patient data and results is not collected for data privacy reasons. The technical specifications and the items included with this product are exactly the same as : GXII-2-L-10C GeneXpert II, 2-Modules, Laptop, 10-color plus the Trade in program specifications detailed above. For a successful trade in, we kindly request you to share the serial numbers of the R1 GeneXpert system(s) with your GDF focal point (Country Supply Officer) as they will be used to place the Trade-in order. | Cepheid HBDC |  | 10952 | USD | 0 | General cargo UN 3481 | 14 | 59 | 50 | 43 | 0.12685 | GXII-2-L-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXIV-2-L-10C | GeneXpert IV, 2-Modules, Laptop, 10-color | 1 unit | N/A | GeneXpert 2-modules instrument with laptop computer. 10-Color assay targets as many as 8 genes for detection of M. tuberculosis and resistance to as many as 6 drugs: isoniazid, fluoroquinolones, amikacin, kanamycin, capreomycin and ethionamide. Multiplexing technology for an expanding test menu including Xpert MTB/XDR. This new 10-color system could also be connected to an existing GeneXpert System and replace an old computer. If that’s the case, make sure to add a separate DAISYKIT part to your new system order. This item includes: cables, Scanner, Laptop, Monitor, GeneXpert Software; WIN10 Computer, Barcode Scanner, latest GeneXpert Software and all accessories. The initial period of the warranty is 24 months from the invoice date of the system. To benefit from the second year of the initial warranty, an Xpert Check kit must be procured separately to calibrate the system at the end of the first year of the initial warranty. The initial warranty includes: Remote assistance; Onsite interventions are not included and will be quoted separately by the assigned authorized service provider and Preventative Maintenance is handled by the customer. Remote installation is included and onsite installation is available upon request. Installation & Training and Super-User training programme is Free of Charge for countries with Cepheid’s representation or Cepheid’s Authorized Service Providers (ASP). Requests for Super User trainings can be sent to: training@cepheid.com or to any Cepheid representative in country or region. | Cepheid HBDC |  | 13280 | USD | 0 | General cargo UN 3481 | 18 | 61 | 47 | 48 | 0.14 | GXIV-2-L-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXIV-2-L-10C (Trade-in IV-2) | GeneXpert IV, 2-Modules, L10C, trade-in for R1 GeneXpert IV-2 | 1 unit | N/A | Trade in old R1 GeneXpert IV-2 for a new R2 IV-2-Laptop-10C. First Generation of GeneXpert systems (R1 model) are eligible to the Cepheid’s Global Access Trade-In Program that offers a new GeneXpert system (R2 model) equipped with 10-color technology at a preferred price Cepheid recycles the GeneXpert instrument (R1 model) with 6-color technology. The original computer with all patient data and results is not collected for data privacy reasons. The technical specifications and the items included with this product are exactly the same as : GXIV-2-L-10C GeneXpert IV, 2-Modules, Laptop, 10-color plus the Trade in program specifications detailed above. For a successful trade in, we kindly request you to share the serial numbers of the R1 GeneXpert system(s) with your GDF focal point (Country Supply Officer) as they will be used to place the Trade-in order. | Cepheid HBDC |  | 10952 | USD | 0 | General cargo UN 3481 | 18 | 61 | 47 | 48 | 0.14 | GXIV-2-L-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXIV-4-L-10C | GeneXpert IV, 4-Modules, Laptop, 10-color | 1 unit | N/A | GeneXpert 4-modules instrument with laptop computer. 10-Color assay targets as many as 8 genes for detection of M. tuberculosis and resistance to as many as 6 drugs: isoniazid, fluoroquinolones, amikacin, kanamycin, capreomycin and ethionamide. Multiplexing technology for an expanding test menu including Xpert MTB/XDR. This new 10-color system could also be connected to an existing GeneXpert System and replace an old computer. If that’s the case, make sure to add a separate DAISYKIT part to your new system order. This item includes: cables, Scanner, Laptop, Monitor, GeneXpert Software; WIN10 Computer, Barcode Scanner, latest GeneXpert Software and all accessories. The initial period of the warranty is 24 months from the invoice date of the system. To benefit from the second year of the initial warranty, an Xpert Check kit must be procured separately to calibrate the system at the end of the first year of the initial warranty. The initial warranty includes: Remote assistance; Onsite interventions are not included and will be quoted separately by the assigned authorized service provider and Preventative Maintenance is handled by the customer. Remote installation is included and onsite installation is available upon request. Installation & Training and Super-User training programme is Free of Charge for countries with Cepheid’s representation or Cepheid’s Authorized Service Providers (ASP). Requests for Super User trainings can be sent to: training@cepheid.com or to any Cepheid representative in country or region. | Cepheid HBDC |  | 19500 | USD | 0 | General cargo UN 3481 | 23 | 61 | 47 | 66 | 0.19 | GXIV-4-L-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXIV-4-L-10C (Trade-in IV-2) | GeneXpert IV, 4-Modules, L10C, trade-in for R1 GeneXpert IV-2 | 1 unit | N/A | Trade in old R1 GeneXpert IV-2 for a new R2 IV-4-Laptop-10C. First Generation of GeneXpert systems (R1 model) are eligible to the Cepheid’s Global Access Trade-In Program that offers a new GeneXpert system (R2 model) equipped with 10-color technology at a preferred price Cepheid recycles the GeneXpert instrument (R1 model) with 6-color technology. The original computer with all patient data and results is not collected for data privacy reasons. The technical specifications and the items included with this product are exactly the same as : GXIV-4-L-10C GeneXpert IV, 4-Modules, Laptop, 10-color plus the Trade in program specifications detailed above. For a successful trade in, we kindly request you to share the serial numbers of the R1 GeneXpert system(s) with your GDF focal point (Country Supply Officer) as they will be used to place the Trade-in order. | Cepheid HBDC |  | 15575 | USD | 0 | General cargo UN 3481 | 23 | 61 | 47 | 66 | 0.19 | GXIV-4-L-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXIV-4-L-10C (Trade-in IV-4) | GeneXpert IV, 4-Modules, L10C, trade-in for R1 GeneXpert IV-4 | 1 unit | N/A | Trade in old R1 GeneXpert IV-4 for a new R2 IV-4-Laptop-10C. First Generation of GeneXpert systems (R1 model) are eligible to the Cepheid’s Global Access Trade-In Program that offers a new GeneXpert system (R2 model) equipped with 10-color technology at a preferred price Cepheid recycles the GeneXpert instrument (R1 model) with 6-color technology. The original computer with all patient data and results is not collected for data privacy reasons. The technical specifications and the items included with this product are exactly the same as : GXIV-4-L-10C GeneXpert IV, 4-Modules, Laptop, 10-color plus the Trade in program specifications detailed above. For a successful trade in, we kindly request you to share the serial numbers of the R1 GeneXpert system(s) with your GDF focal point (Country Supply Officer) as they will be used to place the Trade-in order. | Cepheid HBDC |  | 1550 | USD | 0 | General cargo UN 3481 | 23 | 61 | 47 | 66 | 0.19 | GXIV-4-L-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXXVI-8-L-10C | GeneXpert XVI, 8-Modules, Laptop, 10-color | 1 unit | N/A | GeneXpert 8-modules instrument with laptop computer. 10-Color assay targets as many as 8 genes for detection of M. tuberculosis and resistance to as many as 6 drugs: isoniazid, fluoroquinolones, amikacin, kanamycin, capreomycin and ethionamide. Multiplexing technology for an expanding test menu including Xpert MTB/XDR. This new 10-color system could also be connected to an existing GeneXpert System and replace an old computer. If that’s the case, make sure to add a separate DAISYKIT part to your new system order. This item includes: cables, Scanner, Laptop, Monitor, GeneXpert Software; WIN10 Computer, Barcode Scanner, latest GeneXpert Software and all accessories. The initial period of the warranty is 24 months from the invoice date of the system. To benefit from the second year of the initial warranty, an Xpert Check kit must be procured separately to calibrate the system at the end of the first year of the initial warranty. The initial warranty includes: Remote assistance; Onsite interventions are not included and will be quoted separately by the assigned authorized service provider and Preventative Maintenance is handled by the customer. Remote installation is included and onsite installation is available upon request. Installation & Training and Super-User training programme is Free of Charge for countries with Cepheid’s representation or Cepheid’s Authorized Service Providers (ASP). Requests for Super User trainings can be sent to: training@cepheid.com or to any Cepheid representative in country or region. | Cepheid HBDC |  | 43500 | USD | 0 | General cargo UN 3481 | 88 | 84 | 71 | 103 | 0.61429 | GXXVI-8-L-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXXVI-8-L-10C (Trade-in IV-2) | GeneXpert XVI, 8-Modules, L10C, trade-in for R1 GeneXpert IV-2 | 1 unit | N/A | Trade in old R1 GeneXpert IV-2 for a new R2 XVI-8-Laptop-10C. First Generation of GeneXpert systems (R1 model) are eligible to the Cepheid’s Global Access Trade-In Program that offers a new GeneXpert system (R2 model) equipped with 10-color technology at a preferred price Cepheid recycles the GeneXpert instrument (R1 model) with 6-color technology. The original computer with all patient data and results is not collected for data privacy reasons. The technical specifications and the items included with this product are exactly the same as : GXXVI-8-L-10C GeneXpert XVI, 8-Modules, Laptop, 10-color plus the Trade in program specifications detailed above. For a successful trade in, we kindly request you to share the serial numbers of the R1 GeneXpert system(s) with your GDF focal point (Country Supply Officer) as they will be used to place the Trade-in order. | Cepheid HBDC |  | 33800 | USD | 0 | General cargo UN 3481 | 88 | 84 | 71 | 103 | 0.61429 | GXXVI-8-L-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXXVI-8-L-10C (Trade-in IV-4) | GeneXpert XVI, 8-Modules, L10C, trade-in for R1 GeneXpert IV-4 | 1 unit | N/A | Trade in old R1 GeneXpert IV-4 for a new R2 XVI-8-Laptop-10C. First Generation of GeneXpert systems (R1 model) are eligible to the Cepheid’s Global Access Trade-In Program that offers a new GeneXpert system (R2 model) equipped with 10-color technology at a preferred price Cepheid recycles the GeneXpert instrument (R1 model) with 6-color technology. The original computer with all patient data and results is not collected for data privacy reasons. The technical specifications and the items included with this product are exactly the same as : GXXVI-8-L-10C GeneXpert XVI, 8-Modules, Laptop, 10-color plus the Trade in program specifications detailed above. For a successful trade in, we kindly request you to share the serial numbers of the R1 GeneXpert system(s) with your GDF focal point (Country Supply Officer) as they will be used to place the Trade-in order. | Cepheid HBDC |  | 34975 | USD | 0 | General cargo UN 3481 | 88 | 84 | 71 | 103 | 0.61429 | GXXVI-8-L-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXXVI-8-L-10C (Trade-in XVI-8) | GeneXpert XVI, 8-Modules, L10C, trade-in for R1 GeneXpert XVI-8 | 1 unit | N/A | Trade in old R1 GeneXpert XVI-8 for a new R2 XVI-8-Laptop-10C. First Generation of GeneXpert systems (R1 model) are eligible to the Cepheid’s Global Access Trade-In Program that offers a new GeneXpert system (R2 model) equipped with 10-color technology at a preferred price Cepheid recycles the GeneXpert instrument (R1 model) with 6-color technology. The original computer with all patient data and results is not collected for data privacy reasons. The technical specifications and the items included with this product are exactly the same as : GXXVI-8-L-10C GeneXpert XVI, 8-Modules, Laptop, 10-color plus the Trade in program specifications detailed above. For a successful trade in, we kindly request you to share the serial numbers of the R1 GeneXpert system(s) with your GDF focal point (Country Supply Officer) as they will be used to place the Trade-in order. | Cepheid HBDC |  | 35150 | USD | 0 | General cargo UN 3481 | 88 | 84 | 71 | 103 | 0.61429 | GXXVI-8-L-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXXVI-16-L-10C | GeneXpert XVI, 16-Modules, Laptop, 10-color | 1 unit | N/A | GeneXpert 16-modules instrument with laptop computer. 10-Color assay targets as many as 8 genes for detection of M. tuberculosis and resistance to as many as 6 drugs: isoniazid, fluoroquinolones, amikacin, kanamycin, capreomycin and ethionamide. Multiplexing technology for an expanding test menu including Xpert MTB/XDR. This new 10-color system could also be connected to an existing GeneXpert System and replace an old computer. If that’s the case, make sure to add a separate DAISYKIT part to your new system order. This item includes: cables, Scanner, Laptop, Monitor, GeneXpert Software; WIN10 Computer, Barcode Scanner, latest GeneXpert Software and all accessories. The initial period of the warranty is 24 months from the invoice date of the system. To benefit from the second year of the initial warranty, an Xpert Check kit must be procured separately to calibrate the system at the end of the first year of the initial warranty. The initial warranty includes: Remote assistance; Onsite interventions are not included and will be quoted separately by the assigned authorized service provider and Preventative Maintenance is handled by the customer. Remote installation is included and onsite installation is available upon request. Installation & Training and Super-User training programme is Free of Charge for countries with Cepheid’s representation or Cepheid’s Authorized Service Providers (ASP). Requests for Super User trainings can be sent to: training@cepheid.com or to any Cepheid representative in country or region. | Cepheid HBDC |  | 72350 | USD | 0 | General cargo UN 3481 | 98 | 84 | 71 | 101 | 0.60236 | GXXVI-16-L-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXXVI-16-L-10C (Trade-in IV-2) | GeneXpert XVI, 16-Modules, L10C, trade-in for R1 GeneXpert IV-2 | 1 unit | N/A | Trade in old R1 GeneXpert IV-2 for a new R2 XVI-16-Laptop-10C. First Generation of GeneXpert systems (R1 model) are eligible to the Cepheid’s Global Access Trade-In Program that offers a new GeneXpert system (R2 model) equipped with 10-color technology at a preferred price Cepheid recycles the GeneXpert instrument (R1 model) with 6-color technology. The original computer with all patient data and results is not collected for data privacy reasons. The technical specifications and the items included with this product are exactly the same as : GXXVI-16-L-10C GeneXpert XVI, 16-Modules, Laptop, 10-color plus the Trade in program specifications detailed above. For a successful trade in, we kindly request you to share the serial numbers of the R1 GeneXpert system(s) with your GDF focal point (Country Supply Officer) as they will be used to place the Trade-in order. | Cepheid HBDC |  | 56880 | USD | 0 | General cargo UN 3481 | 98 | 84 | 71 | 101 | 0.60236 | GXXVI-16-L-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXXVI-16-L-10C (Trade-in IV-4) | GeneXpert XVI, 16-Modules, L10C, trade-in for R1 GeneXpert IV-4 | 1 unit | N/A | Trade in old R1 GeneXpert IV-4 for a new R2 XVI-16-Laptop-10C. First Generation of GeneXpert systems (R1 model) are eligible to the Cepheid’s Global Access Trade-In Program that offers a new GeneXpert system (R2 model) equipped with 10-color technology at a preferred price Cepheid recycles the GeneXpert instrument (R1 model) with 6-color technology. The original computer with all patient data and results is not collected for data privacy reasons. The technical specifications and the items included with this product are exactly the same as : GXXVI-16-L-10C GeneXpert XVI, 16-Modules, Laptop, 10-color plus the Trade in program specifications detailed above. For a successful trade in, we kindly request you to share the serial numbers of the R1 GeneXpert system(s) with your GDF focal point (Country Supply Officer) as they will be used to place the Trade-in order. | Cepheid HBDC |  | 55880 | USD | 0 | General cargo UN 3481 | 98 | 84 | 71 | 101 | 0.60236 | GXXVI-16-L-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXXVI-16-L-10C (Trade-in XVI-8) | GeneXpert XVI, 16-Modules, L10C, trade-in for R1 GeneXpert XVI-8 | 1 unit | N/A | Trade in old R1 GeneXpert XVI-8 for a new R2 XVI-16-Laptop-10C. First Generation of GeneXpert systems (R1 model) are eligible to the Cepheid’s Global Access Trade-In Program that offers a new GeneXpert system (R2 model) equipped with 10-color technology at a preferred price Cepheid recycles the GeneXpert instrument (R1 model) with 6-color technology. The original computer with all patient data and results is not collected for data privacy reasons. The technical specifications and the items included with this product are exactly the same as : GXXVI-16-L-10C GeneXpert XVI, 16-Modules, Laptop, 10-color plus the Trade in program specifications detailed above. For a successful trade in, we kindly request you to share the serial numbers of the R1 GeneXpert system(s) with your GDF focal point (Country Supply Officer) as they will be used to place the Trade-in order. | Cepheid HBDC |  | 57497.5 | USD | 0 | General cargo UN 3481 | 98 | 84 | 71 | 101 | 0.60236 | GXXVI-16-L-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
GXXVI-16-L-10C (Trade-in XVI-16) | GeneXpert XVI, 16-Modules L10C, trade-in for R1 GeneXpert XVI-16 | 1 unit | N/A | Trade in old R1 GeneXpert XVI-16 for a new R2 XVI-16-Laptop-10C. First Generation of GeneXpert systems (R1 model) are eligible to the Cepheid’s Global Access Trade-In Program that offers a new GeneXpert system (R2 model) equipped with 10-color technology at a preferred price Cepheid recycles the GeneXpert instrument (R1 model) with 6-color technology. The original computer with all patient data and results is not collected for data privacy reasons. The technical specifications and the items included with this product are exactly the same as : GXXVI-16-L-10C GeneXpert XVI, 16-Modules, Laptop, 10-color plus the Trade in program specifications detailed above. For a successful trade in, we kindly request you to share the serial numbers of the R1 GeneXpert system(s) with your GDF focal point (Country Supply Officer) as they will be used to place the Trade-in order. | Cepheid HBDC |  | 57115 | USD | 0 | General cargo UN 3481 | 98 | 84 | 71 | 101 | 0.60236 | GXXVI-16-L-10C | 90275000 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
1.1.12 GeneXpert Satellite (10 color)
900-0793 | GeneXpert II, 1-Module, Satellite, 10-color | 1 unit | N/A | GeneXpert 1-module satellite equipment with 10-Color assay targets as many as eight genes for detection of M. tuberculosis and resistance to as many as six drugs: isoniazid, fluoroquinolones, amikacin, kanamycin, capreomycin and ethionamide. Multiplexing technology for an expanding test menu including Xpert MTB/XDR. Remote installation included and onsite installation available upon request. A separate order it is required for the DAISYKIT and one Software kit Dx 6.2 or higher per site for Xpert MTB/XDR testing. Also the computer should be check as follows: • Computers running WINXP operating systems must be replaced with a new WIN10 computer. • Cepheid recommends replacing computers running WIN7 with new WIN10 computers, due to Microsoft’s discontinuation of WIN7 support in 2021. This item includes: 1 x GeneXpert stellite (withouth the computer and the barcode scanner) | Cepheid HBDC |  | 6420 | USD | 0 | General Cargo | 6 | 42 | 41 | 32 | 0.074088 | 900-0793 | 90279050 | https://www.cepheid.com/en-CH/systems/genexpert-family-of-systems/genexpert-system.html
DAISYKIT | GeneXpert Daisy Chain Accessory Kit | 1 unit | N/A | The daisy kit allows connection of 2 instruments to a same computer. A satellite instrument is a GeneXpert without the computer and barcode scanner. With the use of the Daisy kit this instrument can be connected to another GeneXpert system. this option can be considered when the existing 6 color GeneXpert systems are in place but would like additional capacity with Xpert MTB/XDR testing capabilities This item includes:1 x NETGEAR 5 port network switch and adapter. Used to connect to 2 GeneXpert Instruments (22 modules maximum) (800-0217); 2 x Ethernet cable. Used to connect 2 instruments (if more than 2 instruments are used, more cables need to be purchased (700-0555); 1 x power plug adaptor. Only used with the optional Intl UPS 850-0387 - discard if not used- (100-5735); 1 x power plug adaptor type B to type A. Only used with the optional North America UPS 850-0386 -discard if not used- (100-5330) and 1x universal power adaptor. Replaces the NETGEAR power adapter to allow for use outside of North America (100-5576) | Cepheid HBDC |  | 75 | USD | 0 | General cargo | 5 | 41 | 39 | 29 | 0.046371 | DAISYKIT | 90275090 | Not available
1.1.13 GeneXpert reagents (6 and 10 color)
GXMTB/RIF-ULTRA-50 | GeneXpert MTB/RIF Ultra kit of 50 tests | 50 tests | 14 months | Xpert MTB/RIF Ultra Assay for detection of MTB and rifampin-resistance associated mutations of the rpoB gene, where MTB complex is detected. • Results in <80 minutes; • Increased accuracy of Rifampicin results; Improved detection of mixed infections; • Compatibility with multiple specimen types: sputum sediment and fresh sputum; • Storage conditions: Store the Xpert MTB/RIF Ultra Assay cartridges at 2–28 °C; This item includes: The Xpert MTB/RIF Ultra Assay kits contains sufficient reagents to process 50 samples/spcimens. Xpert MTB/RIF Ultra Assay Cartridges with Integrated reaction Tubes (50), Sample Reagent Bottles (50), disposable Transfer Pipettes (60) and CD (1). | Cepheid HBDC |  | 398.5 | USD | 0 | Cool cargo (+2-+8°C) UN 2920 – CLASS 8(3)) in excepted quantity (EQ) | 3.35 | 32.5 | 25 | 18 | 0.01 | GXMTB/RIF-ULTRA-50 | 3822900090 | https://www.cepheid.com/en/tests/Critical-Infectious-Diseases/Xpert-MTB-RIF-Ultra
GXMTB/XDR-10 | GeneXpert MTB/XDR, 10 tests | 1 box | 8 months | Xpert MTB/XDR assay for the detection of extensively drug resistant (XDR) MTB complex DNA in unprocessed sputum samples. The Xpert MTB/XDR kit contains sufficient reagents to process 10 patient or quality-control specimens. • Results in <90 minutes; • Storage conditions: Store the Xpert MTB/RIF Ultra Assay cartridges at 2–28 °C; This item includes: The Xpert MTB/XDR kit contains sufficient reagents to process 10 patient/specimens. Xpert MTB/XDR cartridges with Integrated reaction Tubes (10), Sample Reagent Bottles (10), disposable Transfer Pipettes (10) and CD (1). | Cepheid HBDC |  | 149 | USD | 0 | Cool cargo (+2-+8°C) UN 2920 – CLASS 8(3)) in excepted quantity (EQ) | 0.7 | 16 | 24 | 7 | 0.003 | GXMTB/XDR-10 | 3822900090 | https://www.cepheid.com/en/tests/Critical-Infectious-Diseases/Xpert-MTB-XDR
PCK-MTB/XDR-50-L | Pack Xpert 50 MTB/XDR tests and GX4 10C Laptop | 1 pack | 8 months | GeneXpert System customized packages with 5 boxes (50 cartridges) of Xpert MTB XDR and 1 GeneXpert IV-4 10C Laptop System. | Cepheid HBDC |  | 17500 | USD | 0 | Cool cargo (+2-+8°C) UN 2920 – CLASS 8(3)) in excepted quantity (EQ) | 27 | 120 | 100 | 60 | 0.72 | PCK-MTB/XDR-50-L | 90275000 | Not available
PCK-MTB/XDR-50-D | Pack Xpert 50 MTB/XDR tests and GX4 10C Desktop | 1 pack | 8 months | GeneXpert System customized packages with 5 boxes (50 cartridges) of Xpert MTB XDR and 1 GeneXpert IV-4 10C Desktop System. | Cepheid HBDC |  | 17500 | USD | 0 | Cool cargo (+2-+8°C) UN 2920 – CLASS 8(3)) in excepted quantity (EQ) | 37 | 120 | 100 | 60 | 0.72 | PCK-MTB/XDR-50-D | 90275000 | Not available
1.1.14 GeneXpert consumables (6 and 10 color)
XPERTCHECKCE- 5 | GeneXpert Check calibration kit | 1 box | 12 months | GeneXpert Calibration Packs for GeneXpert System. It will be used to verify the optical & thermal systems and will perform a series of system level tests to ensure full functionality of modules for the coming year. The Xpert Check needs to be performed every year. For a site with a GXXVI 16 module instrument, 4 packs are required. This item contains: 5x Xpert check cartridges with integrated reaction tubes (each one with 1 bead per cartridges, 1.0 ml 1x per cartridges, 4x I-CORE Lens cleaning brush, 1x PI/Software (P/N 950-0413) and 1x Data CD. | Cepheid HBDC |  | 450 | USD | 0 | Cool cargo (+2-+8°C) UN 2920 – CLASS 8(3)) in excepted quantity (EQ) | 0.9 | 37 | 27 | 19 | 0.018981 | XPERTCHECKCE- 5 | 3822900000 | https://www.cepheid.com/Package%20Insert%20Files/c9b3f4-301-1520-Rev-C-Xpert-Calibration-PI---US-Version-Print-.pdf
900-0508 | GeneXpert new module for 6-color equipment | 1 unit | N/A | GeneXpert new module for 6-color equipment | Cepheid HBDC |  | 3360 | USD | 0 | General Cargo | 2.5 | 36 | 36 | 20 | 0.02592 | 900-0508 | 90275000 | Not available
900-0491RS | GeneXpert Certified Pre-Owned Module | 1 unit | N/A | A refurbished module is provided to swap out a module requiring calibration or repair. To be purchased when a GeneXpert is not under initial or extended warranty. This item includes: 1 x GeneXpert pre-owned module for 6-color system | Cepheid HBDC |  | 900 | USD | 0 | General Cargo | 2.5 | 36 | 36 | 20 | 0.02592 | 900-0491RS | 90275000 | Not available
700-4510 | GeneXpert dust filter for 6-color equipment GXIV-4 (4 modules) | 1 unit | N/A | GeneXpert dust filter for 6-color equipment GXIV-4 (4 modules), for use in dusty environments to prevent dust from entering the GeneXpert system. | Cepheid HBDC |  | 350 | USD | 0 | General Cargo | 0.1 | 10 | 6 | 2 | 0.00012 | 700-4510 | 90275000 | Not available
900-0807 | GeneXpert, Module Kit - 10-color | 1 unit | N/A | This item includes: The GeneXpert Module A separate order it is required for Software kit Dx 6.2 or higher per site for Xpert MTB/XDR testing. All modules of the same system should be converted (hybrid systems of 6- and 10-color modules are not supported). i.e. for a GeneXpert IV, 4 new modules should be ordered. There is no swap or discount option for the existing 6-color modules Also the computer should be check as follows: • Computers running WINXP operating systems must be replaced with a new WIN10 computer. • Cepheid recommends replacing computers running WIN7 with new WIN10 computers, due to Microsoft’s discontinuation of WIN7 support in 2021. | Cepheid HBDC |  | 3860 | USD | 0 | General cargo | 2.4 | 40 | 38 | 15 | 0.0228 | 900-0807 | 90275090 | Not available
CMPTR-DT-GXDX | GeneXpert Desktop Computer WIN 10 | 1 unit | N/A | Desktop Computer with windows 10 for GeneXpert system. If the desktop computer of GeneXpert Systems is running the Microsoft Windows 7 or Windows XP operating system, the replacement by a Desktop computer with Windows 10 Operating Systems should be done. This item includes: 1x Desktop Computer with windows 10 for GeneXpert system (only the desktop computer itself is included). | Cepheid HBDC |  | 1325 | USD | 0 | General cargo | 6.9 | 53 | 39 | 21 | 0.0434 | CMPTR-DT-GXDX | 8471500000 | Not available
CMPTR-LT-GXDX | GeneXpert Laptop Computer WIN 10 | 1 unit | N/A | Laptop Computer with windows 10 for GeneXpert system. If the laptop computer of GeneXpert Systems is running the Microsoft Windows 7 or Windows XP operating system, the replacement by a Laptop computer with Windows 10 Operating Systems should be done. This item includes: 1x Laptop Computer with windows 10 for GeneXpert system (only the laptop computer itself is included). | Cepheid HBDC |  | 1780 | USD | 0 | General cargo | 3.5 | 52 | 34 | 25 | 0.0442 | CMPTR-LT-GXDX | 8471500000 | Not available
1.1.15 GeneXpert services (same for 6 and 10 color)
ACCESS_CARE (Tier 1) | GeneXpert AccessCare 3.0 Xpert SLA Tier 1 | 1 unit | N/A | Only for countries with a signed Service Level Agreements (SLA) with Cepheid for Acccesscare. Accesscare is based on the number of modules in the country and the number of Xpert tests volumes purchased/committed per year. The Tier 1 surcharge applies when the consumption of tests/module/day is between 3.32 to 5 tests. The Tier price that will apply is defined by Cepheid and the country during the assessment process of country eligibility to AccessCare The SLA includes: 1. Provision of monthly reporting Key Performance Indicators (KPIs) results against targets: system availability (>=95%), module availability (>=95%) and timely critical part replacement (>=90%) 2. Onsite Repair and Support Services only by Cepheid or Cepheid's Authorized Service Provider (As described in the SLA with the specific country). 3. Online system of monitoring service request status 4. Replacement of any failed 6 Color Module under Access Care for a new generation 10 Color module is free of charge for new generation systems and with computer under Microsoft Windows 10 Operating System only (any replacement of computer running Windows 7 or Windows XP operating system might be subject to additional cost). Not applicable to first generation of GeneXpert (R1 model). 5. The to the 10-Color Hybrid systems approach allowing a GeneXpert system to run a combination of 6-color (6C) and 10-color (10C) modules, called “10C Hybrids”. A mandatory prerequisite for a Hybrid system is a computer running the Windows 10 Operating System and if not yet installed, to upgrade Dx software to version 6.5 or higher (free of charge for a computer running the Microsoft Windows 10 Operating System. | Cepheid HBDC |  | 0.9 | USD | 0 | N/A | N/A | N/A | N/A | N/A | N/A | ACCESS_CARE (Tier 1) | N/A | Not available
ACCESS_CARE (Tier 2) | GeneXpert AccessCare 3.0 Xpert SLA Tier 2 | 1 unit | N/A | Only for countries with a signed Service Level Agreements (SLA) with Cepheid for Acccesscare. Accesscare is based on the number of modules in the country and the number of Xpert tests volumes purchased/committed per year. The Tier 2 surcharge applies when the consumption of tests/module/day is is minimum of 2.28 tests. The Tier price that will apply is defined by Cepheid and the country during the assessment process of country eligibility to AccessCare The SLA includes: 1. Provision of monthly reporting Key Performance Indicators (KPIs) results against targets: system availability (>=95%), module availability (>=95%) and timely critical part replacement (>=90%) 2. Onsite Repair and Support Services only by Cepheid or Cepheid's Authorized Service Provider (As described in the SLA with the specific country). 3. Online system of monitoring service request status 4. Replacement of any failed 6 Color Module under Access Care for a new generation 10 Color module is free of charge for new generation systems and with computer under Microsoft Windows 10 Operating System only (any replacement of computer running Windows 7 or Windows XP operating system might be subject to additional cost). Not applicable to first generation of GeneXpert (R1 model). 5. The to the 10-Color Hybrid systems approach allowing a GeneXpert system to run a combination of 6-color (6C) and 10-color (10C) modules, called “10C Hybrids”. A mandatory prerequisite for a Hybrid system is a computer running the Windows 10 Operating System and if not yet installed, to upgrade Dx software to version 6.5 or higher (free of charge for a computer running the Microsoft Windows 10 Operating System. | Cepheid HBDC |  | 1.2 | USD | 0 | N/A | N/A | N/A | N/A | N/A | N/A | ACCESS_CARE (Tier 2) | N/A | Not available
ACCESS_CARE (Tier 3) | GeneXpert AccessCare 3.0 Xpert SLA Tier 3 | 1 unit | N/A | Only for countries with a signed Service Level Agreements (SLA) with Cepheid for Acccesscare. Accesscare is based on the number of modules in the country and the number of Xpert tests volumes purchased/committed per year. The Tier 3 surcharge applies when the consumption of tests/module/day is is minimum of 1.72 tests. The Tier price that will apply is defined by Cepheid and the country during the assessment process of country eligibility to AccessCare The SLA includes: 1. Provision of monthly reporting Key Performance Indicators (KPIs) results against targets: system availability (>=95%), module availability (>=95%) and timely critical part replacement (>=90%) 2. Onsite Repair and Support Services only by Cepheid or Cepheid's Authorized Service Provider (As described in the SLA with the specific country). 3. Online system of monitoring service request status 4. Replacement of any failed 6 Color Module under Access Care for a new generation 10 Color module is free of charge for new generation systems and with computer under Microsoft Windows 10 Operating System only (any replacement of computer running Windows 7 or Windows XP operating system might be subject to additional cost). Not applicable to first generation of GeneXpert (R1 model). 5. The to the 10-Color Hybrid systems approach allowing a GeneXpert system to run a combination of 6-color (6C) and 10-color (10C) modules, called “10C Hybrids”. A mandatory prerequisite for a Hybrid system is a computer running the Windows 10 Operating System and if not yet installed, to upgrade Dx software to version 6.5 or higher (free of charge for a computer running the Microsoft Windows 10 Operating System. | Cepheid HBDC |  | 1.5 | USD | 0 | N/A | N/A | N/A | N/A | N/A | N/A | ACCESS_CARE (Tier 3) | N/A | Not available
WXPLUS02Y1 (GXII-2) | 1 Year Warranty Extension Plus for GXII-2 | 1 unit | N/A | The Warranty Extension Plus has the same SLA than AccessCare but it is a One-time payment per year for one or multiple systems via one contract. The price of the Warranty Extension Plus is per type of system and per duration. The Extension Plus includes: 1. Provision of monthly reporting Key Performance Indicators (KPIs) results against targets: system availability (>=95%), module availability (>=95%) and timely critical part replacement (>=90%) 2. Onsite Repair and Support Services only by Cepheid or Cepheid's Authorized Service Provider (As described in the SLA with the specific country). 3. Online system of monitoring service request status 4. Replacement of any failed 6 Color Module under Access Care for a new generation 10 Color module is free of charge for new generation systems and with computer under Microsoft Windows 10 Operating System only (any replacement of computer running Windows 7 or Windows XP operating system might be subject to additional cost). Not applicable to first generation of GeneXpert (R1 model). 5. The Warranty Extension Plus only apply to 10-color GeneXpert system 6. The to the 10-Color Hybrid systems approach allowing a GeneXpert system to run a combination of 6-color (6C) and 10-color (10C) modules, called “10C Hybrids”. A mandatory prerequisite for a Hybrid system is a computer running the Windows 10 Operating System and if not yet installed, to upgrade Dx software to version 6.5 or higher (free of charge for a computer running the Microsoft Windows 10 Operating System. | Cepheid HBDC |  | 2484 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WXPLUS02Y1(GXII-2) | N/A | Not available
WXPLUS02Y3 (GXII-2) | 3 Years Warranty Extension Plus for GXII-2 | 1 unit | N/A | The Warranty Extension Plus has the same SLA than AccessCare but it is a One-time payment per year for one or multiple systems via one contract. The price of the Warranty Extension Plus is per type of system and per duration. The Extension Plus includes: 1. Provision of monthly reporting Key Performance Indicators (KPIs) results against targets: system availability (>=95%), module availability (>=95%) and timely critical part replacement (>=90%) 2. Onsite Repair and Support Services only by Cepheid or Cepheid's Authorized Service Provider (As described in the SLA with the specific country). 3. Online system of monitoring service request status 4. Replacement of any failed 6 Color Module under Access Care for a new generation 10 Color module is free of charge for new generation systems and with computer under Microsoft Windows 10 Operating System only (any replacement of computer running Windows 7 or Windows XP operating system might be subject to additional cost). Not applicable to first generation of GeneXpert (R1 model). 5. The Warranty Extension Plus only apply to 10-color GeneXpert system 6. The to the 10-Color Hybrid systems approach allowing a GeneXpert system to run a combination of 6-color (6C) and 10-color (10C) modules, called “10C Hybrids”. A mandatory prerequisite for a Hybrid system is a computer running the Windows 10 Operating System and if not yet installed, to upgrade Dx software to version 6.5 or higher (free of charge for a computer running the Microsoft Windows 10 Operating System. | Cepheid HBDC |  | 6895 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WXPLUS02Y3(GXII-2) | N/A | Not available
WXPLUS02Y1 (GXIV-2) | 1 Year Warranty Extension Plus for GXIV-2 | 1 unit | N/A | The Warranty Extension Plus has the same SLA than AccessCare but it is a One-time payment per year for one or multiple systems via one contract. The price of the Warranty Extension Plus is per type of system and per duration. The Extension Plus includes: 1. Provision of monthly reporting Key Performance Indicators (KPIs) results against targets: system availability (>=95%), module availability (>=95%) and timely critical part replacement (>=90%) 2. Onsite Repair and Support Services only by Cepheid or Cepheid's Authorized Service Provider (As described in the SLA with the specific country). 3. Online system of monitoring service request status 4. Replacement of any failed 6 Color Module under Access Care for a new generation 10 Color module is free of charge for new generation systems and with computer under Microsoft Windows 10 Operating System only (any replacement of computer running Windows 7 or Windows XP operating system might be subject to additional cost). Not applicable to first generation of GeneXpert (R1 model). 5. The Warranty Extension Plus only apply to 10-color GeneXpert system 6. The to the 10-Color Hybrid systems approach allowing a GeneXpert system to run a combination of 6-color (6C) and 10-color (10C) modules, called “10C Hybrids”. A mandatory prerequisite for a Hybrid system is a computer running the Windows 10 Operating System and if not yet installed, to upgrade Dx software to version 6.5 or higher (free of charge for a computer running the Microsoft Windows 10 Operating System. | Cepheid HBDC |  | 2484 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WXPLUS02Y1 | N/A | Not available
WXPLUS02Y3 (GXIV-2) | 3 Years Warranty Extension Plus for GXIV-2 | 1 unit | N/A | The Warranty Extension Plus has the same SLA than AccessCare but it is a One-time payment per year for one or multiple systems via one contract. The price of the Warranty Extension Plus is per type of system and per duration. The Extension Plus includes: 1. Provision of monthly reporting Key Performance Indicators (KPIs) results against targets: system availability (>=95%), module availability (>=95%) and timely critical part replacement (>=90%) 2. Onsite Repair and Support Services only by Cepheid or Cepheid's Authorized Service Provider (As described in the SLA with the specific country). 3. Online system of monitoring service request status 4. Replacement of any failed 6 Color Module under Access Care for a new generation 10 Color module is free of charge for new generation systems and with computer under Microsoft Windows 10 Operating System only (any replacement of computer running Windows 7 or Windows XP operating system might be subject to additional cost). Not applicable to first generation of GeneXpert (R1 model). 5. The Warranty Extension Plus only apply to 10-color GeneXpert system 6. The to the 10-Color Hybrid systems approach allowing a GeneXpert system to run a combination of 6-color (6C) and 10-color (10C) modules, called “10C Hybrids”. A mandatory prerequisite for a Hybrid system is a computer running the Windows 10 Operating System and if not yet installed, to upgrade Dx software to version 6.5 or higher (free of charge for a computer running the Microsoft Windows 10 Operating System. | Cepheid HBDC |  | 6895 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WXPLUS02Y3 | N/A | Not available
WXPLUS04Y1 | 1 Year Warranty Extension Plus for GXIV-4 | 1 unit | N/A | The Warranty Extension Plus has the same SLA than AccessCare but it is a One-time payment per year for one or multiple systems via one contract. The price of the Warranty Extension Plus is per type of system and per duration. The Extension Plus includes: 1. Provision of monthly reporting Key Performance Indicators (KPIs) results against targets: system availability (>=95%), module availability (>=95%) and timely critical part replacement (>=90%) 2. Onsite Repair and Support Services only by Cepheid or Cepheid's Authorized Service Provider (As described in the SLA with the specific country). 3. Online system of monitoring service request status 4. Replacement of any failed 6 Color Module under Access Care for a new generation 10 Color module is free of charge for new generation systems and with computer under Microsoft Windows 10 Operating System only (any replacement of computer running Windows 7 or Windows XP operating system might be subject to additional cost). Not applicable to first generation of GeneXpert (R1 model). 5. The Warranty Extension Plus only apply to 10-color GeneXpert system 6. The to the 10-Color Hybrid systems approach allowing a GeneXpert system to run a combination of 6-color (6C) and 10-color (10C) modules, called “10C Hybrids”. A mandatory prerequisite for a Hybrid system is a computer running the Windows 10 Operating System and if not yet installed, to upgrade Dx software to version 6.5 or higher (free of charge for a computer running the Microsoft Windows 10 Operating System. | Cepheid HBDC |  | 3790 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WXPLUS04Y1 | N/A | Not available
WXPLUS04Y3 | 3 Years Warranty Extension Plus for GXIV-4 | 1 unit | N/A | The Warranty Extension Plus has the same SLA than AccessCare but it is a One-time payment per year for one or multiple systems via one contract. The price of the Warranty Extension Plus is per type of system and per duration. The Extension Plus includes: 1. Provision of monthly reporting Key Performance Indicators (KPIs) results against targets: system availability (>=95%), module availability (>=95%) and timely critical part replacement (>=90%) 2. Onsite Repair and Support Services only by Cepheid or Cepheid's Authorized Service Provider (As described in the SLA with the specific country). 3. Online system of monitoring service request status 4. Replacement of any failed 6 Color Module under Access Care for a new generation 10 Color module is free of charge for new generation systems and with computer under Microsoft Windows 10 Operating System only (any replacement of computer running Windows 7 or Windows XP operating system might be subject to additional cost). Not applicable to first generation of GeneXpert (R1 model). 5. The Warranty Extension Plus only apply to 10-color GeneXpert system 6. The to the 10-Color Hybrid systems approach allowing a GeneXpert system to run a combination of 6-color (6C) and 10-color (10C) modules, called “10C Hybrids”. A mandatory prerequisite for a Hybrid system is a computer running the Windows 10 Operating System and if not yet installed, to upgrade Dx software to version 6.5 or higher (free of charge for a computer running the Microsoft Windows 10 Operating System. | Cepheid HBDC |  | 10500 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WXPLUS04Y3 | N/A | Not available
WXPLUS08Y1 | 1 Year Warranty Extension Plus for GXXVI-8 | 1 unit | N/A | The Warranty Extension Plus has the same SLA than AccessCare but it is a One-time payment per year for one or multiple systems via one contract. The price of the Warranty Extension Plus is per type of system and per duration. The Extension Plus includes: 1. Provision of monthly reporting Key Performance Indicators (KPIs) results against targets: system availability (>=95%), module availability (>=95%) and timely critical part replacement (>=90%) 2. Onsite Repair and Support Services only by Cepheid or Cepheid's Authorized Service Provider (As described in the SLA with the specific country). 3. Online system of monitoring service request status 4. Replacement of any failed 6 Color Module under Access Care for a new generation 10 Color module is free of charge for new generation systems and with computer under Microsoft Windows 10 Operating System only (any replacement of computer running Windows 7 or Windows XP operating system might be subject to additional cost). Not applicable to first generation of GeneXpert (R1 model). 5. The Warranty Extension Plus only apply to 10-color GeneXpert system 6. The to the 10-Color Hybrid systems approach allowing a GeneXpert system to run a combination of 6-color (6C) and 10-color (10C) modules, called “10C Hybrids”. A mandatory prerequisite for a Hybrid system is a computer running the Windows 10 Operating System and if not yet installed, to upgrade Dx software to version 6.5 or higher (free of charge for a computer running the Microsoft Windows 10 Operating System. | Cepheid HBDC |  | 7436 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WXPLUS08Y1 | N/A | Not available
WXPLUS08Y3 | 3 Years Warranty Extension Plus for GXXVI-8 | 1 unit | N/A | The Warranty Extension Plus has the same SLA than AccessCare but it is a One-time payment per year for one or multiple systems via one contract. The price of the Warranty Extension Plus is per type of system and per duration. The Extension Plus includes: 1. Provision of monthly reporting Key Performance Indicators (KPIs) results against targets: system availability (>=95%), module availability (>=95%) and timely critical part replacement (>=90%) 2. Onsite Repair and Support Services only by Cepheid or Cepheid's Authorized Service Provider (As described in the SLA with the specific country). 3. Online system of monitoring service request status 4. Replacement of any failed 6 Color Module under Access Care for a new generation 10 Color module is free of charge for new generation systems and with computer under Microsoft Windows 10 Operating System only (any replacement of computer running Windows 7 or Windows XP operating system might be subject to additional cost). Not applicable to first generation of GeneXpert (R1 model). 5. The Warranty Extension Plus only apply to 10-color GeneXpert system 6. The to the 10-Color Hybrid systems approach allowing a GeneXpert system to run a combination of 6-color (6C) and 10-color (10C) modules, called “10C Hybrids”. A mandatory prerequisite for a Hybrid system is a computer running the Windows 10 Operating System and if not yet installed, to upgrade Dx software to version 6.5 or higher (free of charge for a computer running the Microsoft Windows 10 Operating System. | Cepheid HBDC |  | 20846 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WXPLUS08Y3 | N/A | Not available
WXPLUS16Y1 | 1 Year Warranty Extension Plus for GXXVI-16 | 1 unit | N/A | The Warranty Extension Plus has the same SLA than AccessCare but it is a One-time payment per year for one or multiple systems via one contract. The price of the Warranty Extension Plus is per type of system and per duration. The Extension Plus includes: 1. Provision of monthly reporting Key Performance Indicators (KPIs) results against targets: system availability (>=95%), module availability (>=95%) and timely critical part replacement (>=90%) 2. Onsite Repair and Support Services only by Cepheid or Cepheid's Authorized Service Provider (As described in the SLA with the specific country). 3. Online system of monitoring service request status 4. Replacement of any failed 6 Color Module under Access Care for a new generation 10 Color module is free of charge for new generation systems and with computer under Microsoft Windows 10 Operating System only (any replacement of computer running Windows 7 or Windows XP operating system might be subject to additional cost). Not applicable to first generation of GeneXpert (R1 model). 5. The Warranty Extension Plus only apply to 10-color GeneXpert system 6. The to the 10-Color Hybrid systems approach allowing a GeneXpert system to run a combination of 6-color (6C) and 10-color (10C) modules, called “10C Hybrids”. A mandatory prerequisite for a Hybrid system is a computer running the Windows 10 Operating System and if not yet installed, to upgrade Dx software to version 6.5 or higher (free of charge for a computer running the Microsoft Windows 10 Operating System. | Cepheid HBDC |  | 10218 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WXPLUS16Y1 | N/A | Not available
WXPLUS16Y3 | 3 Years Warranty Extension Plus for GXXVI-16 | 1 unit | N/A | The Warranty Extension Plus has the same SLA than AccessCare but it is a One-time payment per year for one or multiple systems via one contract. The price of the Warranty Extension Plus is per type of system and per duration. The Extension Plus includes: 1. Provision of monthly reporting Key Performance Indicators (KPIs) results against targets: system availability (>=95%), module availability (>=95%) and timely critical part replacement (>=90%) 2. Onsite Repair and Support Services only by Cepheid or Cepheid's Authorized Service Provider (As described in the SLA with the specific country). 3. Online system of monitoring service request status 4. Replacement of any failed 6 Color Module under Access Care for a new generation 10 Color module is free of charge for new generation systems and with computer under Microsoft Windows 10 Operating System only (any replacement of computer running Windows 7 or Windows XP operating system might be subject to additional cost). Not applicable to first generation of GeneXpert (R1 model). 5. The Warranty Extension Plus only apply to 10-color GeneXpert system 6. The to the 10-Color Hybrid systems approach allowing a GeneXpert system to run a combination of 6-color (6C) and 10-color (10C) modules, called “10C Hybrids”. A mandatory prerequisite for a Hybrid system is a computer running the Windows 10 Operating System and if not yet installed, to upgrade Dx software to version 6.5 or higher (free of charge for a computer running the Microsoft Windows 10 Operating System. | Cepheid HBDC |  | 27795 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WXPLUS16Y3 | N/A | Not available
WX01RG12 | GeneXpert Edge warranty extension 1 year | 1 unit | N/A | 1-year warranty extension for a 1-module GeneXpert Edge that can be procured anytime. Warranties only cover parts and shipment, and do not include costs for in-country travel or time of service providers, which need to be budgeted separately. The cost of shipping parts is covered by Cepheid (CIP) – additional charges may be incurred to the customer Remote assistance is included Onsite interventions are not included and will be quoted separately by the assigned authorized service provider. Warranty Extensions are recommended for Countries with NO Cepheid presence or representative. This item includes: 1 GeneXpert Edge Preventive Maintenance (PM) kit (XPERTCHECK-CE- 5 Edge - Each kit contains 5 cartridges) for beginning of year. | Cepheid HBDC |  | 1250 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WX01RG12 | N/A | Not available
WX02RG12 | GeneXpert warranty extension 1 year for 2 modules | 1 unit | N/A | 1-year warranty extension for a 2-module GeneXpert system that can be procured anytime. Warranties only cover parts and shipment, and do not include costs for in-country travel or time of service providers, which need to be budgeted separately. The cost of shipping parts is covered by Cepheid (CIP) – additional charges may be incurred to the customer Remote assistance is included Onsite interventions are not included and will be quoted separately by the assigned authorized service provider. Warranty Extensions are recommended for Countries with NO Cepheid presence or representative. This item includes 1 GeneXpert Calibration kit (XPERTCHECK-CE- 5 - Each kit contains 5 cartridges) for year 1. This means that for year 2, another kit must be ordered separately. | Cepheid HBDC |  | 1896 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WX02RG12 | N/A | Not available
WX04RG12 | GeneXpert warranty extension 1 year for 4 modules | 1 unit | N/A | 1-year warranty extension for a 4-module GeneXpert system that can be procured anytime. Warranties only cover parts and shipment, and do not include costs for in-country travel or time of service providers, which need to be budgeted separately. The cost of shipping parts is covered by Cepheid (CIP) – additional charges may be incurred to the customer Remote assistance is included Onsite interventions are not included and will be quoted separately by the assigned authorized service provider. Warranty Extensions are recommended for Countries with NO Cepheid presence or representative. This item includes 1 GeneXpert Calibration kit (XPERTCHECK-CE- 5 - Each kit contains 5 cartridges) for year 1. This means that for year 2, another kit must be ordered separately. | Cepheid HBDC |  | 2898 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WX04RG12 | N/A | Not available
WV08RG12 | GeneXpert warranty extension 1 year for 8 modules | 1 unit | N/A | 1-year warranty extension for a 8-module GeneXpert system that can be procured anytime. Warranties only cover parts and shipment, and do not include costs for in-country travel or time of Cepheid's authorized service providers, which need to be budgeted separately. The cost of shipping parts is covered by Cepheid (CIP) – additional charges may be incurred to the customer Remote assistance is included Onsite interventions are not included and will be quoted separately by the assigned Cepheid's authorized service provider. Warranty Extensions are recommended for Countries with NO Cepheid presence or representative. This item includes 1 GeneXpert Calibration kit (XPERTCHECK-CE- 5 - Each kit contains 5 cartridges) for year 1. This means that for year 2, another kit must be ordered separately. | Cepheid HBDC |  | 5676 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WV08RG12 | N/A | Not available
WV16RG12 | GeneXpert warranty extension 1 year for 16 modules | 1 unit | N/A | 1-year warranty extension for a 16-module GeneXpert system that can be procured anytime. Warranties only cover parts and shipment, and do not include costs for in-country travel or time of service providers, which need to be budgeted separately. The cost of shipping parts is covered by Cepheid (CIP) – additional charges may be incurred to the customer Remote assistance is included Onsite interventions are not included and will be quoted separately by the assigned authorized service provider. Warranty Extensions are recommended for Countries with NO Cepheid presence or representative. This item includes 4 GeneXpert Calibration kits (XPERTCHECK-CE- 5-Each kit contains 5 cartridges) for year 1. This means that for year 2 (of the 2 first years of Warranty, 4 other kits must be ordered separately). | Cepheid HBDC |  | 7800 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WV16RG12 | N/A | Not available
WX01UP36 | GeneXpert Edge warranty extension 3 years with system | 1 unit | N/A | 3-year warranty extension for a 1-module GeneXpert Edge, paid upfront together with the system purchase. Warranties only cover parts and shipment, and do not include costs for in-country travel or time of service providers, which need to be budgeted separately. The cost of shipping parts is covered by Cepheid (CIP) – additional charges may be incurred to the customer Remote assistance is included Onsite interventions are not included and will be quoted separately by the assigned authorized service provider. Warranty Extensions are recommended for Countries with NO Cepheid presence or representative. This item includes: 4 GeneXpert Edge Preventive Maintenance (PM) kit (XPERTCHECK-CE- 5 Edge- Each kit contains 5 cartridges) for the end of the 1st year of warranty (to validate the initial 2 year warranty), for the beginning of year 3, for the beginning of year 4 and for the beginning of year 5. | Cepheid HBDC |  | 3150 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WX01UP36 | N/A | Not available
WX02UP36 | GeneXpert warranty extension 3 years for 2 module -with system | 1 unit | N/A | 3-year warranty extension for a 2-module GeneXpert system, paid upfront together with the system purchase. Warranties only cover parts and shipment, and do not include costs for in-country travel or time of service providers, which need to be budgeted separately. The cost of shipping parts is covered by Cepheid (CIP) – additional charges may be incurred to the customer Remote assistance is included Onsite interventions are not included and will be quoted separately by the assigned authorized service provider. Warranty Extensions are recommended for Countries with NO Cepheid presence or representative. This item includes: 4 GeneXpert Calibration kit (XPERTCHECK-CE-5- Each kit contains 5 cartridges) for the end of the 1st year of warranty (to validate the initial 2 year warranty), for the beginning of year 3, for the beginning of year 4 and for the beginning of year 5. | Cepheid HBDC |  | 4500 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WX02UP36 | N/A | Not available
WX04UP36 | GeneXpert warranty extension 3 years for 4 module -with system | 1 unit | N/A | 3-year warranty extension for a 4-module GeneXpert system, paid upfront together with the system purchase. Warranties only cover parts and shipment, and do not include costs for in-country travel or time of service providers, which need to be budgeted separately. The cost of shipping parts is covered by Cepheid (CIP) – additional charges may be incurred to the customer Remote assistance is included Onsite interventions are not included and will be quoted separately by the assigned authorized service provider. Warranty Extensions are recommended for Countries with NO Cepheid presence or representative. This item includes: 4 GeneXpert Calibration kit (XPERTCHECK-CE-5 - Each kit contains 5 cartridges) for the end of the 1st year of warranty (to validate the initial 2 year warranty), for the beginning of year 3, for the beginning of year 4 and for the beginning of year 5. | Cepheid HBDC |  | 6840 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WX04UP36 | N/A | Not available
WV08UP36 | GeneXpert warranty extension 3 years for 8 module -with system | 1 unit | N/A | 3-year warranty extension for a 8-module GeneXpert system, paid upfront together with the system purchase. Warranties only cover parts and shipment, and do not include costs for in-country travel or time of Cepheid's authorized service providers, which need to be budgeted separately. The cost of shipping parts is covered by Cepheid (CIP) – additional charges may be incurred to the customer Remote assistance is included Onsite interventions are not included and will be quoted separately by the assigned Cepheid's authorized service provider. Warranty Extensions are recommended for Countries with NO Cepheid presence or representative. This item includes: 4 GeneXpert Calibration kit (XPERTCHECK-CE-5 - Each kit contains 5 cartridges) for the end of the 1st year of warranty (to validate the initial 2 year warranty), for the beginning of year 3, for the beginning of year 4 and for the beginning of year 5. | Cepheid HBDC |  | 13878 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WV08UP36 | N/A | Not available
WV16UP36 | GeneXpert warranty extension 3 years for 16 module -with system | 1 unit | N/A | 3-year warranty extension for a 16-module GeneXpert system, paid upfront together with the system purchase. Warranties only cover parts and shipment, and do not include costs for in-country travel or time of service providers, which need to be budgeted separately. The cost of shipping parts is covered by Cepheid (CIP) – additional charges may be incurred to the customer Remote assistance is included Onsite interventions are not included and will be quoted separately by the assigned authorized service provider. Warranty Extensions are recommended for Countries with NO Cepheid presence or representative. This item includes: 16 GeneXpert Calibration kit (XPERTCHECK-CE-5- Each kit contains 5 cartridges) for the end of the 1st year of warranty (to validate the initial 2 year warranty), for the beginning of year 3, for the beginning of year 4 and for the beginning of year 5. | Cepheid HBDC |  | 18504 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WX16UP36 | N/A | Not available
WX01RG36 | GeneXpert Edge warranty extension 3 years purchased later | 1 unit | N/A | 3-year warranty extension for a 1-module GeneXpert Edge. Warranties only cover parts and shipment, and do not include costs for in-country travel or time of service providers, which need to be budgeted separately. The cost of shipping parts is covered by Cepheid (CIP) – additional charges may be incurred to the customer Remote assistance is included Onsite interventions are not included and will be quoted separately by the assigned authorized service provider. Warranty Extensions are recommended for Countries with NO Cepheid presence or representative. This item includes: 3 GeneXpert Preventive Maintenance (PM) kit (XPERTCHECK-CE- 5 Edge -Each kit contains 5 cartridges) for beginning of each of the 3 years | Cepheid HBDC |  | 3400 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WX01RG36 | N/A | Not available
WX02RG36 | GeneXpert warranty extension 3 years 2 modules -purchased later | 1 unit | N/A | 3-year warranty extension for a 2-module GeneXpert system, paid before the end of the 2 first years of warranty. Warranties only cover parts and shipment, and do not include costs for in-country travel or time of service providers, which need to be budgeted separately. The cost of shipping parts is covered by Cepheid (CIP) – additional charges may be incurred to the customer Remote assistance is included Onsite interventions are not included and will be quoted separately by the assigned authorized service provider. Warranty Extensions are recommended for Countries with NO Cepheid presence or representative. This item includes: 3 GeneXpert Calibration kit (XPERTCHECK-CE-5 - Each kit contains 5 cartridges) for beginning of each of the 3 years | Cepheid HBDC |  | 5184 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WX02RG36 | N/A | Not available
WX04RG36 | GeneXpert warranty extension 3 years 4 modules -purchased later | 1 unit | N/A | 3-year warranty extension for a 4-module GeneXpert system, paid before the end of the two first years of warranty. Warranties only cover parts and shipment, and do not include costs for in-country travel or time of service providers, which needs to be budgeted separately. The cost of shipping parts is covered by Cepheid (CIP) – additional charges may be incurred to the customer Remote assistance is included Onsite interventions are not included and will be quoted separately by the assigned authorized service provider. Warranty Extensions are recommended for Countries with NO Cepheid presence or representative. This item includes: 3 GeneXpert Calibration kit (XPERTCHECK-CE-5 - Each kit contains 5 cartridges) for beginning of each of the 3 years. | Cepheid HBDC |  | 7902 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WX04RG36 | N/A | Not available
WV08RG36 | GeneXpert warranty extension 3 years 8 modules -purchased later | 1 unit | N/A | 3-year warranty extension for a 8-module GeneXpert system, paid before the end of the two first years of warranty. Warranties only cover parts and shipment, and do not include costs for in-country travel or time of Cepheid's authorized service providers, which needs to be budgeted separately. The cost of shipping parts is covered by Cepheid (CIP) – additional charges may be incurred to the customer Remote assistance is included Onsite interventions are not included and will be quoted separately by the Cepheid's assigned authorized service provider. Warranty Extensions are recommended for Countries with NO Cepheid presence or representative. This item includes: 3 GeneXpert Calibration kit (XPERTCHECK-CE-5 - Each kit contains 5 cartridges) for beginning of each of the 3 years | Cepheid HBDC |  | 15673 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WV08RG36 | N/A | Not available
WV16RG36 | GeneXpert warranty extension 3 years 16 modules -purchased later | 1 unit | N/A | 3-year warranty extension for a 16-module GeneXpert system, paid before the end of the two first years of warranty. Warranties only cover parts and shipment, and do not include costs for in-country travel or time of service providers, which needs to be budgeted separately. The cost of shipping parts is covered by Cepheid (CIP) – additional charges may be incurred to the customer Remote assistance is included Onsite interventions are not included and will be quoted separately by the assigned authorized service provider. Warranty Extensions are recommended for Countries with NO Cepheid presence or representative. This item includes: 12 GeneXpert Calibration kit (XPERTCHECK-CE-5 - Each kit contains 5 cartridges) for beginning of each of the 3 years. | Cepheid HBDC |  | 20898 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | WX16RG36 | N/A | Not available
1.2 RAPID ANTIGEN DETECTION
1.2.1 Rapid antigen detection tests (LAM)
106749 | Determine TB LAM Ag, 25 tests | 25 tests | 14 months | Determine TB LAM Ag is an in vitro, visually read, qualitative immunoassay for the detection of LAM antigen of Mycobacteria in human urine as an aid in the diagnosis of active mycobacterial infection in HIV positive individuals with clinical symptoms of tuberculosis. · This item includes: 25 Determine TB LAM Ag Tests; 1 Determine TB LAM Ag Reference Card and 1 Determine TB LAM Ag IFU | Abbott Rapid Dx International |  | 92.5 | USD | 0 | General cargo | 0.8269 | 28 | 16 | 1 | 0.0005 | 7D2741 | 38220000 | https://www.globalpointofcare.abbott/en/product-details/determine-tb-lam.html
1.3 SMEAR MICROSCOPY
1.3.1 Smear microscopy kits
106938 | Ziehl Neelsen light microscopy supplies Kit | 1 kit | 34 months | The kit is composed of the following ready to use reagents and consumables in quantities adequate for 1,000 smears. The kit is packed in 3 boxes (1 box with reagents, 1 box with consumables and 1 box with paper towels) Reagents: ZN carbol-fuchsin stain solution 1L: 6L ; ZN Acid-Alcohol decolorization solution 1L: 6L; ZN Methylene blue counterstaining solution 1L: 6L; Immersion oil: 4 x 25ml (100ml); Chlorine disinfectant: Pellets for 100 L (100 tablets). Consumables: Microscope slide: 20 x 50 units (1,000 units); Lens cleaning tissues: 100 tissues; Pencil: 10 pencils; Waterproof marker pen: 2 pens; Waste plastic bag: 500 bags; Wooden applicator stick: 1000 sticks; Nitrile Gloves sizes S, M, L: 800 gloves Paper towels: Paper towels: 20 x 150 units (3,000) | Svizera Europe B.V. |  | 323 | USD | 0 | Dangerous cargo ZN carbol-fuchsin stain solution: UN2821 6.1 III ZN Acid-Alcohol decolorization solution: UN2924 3(8) II Chlorine disinfectant pellets: UN3077 9 III Immersion oil: UN 3082, 9, III | 39 | 40 | 40 | 40 | 0.225 | Reagents: ZN carbol-fuchsin stain solution: Gainland International Ltd. S0170; ZN Acid-Alcohol decolorization solution: Gainland International Ltd. S1075.1; ZN Methylene blue counterstaining solution: Gainland International Ltd. S0830.1; Immersion oil: Gainland International Ltd. M250; Chlorine disinfectant pellets: Gainland International Ltd. T0025. Consumables: Microscope slides: Yancheng Feizhou Glass Co., Ltd. 7105-1; Lens cleaning tissues: Hangzhou Ocome Filtech Co.,LTD. HZO-L1015; Pencils Shenzhen: Shiechen Industial Company LTD. 1002; Waterproof marker pens: Wenzhou Junbo Stationery Co., LTD. 1125; Waste plastic bags: QINGDAO SEVENLEAVES PACKAGING CO., LTD. 91370285MA3C74CU47; Wooden applicator sticks: Shucheng County Ancheng Bamboo and Wood Handicraft Factory. BBS223; Nitrile Gloves sizes: Top Glove Sdn Bhd. ENW50. Paper towels: Megall Paper (Qingdao) Co., Ltd. SMD-PTR150 | 38221900 | http://www.svizera.org/files/pdfs/Productlijsten/productlist_diagnostic_svizera-europe_2022.pdf
106939 | Ziehl Neelsen Staining Kit, Hot Staining | 1 kit | 34 months | This kit is composed of the following ready to use reagents and with quantities adequate for 150 - 200 smears. 1. ZN carbol-fuchsin stain solution 1L: 1% Basic Fuchsin; 5% Phenol (C6H6O); 10% Ethanol (C2H60); 85% distilled Water. 2. ZN Acid-Alcohol decolorization solution 1L: 3.0% Hydrochloric acid (HCl); 97% Ethanol ( C2H60). 3. ZN Methylene blue counterstaining solution 1L: 0.1 to 0.3% Methylene blue in 100% water (volume %). | Svizera Europe B.V. |  | 35 | USD | 0 | Dangerous cargo ZN carbol-fuchsin stain solution: UN2821 6.1 III ZN Acid-Alcohol decolorization solution: UN2924 3(8) III | 3.5 | 35 | 15 | 30 | 0.0157 | ZN carbol-fuchsin stain solution: Gainland International Ltd. S0170 ZN Acid-Alcohol decolorization solution: Gainland International Ltd. S1075.1 ZN Methylene blue counterstaining solution: Gainland International Ltd. S0830.1 | 38221900 | http://www.svizera.org/files/pdfs/Productlijsten/productlist_diagnostic_svizera-europe_2022.pdf
106940 | Auramine microscopy supplies Kit | 1 kit | 34 months | This kit is composed of the following ready to use reagents and consumables in a quantity adequate for 1,000 smears. The kit is packed in 3 boxes (1 box with reagents, 1 box with consumables and 1 box with paper towels) Reagents: Carbolic Auramine stain solution 1L: 6 x 1L; Auramine Acid alcohol decolorisation solution 1L: 6 x 1L; Auramine Methylene blue counterstaining solution 1L: 6 x 1L; Immersion oil 25ml: 4 x 25ml (100ml); Chlorine disinfectant pellets: 100 Pellets for 100 L. Consumables: Microscope slides 50 units: 20 x 50 units (1,000 units); Lens cleaning tissues: 100 tissues; Pencils: 10 units; Waterproof marker pens: 2 units; Waste plastic bags: 500 bags; Wooden applicator sticks: 1,000 units; Nitrile Gloves sizes S, M, L - 800 gloves. Paper towels 150 units: 20 x 150 units (3,000) | Svizera Europe B.V. |  | 323 | USD | 0 | Dangerous cargo Auramine stain solution: UN2821, 6.1, III Auramine Acid alcohol decolorisation solution: UN2924 3(8) II Chlorine disinfectant pellets: UN3077 9 III Immersion oil: UN 3082, 9, III | 39 | 40 | 40 | 40 | 0.225 | Reagents: Carbolic Auramine stain solution: Gainland International Ltd. S0090; Auramine Acid alcohol decolorisation solution: Gainland International Ltd. S1075.2; Auramine Methylene blue counterstaining solution: Gainland International Ltd. S0830.2, Immersion oil: Gainland International Ltd. M250; Chlorine disinfectant pellets: Gainland International Ltd. T0025 Consumables: Microscope slides 20x50 units: Yancheng Feizhou Glass Co., Ltd. 7105-1; Lens cleaning tissues Hangzhou Ocome Filtech Co.,LTD. HZO-L1015; Pencils Shenzhen Shiechen Industial Company LTD. 1002; Waterproof marker pens: Wenzhou Junbo Stationery Co., LTD. 1125; Waste plastic bags QINGDAO SEVENLEAVES PACKAGING CO., LTD. 91370285MA3C74CU47; Wooden applicator sticks Shucheng County Ancheng Bamboo and Wood Handicraft Factory. BBS223; Nitrile Gloves sizes: Top Glove Sdn Bhd. ENW50. Paper towels: 150 units Megall Paper (Qingdao) Co., Ltd. SMD-PTR150 | 38221900 | http://www.svizera.org/files/pdfs/Productlijsten/productlist_diagnostic_svizera-europe_2022.pdf
106941 | Auramine staining kit | 1 kit | 34 months | This kit must be composed of the following ready to use reagents and with quantities adequate for 150 - 200 smears. 1. Carbolic Auramine stain solution 1L: Color index 41000, CAS Number: 2465-27-2, molecular formular: C17H22ClN3, [C17H21N3 HCl]. 2. Auramine Acid alcohol decolorisation solution 1L: 0.5% HCl and 99.5% ethanol (C2H60). 3. Auramine Methylene blue counterstaining solution 1L: 0.3% methylene blue and 100% (volume%) water. Color Index: 52015, CAS: 1934-16-3, molecular formula: C16H18ClN3S. | Svizera Europe B.V. |  | 35 | USD | 0 | Dangerous cargo UN 2821 & PG III | 3.5 | 35 | 15 | 30 | 0.0157 | Carbolic Auramine stain solution: Gainland International Ltd.S0090 Auramine Acid alcohol decolorisation solution: Gainland International Ltd. S1075.2 Auramine Methylene blue counterstaining solution:Gainland International Ltd. S0830.2 | 38221900 | http://www.svizera.org/files/pdfs/Productlijsten/productlist_diagnostic_svizera-europe_2022.pdf
106942 | Ziehl Neelsen / Auramine microscopy starter supplies Kit | 1 kit | 34 months | This kit is composed of the following ready to use consumables and with quantities adequate for the kit. The kit is packed in 2 boxes (1 box with consumables and 1 box with paper towels) Consumables: Glass alcohol lamp (2 units); Cotton wick for spirit lamp (2x 1meter); Mechanical Timer: 1 unit; Slide staining rack: 2 units; Slide drying rack: 2 units; Slide storage box: 6 units; Slide-holding forceps: 2 units; Reagent bottle 1L: 4 units; Staining bottle 250 ml: 3 units; Staining bottle 500 ml: 3 units; Beaker 250 mL: 1 unit; Beaker 1000 mL: 1 unit; Measuring cylinder 10 ml: 1 unit; Goggles: 2 units; Funnels sizes S, M, L: 3 units, 1 unit of each size; Folded paper filters for funnels 45mm, 120mm, 155mm: 100 units; Pencils: 10 units; Pencil sharpener: 1 unit; Stand for waste plastic bags: 3 units; Waste plastic bag: 500 units; Paper towel dispenser: 1 unit; Handbook Lab Diagnosis of TB by Sputum Microscopy: 1 unit; Standard Operating Procedures: 1 unit. Paper towels: 150 units: 20 x 50 units (3000) | Svizera Europe B.V. |  | 110 | USD | 0 | General cargo | 23 | 60 | 40 | 35 | 0.084 | Consumables: Glass alcohol lamp: Wuhan Pharmchem Co., Ltd 6201; Cotton wick for spirit lamp 1m Wuhan Pharmchem Co., Ltd. 6202; Mechanical timer: SUZHOU SINOMED. SMD-MT301; Slide staining rack Zhejiang Hanjun Metal Products Co., Ltd. YF-1027; Slide drying rack: Dongguan Jiesheng Hardware Indusry 330281600108451; Slide storage box: SUZHOU SINOMED SMD-STB100; Slide-holding forceps Atlas Surgical ASSI.2460; Reagent bottle: Nantong Hailun Bio-Medical Apparatuses Manufacturing Co., Ltd. HL-K46 (Amber COLOR); Staining bottle 250 ml: SUZHOU SINOMED SMD-SB250; Staining bottle 500 ml: SUZHOU SINOMED STM-SB500; Beaker 250 mL: NINGBO YINZHOU JOAN LAB EQUIPMENT CO., LTD. 1102-250; Beaker 1000 mL: NINGBO YINZHOU JOAN LAB EQUIPMENT CO., LTD 1102-1000; Measuring cylinder: NINGBO YINZHOU JOAN LAB EQUIPMENT CO., LTD. 1601-10; Goggles SUZHOU SINOMED STM-GOGAF; Funnels: SUZHOU SINOMED SMD-FUNS,SMD-FUNM,SMD-FUNL; Folded paper filters for funnels: Hangzhou Ocome Filtech Co.,LTD. HZO-F45, HZO-F110, HZO-F150; Pencils: Shenzhen Shiechen Industial Company LTD. 1002; Pencil sharpener: Shenzhen Shiechen Industial Company LTD. 301; Stand for waste bags: Zhejiang Hanjun Metal Products Co.,Ltd YF-0473; Waste plastic bag: QINGDAO SEVENLEAVES PACKAGING CO., LTD. 91370285MA3C74CU47; Paper towel dispenser: SUZHOU SINOMED SMD-PTD. Paper towels: Megall Paper (Qingdao) Co., Ltd. SMD-PTR150 | 90189084 | http://www.svizera.org/files/pdfs/Productlijsten/productlist_diagnostic_svizera-europe_2022.pdf
1.3.2 Smear microscopy equipment
106535 | Bright Field Microscope (CX23) | 1 unit | N/A | This microscope model is easy to use, and offers outstanding optical performance with a large 20mm field of view. Binocular 30⁰ inclined and rotatable 360⁰. Binocular microscope with 2 years warranty. Power supply: 100-240VAC, 50/60Hz. This item includes: - Binocular tube; stand with quadruple revolving nose-piece; - Eyepieces 10x with Helicoid adjustment (FN 20, anti fungus); right-handle- double plate stage; - Abbe condenser; power cord; Plan objectives (anti-fungus) 4x, 10x, 40x, 100x; - With AC adapter and fixing belt for transportation. Connectable to a digital camera. This item includes a dust cover and a Lens tissue paper The reflecting mirror has to be procured separately. | The Medical Export Group |  | 825.4 | USD | 0 | General Cargo | 6 | 40 | 50 | 50 | 0.02 | Olympus CX23 | 90189084 | https://www.olympus-lifescience.com/en/microscopes/upright/cx23/
106629 | Bright field Microscope, reflecting mirror (CX23) | 1 unit | N/A | Mirror unit for microscope Olympus CH20/CX21/CX22 and CX23; Spare part of the Binocular bright field microscope - ZN | The Medical Export Group |  | 22.86 | USD | 0 | General Cargo | 0.1 | 75 | 67 | 75 | 0.002 | Olympus CX23 reflecting mirror | 90189084 | https://www.olympus-lifescience.com/en/microscopes/upright/cx23/
106630 | Bright Field Microscope, transport case (CX23) | 1 unit | N/A | Wooden storage/transport box for Olympus microscopes models CH20/CX21/CX22 and CX/23. Supplied with carrying handle, size 26.4x30.5x47.4cm; If the microscope will stay in the lab, a transport case may not be necessary | The Medical Export Group |  | 161.9 | USD | 0 | General Cargo | 6.25 | 26.4 | 30.5 | 47.4 | 0.0383 | Olympus CX23 transport case | 90189084 | https://www.olympus-lifescience.com/en/microscopes/upright/cx23/
106943 | Bright field microscope (Primo Star 1) | 1 unit | N/A | This microscope model is easy to use, and offers outstanding optical performance with a large 20mm field of view. Binocular 30⁰ inclined and rotatable 360⁰. Binocular microscope with 2 years warranty. Power supply: 100-240VAC, 50/60Hz. This item includes: - Binocular tube; stand with quadruple revolving nose-piece; - Eyepieces 10x with Helicoid adjustment (FN 20, anti fungus); right-handle- double plate stage; - Abbe condenser; power cord; Plan objectives (anti-fungus) 4x, 10x, 40x, 100x; - With AC adapter and fixing belt for transportation. Connectable to a digital camera. The reflecting mirror has to be procured separately. | Svizera Europe B.V. |  | 897.48 | USD | 0 | General Cargo | 15 | 30 | 61 | 60 | 0.14 | Primo Star 1 | 90118000 | https://www.micro-shop.zeiss.com/en/dk/system/primostar+1-primostar+1-upright+microscopes/10336/
106944 | Bright field Microscope, reflecting mirror (Primo Star 1) | 1 unit | N/A | Mirror unit (spare part) for the bright field microscope | Svizera Europe B.V. |  | 20.5 | USD | 0 | General Cargo | 0.3 | 75 | 67 | 75 | 0.002 | Mirror unit for Carl Zeiss Primo Star (415500-1202-000) | 90119000 | https://www.micro-shop.zeiss.com/en/dk/system/primostar+1-primostar+1-upright+microscopes/10336/
106945 | Bright Field Microscope, transport case (Primo Star 1) | 1 unit | N/A | Wooden storage/transport box for Olympus microscopes models CH20/CX21/CX22 and CX/23. Supplied with carrying handle, size 26.4x30.5x47.4cm; If the microscope will stay in the lab, a transport case may not be necessary | Svizera Europe B.V. |  | 94.01 | USD | 0 | General Cargo | 7.5 | 43 | 28 | 50 | 0.29 | Case for transport and storage for primo star (415500-1827-000) | 90119000 | https://www.micro-shop.zeiss.com/en/dk/system/primostar+1-primostar+1-upright+microscopes/10336/
106950 | LED microscope (Iscope) | 1 unit | N/A | Fluorescence microscope with binocular tube 30°, equipped with quadruple revolving nosepiece (tilted backwards) and objectives “Planachromat” 10x, 20x, 40x and 100x (oil), with white LED illumination for transmitted-bright field and blue LED illumination for epi-fluorescence. Product has 2 years warranty. · The base is sturdy, stable, base body with focus adjustment. Base of epoxy powder coated heat treated, resistant to standard reagents used for staining; · Optical system: Color corrected infinity optics, parafocal distance 45 mm, tube length: 180±20 mm; · Eye pieces: Tube rotatable through 360° with adjustable inter-pupillary distance ranging from at least 48–75 mm, maintaining parafocality; · The nosepiece is with ribbed grip for easy rotation on a precision ball bearing mechanism; · The stage is provided with a spring loaded slide holder for safe and exact positioning of the slide. Connectable to a digital camera. This item includes: dust cover and antistatic optical cleaning brush | Euromex |  | 1920 | USD | 0 | General Cargo | 17 | box 1 60, box 2 70cm | box1 50cm, box2 29cm | box1 38cm, box2 44cm | 0.203 | IScope.1152-PLI/LB-UN1 | 90118000 | https://www.euromex.com/en/products/applications/86129-un1/
106951 | LED microscope, transport case (Iscope) | 1 unit | N/A | Wooden case for transport and storage of the LED microscope. Supplied with carrying handle with appropriate receptacles (with fixing points and belts to retain the microscope and to eliminate shocks during transportation) for safe holding of accessories. Lockable. Warranty of 2 years If the microscope will stay in the lab, a transport case may not be necessary | Euromex |  | 89 | USD | 0 | General Cargo | 8 | 75 | 60 | 60 | 0.27 | IS.4300-SP | 90118000 | https://www.euromex.com/en/products/applications/86129-un1/
106952 | LED microscope, battery supply unit (Iscope) | 1 unit | N/A | Battery supply unit consisting of battery and charger in one unit, for power supply via the line or the battery. This battery can be use for the LED microscope (Iscope) (106950) | Euromex |  | 69 | USD | 0 | Dangerous CAO Dangerous Cargo UN 3480, Cargo Aircraft Only (CAO) | 1 | 20 | 20 | 20 | 0.008 | IS.BATT | 90118000 | https://www.euromex.com/en/products/applications/86129-un1/
106979 | LED microscope (Primo Star 3 iLED) | 1 unit | N/A | Fluorescence microscope with viewing angle 25°/20, equipped with quadruple revolving nosepiece (tilted backwards) and objectives “Plan-achromatic” 10X, 40X and 100X , with white LED illumination for transmitted-bright field and blue LED illumination. -The base is sturdy, stable, base body with focus adjustment with protective plastic coverying. Base of epoxy powder coated heat treated, resistant to standard reagents used for staining; - Optical system : Color corrected infinity optics, parafocal distance 45mm, tube length :180±20mm, the entire optical system shall be ""anti-fungus"" treated; - Eye pieces : Viewing angle 25°/20, ergonomic with Tube rotatable through 360° with adjustable inter-pupillary distance ranging from at least 48-75mm, maintaining parafocality; oculars with visual field numbers of 10x/20 Br.foc. -The nosepiece : mechanical revolving nosepiece to accommodate 4 objectives, inclined backwards (tilted backwards - mechanical rackless stage 75 X 40), with ribbed grip for easy rotation on a pricision ball bearing mechanism; -Complete configuration with fixed-Köhler stand including Transmitted-light illumination LED 1W white-light, 5600K 4-position nosepiece -The stage : uniformely horizontal, mechanical stage with dimensions of 140 x140 mm (w x d) ±5 mm each, it is provided with a spring loaded slide holder for safe and exact positioning of the slide. Connectable to a digital camera. -Country-specific power cable -The 5v USB port offer the possibility to use power bank to supply power to Primostar 3 -Magnetic transmitter is stick on cover plater, no chance to fall down or misalign -The main power cable can be stored at back size panel -5 years warranty This item includes : dust cover and antistatic optical cleaning brush It works at humidity : up to 100%, ambient temperature in the range +5°C to +40° C | Svizera Europe B.V. |  | 1990 | USD | 0 | General Cargo | 15 | 30 | 61 | 60 | 0.14 | Primo Star 3 iLED (415501-0051-000) | 90118000 | https://www.zeiss.com/microscopy/en/products/light-microscopes/widefield-microscopes/primostar-3-iled.html
106534 | LED microscope, transport case (Primo Star iLED) | 1 unit | N/A | Temporarily not available | Svizera Europe B.V. |  | 220 | USD | 0 | General Cargo | 7.5 | 61.5 | 54 | 32.5 | 0.108 | Primo Star iLED transport case (415500-1827-000) | 90119000 | https://www.zeiss.com/microscopy/int/products/light-microscopes/primo-star-iled.html
106980 | Battery supply for microscope(PrimoStar 1 and PrimoStar 3 iLED) | 1 unit | N/A | Battery supply unit consisting of battery and charger in one unit, for power supply via the line or the battery. Power bank Lithium-Ion. 20100mAh 73Wh. This battery can be used for the Bright field microscope (Primo Star 1) and for the LED microscope (Primo Star 3 iLED) | Svizera Europe B.V. |  | 203 | USD | 0 | Dangerous CAO Dangerous Cargo UN 3480, Class 9A, Cargo Aircraft Only (CAO) | 5.61 | 27 | 22 | 14 | 0.0199 | PrimoStar 1 and PrimoStar 3 iLED (000000-0677-975) | 90119000 | https://www.xtpower.de/XT-20000QC3/
1.3.3 Smear microscopy reagents
106276 | Auramine O, 25 g | 1 unit | 16 months | Auramine O for microscopy. CAS No. 2465-27-2; · Dye Content > 80 % | TTM |  | 38.4 | USD | 0 | Dangerous cargo UN 2811, Class 6.1, PG III | 0.045 | 4 | 4 | 9 | 0.000144 | Merck KGaA SAFA861030-25G | 32041300 | https://www.sigmaaldrich.com/catalog/product/aldrich/861030?lang=en®ion=GB
106597 | Denatured alcohol, 1 lt | 1 unit | 23 months | Denatured alcohol for stain solutions and decolorization. CAS No. 64-17-5; · Grade: Purum/analytical ≥99,8 %, p.a., denatured | TTM |  | 8.35 | USD | 0 | Dangerous cargo UN 1170, Class 3, PG II | 5 | 21 | 28 | 32 | 0.018816 | Carl Roth GmbH 0911.1 | 22072000 | https://www.carlroth.com/com/en/a-to-z/ethanol/p/0911.8
106599 | Phenol crystals, 2x250 g | 1 unit | 16 months | Colourless Phenol crystals. CAS No. 108-95-2; · Phenol GR for analysis ACS,Reag. Ph Eur; Grade: Purum/analytical ≥ 99.5%, p.a This item includes: 2 x 250 g Subject to availability from our supplier. Product lead-time may be longer than usual. Please contact your GDF Country Supply Officer for further information | TTM |  | 96.05 | USD | 0 | Dangerous cargo UN 1671, Class 6.1, PG II | 4 | 21 | 28 | 32 | 0.018816 | Merck KGaA 1.00206.0250 | 29071100 | https://www.sigmaaldrich.com/catalog/product/mm/100206?lang=en®ion=GB
106627 | Immersion oil, 100 ml | 4 x 25ml (100ml) | 34 months | Microscope Immersion Oil 100 ml. Immersion oil with refractive index in the range (nD) 1,515 – 1,517. Atoxic and Free from particle with drop size of about 0.1 mL. No vapor or odor and labelled with applying warning symbols. This item includes: 4 bottles of 25 ml each for a total of 100 ml | Svizera Europe B.V. |  | 14.25 | USD | 0 | General cargo UN 3082, 9, III (Excepted Quantity - Not restricted) | 0.15 | 10 | 12 | 10 | 0.0012 | Gainland International Ltd. (M250) | 38221900 | http://www.gccdiagnostics.com/user/M250%20Immersion%20oil%202008%20modification.pdf
1.3.4 Smear microscopy consumables
106525 | Sputum containers with screw caps | 1000 units | N/A | Sample container used to collect human sputum and urine for diagnosis of tuberculosis, made of transparent plastic polypropylene with multi (more than 2) thread closing screw caps. · Volume 80 ml with height: 50 to 70 mm. Opening angle of 270°. · Mouth diameter: 40 - 55 mm, wide opening to easily collect sputum. · Leak-resistant (waterproof) with frosted writing panel which accept pencil. · Secure stackable cups for volume-reducing packing, sample transport and storage For logistics purposes: 176 boxes (11 pallets) will fill up a 20' container | Svizera Europe B.V. |  | 83.5 | USD | 0 | General cargo | 16.95 | 60 | 40 | 50 | 0.12 | Jiangsu Benoy Lab Instrument Co.,Ltd (SMD-SC80) | 39269097 / 90189084 | https://www.sz-sinomedevice.com/sputum-containers-with-screw-caps.html
106609 | Microscope slides, 50 x 20 slides | 20x50 units (1000 units) | N/A | Microscope slides, made of soda-lime glass with frosted ends for marking. · Dimensions: 76 mm × 26 mm and Thickness: 1.0–1.2 mm ; Straight edges and corners with cellophane wrapped. Compliant with ISO 8037–1. This item includes: 50 boxes of 20 units each for a total of 1000 units | Svizera Europe B.V. |  | 16.25 | USD | 0 | General cargo | 6 | 38 | 18 | 8 | 0.0055 | Yancheng Feizhou Glass Co., Ltd. (7105-1) | 70179000 | www.ycfzbl.com
106610 | AFB Quality control slides kit | 1 kit of 6 slides (3 positive and 3 negative control slides) | 22 months | Quality control slides to validate routine staining procedures. Individually packed to preserve control sample integrity. Eliminate need to maintain control cultures or retain positive results. This item includes: · 3 positive slides The positive well (+) contains Mycobacterium scrofulaceum (acid fast bacilli) to produce the positive acid-fast stain reaction. · 3 negative slides: The negative well (-) contains Escherichia coli (a gram+ bacterium), which results in the negative acid-fast stain reaction The nonviable microorganisms stain control log is available electronically: https://www.alphatecsystems.com/wp/wp-content/uploads/2018/09/L003241.D-AFB-Quality-Control-Log.pdf | Svizera Europe B.V. |  | 19.79 | USD | 0 | General cargo | 0.65 | 0.09 | 8.9 | 6.4 | 0.0055 | Alpha-Tec Systems, Inc. (0003246) | 382219 | https://www.alphatecsystems.com/mycobacteriology/qc1-afb-slides/
106568 | Slide storage box | 1 unit | N/A | Slide storage box made of durable plastic with a capacity to hold 100 slides (slides 76 mm ~ 26 mm). · Cork-lined base and with a cover with an index-card holder | Svizera Europe B.V. |  | 1.8 | USD | 0 | General cargo | 0.23 | 48 | 41 | 36 | 0.070848 | SUZHOU SINOMED (SMD-STB100) | 39231010 | https://www.sz-sinomedevice.com/slide-storage-box.html
106564 | Glass alcohol lamp, 200 ml | 1 unit | N/A | Glass alcohol lamp (200 ml) with screw top lid with fuel of 96% denatured alcohol. The cap is made of metal to extinguish the fire | Svizera Europe B.V. |  | 1.45 | USD | 0 | General cargo | 0.24 | 61.5 | 41.5 | 38 | 0.0015 | Wuhan Pharmchem Co., Ltd (6201) | 70179000 | www.cspharmchem.com
106565 | Cotton wick for spirit lamp, 1 meter | 10x10 cm (1 meter) | N/A | Cotton wick for spirit lamp of 1 meter made of cotton. The cotton wick are pre-cut cylindrical wicks in 10 x 10 cm. | Svizera Europe B.V. |  | 0.76 | USD | 0 | General cargo | 0.68 | 50 | 43 | 41 | 0.09 | Wuhan Pharmchem Co., Ltd (6202) | 36031000 | www.cspharmchem.com
106566 | Slide staining rack | 1 unit | N/A | Slide staining rack made of stainless steel 410 which has excellent chemical resistance to collect excess stain. Harbours 30 to 40 of standard slides size (26 x 76 mm) | Svizera Europe B.V. |  | 4.38 | USD | 0 | General cargo | 1.43 | 45.5 | 31.5 | 36 | 0.052 | Zhejiang Hanjun Metal Products Co.,Ltd (YF-1027) | 73239900 | Manufacturer website it is not available
106567 | Slide drying rack | 1 unit | N/A | Slide drying rack made of stainless steel 410 which has excellent chemical resistance. Harbours 30 to 40 of standard slides size (26 x 76 mm) | Svizera Europe B.V. |  | 1.7 | USD | 0 | General cargo | 2 | 120 | 80 | 100 | 0.96 | Dongguan Jiesheng Hardware Industry (330281600108451) | 73261990 | Manufacturer website it is not available
106569 | SIide-holding forceps | 1 unit | N/A | Slide-holding forceps made of stainless steel 410 which has excellent chemical resistance. Able to hold 1 standard slide size (26 x 76 mm). Size: Between 100 to 120 mm length | Svizera Europe B.V. |  | 0.9 | USD | 0 | General cargo | 0.018 | 33 | 24 | 13 | 0.0103 | Atlas Surgical (ASSI.2460) | 90189084 | www.sahaj.com, www.atlassurgicalindia.com
106570 | Mechanical timer | 1 unit | N/A | Mechanical timer spring-powered (Not line or battery powered). Timer range minimum 20, maximum 60 min with 1 minute or shorter increments. Distance of one minute increments >3 mm | Svizera Europe B.V. |  | 1.56 | USD | 0 | General cargo | 0.08 | 38.5 | 28.5 | 32.5 | 0.035661 | SUZHOU SINOMED (SMD-MT301) | 91070000 | https://www.sz-sinomedevice.com/mechanical-timer.html ‎
106576 | Measuring cylinder, 10 ml | 1 unit | N/A | Measuring cylinder made of type 1 borosilicate glass with spout. Autoclavable (121°C) and Acid and alkalin resistant. Wall thickness of 3 mm; 0.2 mL graduation printed on glass. Volume 10 ml | Svizera Europe B.V. |  | 1.1 | USD | 0 | General cargo | 0.58 | 44 | 36 | 48 | 0.076 | NINGBO YINZHOU JOAN LAB EQUIPMENT CO.,LTD (1601-10) | 70172000 | https://www.monotaro.id/s028109709.html
106947 | Folded paper filters for funnels 45 to 155mm | 100 units | N/A | Folded paper filters to be used with funnels of sizes between 45 mm to 155 mm. Filter paper for qualitative analysis made of >95% cellulose with ash content of 0.06 - 0.15%. Thickness: 0.16 mm - 0.33 mm | Svizera Europe B.V. |  | 5.03 | USD | 0 | General cargo | 4.5 | 67 | 18 | 16 | 0.019 | Hangzhou Ocome Filtech Co.,LTD. (HZO-F45, HZO-F110, HZO-F150) | 48232000 | http://www.hzocome.com/eproduct_1.asp?id=472
106655 | Wooden applicator sticks | 1000 units | N/A | Wooden applicator sticks with no swab end made of bamboo. Cylindrical with 3.0 x 120mm size. Length 120mm | Svizera Europe B.V. |  | 1.92 | USD | 0 | General cargo | 0.85 | 45 | 16 | 23 | 0.012937 | Shucheng County Ancheng Bamboo and Wood Handicraft Factory (BBS223) | 44219100 | https://www.anchengfoodservice.com/eco-friendly-natural-bamboo-BBQ-skewer.html
106946 | Lens cleaning tissues | 100 tissues | N/A | Lens cleaning tissues with lint free. Not scratching glass. Size per piece: 100 mm x 150 mm. Box of 100 tissues | Svizera Europe B.V. |  | 0.77 | USD | 0 | General cargo | 0.025 | 52 | 47 | 17 | 0.042 | Hangzhou Ocome Filtech Co.,LTD. (HZO-L1015) | 48239085 | Not available
106571 | Reagent bottle 1 lt | 1 unit | N/A | Reagent bottle made of amber borosilicate to prevent UV light penetrating inside the bottle. Wall thickness of 4mm. Provided with a GL45 screw cap including tight fitting screw cap and pouring ring made of polypropylene. Autoclavable (121 °C) | Svizera Europe B.V. |  | 4.72 | USD | 0 | General cargo | 0.68 | 62 | 42 | 24.3 | 0.0633 | Nantong Hailun Bio-Medical Apparatuses Manufacturing Co., Ltd. HL-K46 (Amber COLOR) | 70179000 | http://www.nthailun.net/productView_en.asp?ID=378&SortID=160
106605 | Graduated pipette with suction piston | 1 unit | N/A | Graduated pipette with piston, capacity: 50 ml, brown graduation: 0,2 ml, with attached syringe for simple and risk-free pipetting, main point ring graduation, zero point at top, graduated to tip. | TTM |  | 37.85 | USD | 0 | General Cargo | 0.02 | 3 | 3 | 50 | 0.00045 | Hirschmann Laborgeräte GmbH 1190175 | 70179000 | https://www.hirschmannlab.com/artikelgruppe/11901?parent=%7B60C3B706-491A-4B9D-9DAD-87313315A49B%7D
106606 | Weighting boats | 12 units | N/A | Weighing boats to weigh different elements in the laboratory; · Square shape, glazed porcelain; · Good chemical resistance and acid resistance; Withstand temperatures up to at least 70°C; · White color, made of Porcelain, glazed; · Flat base with round corners | TTM |  | 51.3 | USD | 0 | General Cargo | 1.3 | 6 | 3 | 15 | 0.00027 | Witeg Labortechnik GmbH 8526371 | 69091100 | https://www.witeg.de/en/detail/index/sArticle/10433
1.4 MOLECULR TESTING: CENTRALIZED
1.4.1 LPA equipment
106050 | Thermoshaker / Twincubator (7025009) | 1 unit | N/A | The Thermoshaker / TwinCubator is for use with reverse line blot strip assays and for manual hybridization and washing steps of up to 12 GenoType test strips. · High quality peltier elements. Minimal risk of contaminations. Pre-installed temperature profile for processing of the GenoType tests. · Number of samples: 1-12. Program storage: 9 programs with up to 20 steps. Temperature range 4-105°C in steps of 1°C. Heating rate: 0.3°C/s; Cooling rate: 0.2°C/s. Power supply: 100-240 V, 50/60 Hz · Countries where US/EU/OSCE/UN sanctions apply are excluded. This item includes: TwinCubator with hybridization block; Power cable; Operator’s manual; Checklist for installation; Decontamination protocol and QC records. Remote support is delivered by HAIN Lifescience / Bruker Support Team (phone number: +49 74 73- 94 51- 744 . E-mail: Diagnostics.support@bruker.com), available from Monday to Friday (8:30am-5:00pm CET). | Hain LifeScience GMBH |  | 2200 | EURO | 0 | General Cargo | 13.2 | 54 | 37 | 28 | 0.055944 | 7025009 | 90181990 | https://www.hain-lifescience.de/en/products/equipment/hybridization/twincubator.html
106064 | GT blot - 48 (1003/1) | 1 unit | N/A | The GT-Blot 48 is for use with reverse line blot strip assays and provides fully automated washing and hybridization of up to 48 GenoType test strips. · Novel heating system and high quality system components. Color-coded reagent containers. Automated recirculation of the remaining fluid. · Number of samples: 1-48. Pipetting volume: 1-4 ml. Dispensing channels: 6. Power supply: 220-240 V Europe, 110-120 V USA, 50/60 Hz. For customers using the GT-Blot48 only: approx. 1 Reagent Set for DNA STRIP Kits is suitable for 4 GenoType Kits of 96 tests. · Countries where US/EU/OSCE/UN sanctions apply are excluded. This item includes: GT-Blot 48; Power cable and spare fuses; USB cable with driver CD; GT-Blot 48 Programmable Software CD; 5 x Plastic trays, 2x48 wells; Square container for cleaning tubings; 1 x 500 ml container; 1 x 250 ml container; 4 x 125 ml containers; Pipetting template; Calibration documents; Operator‘s manual; Short instructions; Checklist for installation and Decontamination protocol. Remote support is delivered by HAIN Lifescience / Bruker Support Team (phone number: +49 74 73- 94 51- 744 . E-mail: Diagnostics.support@bruker.com), available from Monday to Friday (8:30am-5:00pm CET). | Hain LifeScience GMBH |  | 14352 | EURO | 0 | General Cargo | 85 | 100 | 84 | 89 | 0.7476 | 1003/1 | 90181990 | https://www.hain-lifescience.de/en/products/equipment/hybridization/gt-blot-48.html
106438 | GT-Blot 48 Tray for 96 strips (black) (1003/5) | 1 unit | N/A | GT-Blot 48 Tray for 96 strips, black (1003/5) for use in combination with the GT-Blot48 (106064). Remote support is delivered by HAIN Lifescience / Bruker Support Team (phone number: +49 74 73- 94 51- 744 . E-mail: Diagnostics.support@bruker.com), available from Monday to Friday (8:30am-5:00pm CET). | Hain LifeScience GMBH |  | 3 | EURO | 0 | General Cargo | 0.1 | 45 | 16 | 2 | 0.00144 | 1003/5 | 39239000 | https://www.hain-lifescience.de/en/products/equipment/hybridization/gt-blot-48.html
106613 | GTQ-Cycler 96 (7024007) | 1 unit | N/A | High-capacity thermal cycler assisting you processing with the GenoType systems. · Amplification of up to 96 samples. Pre-installed PCR protocols for processing of GenoType tests. Clearly arranged color touchscreen. Detailed GLP report. · Number of samples: 96. Sample volume: 5-100 μl. Program storage: 1,000 programs. Heated lid 112-120°C. Peltier elements: 8. Temperature sensors: 4. Temperature range: 4-99.9°C in steps of 0.1°C. Heating rate up to 5°C/s. Cooling rate up to 3°C/s. Power supply: 100-240 V, 50/60 Hz. · Countries where US/EU/OSCE/UN sanctions apply are excluded. This item includes: GTQ-Cycler 96 including thermal block; Touchscreen pen; Power cable; Operator’s manual; Short instructions; Checklist for installation; Decontamination protocol and QC records. Remote support is delivered by HAIN Lifescience / Bruker Support Team (phone number: +49 74 73- 94 51- 744 . E-mail: Diagnostics.support@bruker.com), available from Monday to Friday (8:30am-5:00pm CET). | Hain LifeScience GMBH |  | 4595 | EURO | 0 | General Cargo | 16.3 | 54 | 43 | 48 | 0.11146 | 7024007 | 84798997 | https://www.hain-lifescience.de/en/products/equipment/pcr/gtq-cycler-96.html
106732 | Thermoblock with 4 blocks | 1 unit | N/A | Thermoblock with a temperature range: ambient to +5°C to 150°C. Heat up speed of 5°C per per min. · User‘s self-compensation function to offset between real temperature and indicated temperature (±10.0°C); Internal PPS bath ensures best insulation, block made of anodized aluminium; · Precise temperature control; Digital timer function – delay time and operation control; timer range of 99hr 59min alarm; · Convenient integrated timer for time-sensitive incubations; Block removal tool with screw thread included; Locking mode; overheat and over-current protection and sensor error detection; · Storage function for the set values of temperature and timer; Various block types (optional); · Safety mechanism: locking mode, overheat and over-current protection and sensor error detection. Thi item includes: 1 unit + 4 blocks: Block of 96 places 0,2 ml PCR tubes, Block 15 places 15 ml conical tubes, Block 8 places 50 ml conical tubes and Block 96 places 0,5 ml PCR tubes | TTM |  | 1489.45 | USD | 0 | General Cargo | 0.8 | 25 | 26.5 | 8.3 | 0.0054988 | Witeg Labortechnik GmbH DH WHB02100 with blocks: DH.WHB510596; DH.WHB531515; DH.WHB535008; DH.WHB510570 | 84199085 | https://www.witeg.de/en/products/laboratory-equipment/heating/dry-bath-incubator/dry-bath-incubator-h-20-1-block-up-to-1500c
106205 | UV transilluminator 365 nm | 1 unit | N/A | UV-Transilluminator 365 nm with adjustable protective shield that can be fixed in any angle and thus protects the user against UV radiation. Stainless steel frame; Intensity in the range 7-8 mW/cm2. • UV-Transilluminator for reading gel - Ethidium bromide; • Unlimited lifetime of the filter for 312 nm; No flickering of tubes; Low amount of infrared light • Intensity selector (70%/100%); Ondulex reflector; Continuously adjustable UV protective shield • Adjustable power of the light in at least 2 positions: low and high power with UV filter; • Filter size at least of 210x 260 mm; No flickering effects; Low heat sample irradiation; • Transparent UV blocking cover made of acrylic; UV screen frames made of Stainless Steel AISI 304 or better; Body made of steel covered with coated paint or better material. | TTM |  | 1400.55 | USD | 0 | General Cargo | 15 | 30 | 38 | 62 | 0.07068 | LTF Labortechnik Vilber ECX-F 20.L | 90182000 | https://www.labortechnik.com/en/uv-transilluminators-low-cost/ecx-compact
106048 | Microcentrifuge | 1 unit | N/A | The microcentrifuge has rotors that are designed to spin microliter and PCR tubes. - RPM: 500 - 15,000 min -1; Adjustable in increments of 10; Max. RCF: 21,382; Max. capacity: 30 x 2.0 ml - Choice of 3 rotors; Impulse key for short cycle mode - Easy operation with keypad and control knob; - 4 program memories for more individuality; - 9 individual acceleration and deceleration stages; - model 200 R coolable from -20 to +40 °C with pre-cooling function. If an UPS is needed for this microcentrifuge, please select the UPS 700 VA. | TTM |  | 2034.75 | USD | 0 | General Cargo | 18 | 49 | 41 | 44 | 0.088396 | Andreas Hettich GmbH 2400 & 2434 & 2031 | 84211920 | https://www.hettichlab.com/downloadcenter/Products/Datasheets/MIKRO200-200R_EN.pdf
106224 | Refrigerated benchtop centrifuge | 1 unit | N/A | Refrigerated centrifuge for daily use in research or analytical laboratories. It has a capacity of 4 x 290 ml at a maximum RCF of 4.863; High-performance cooling in the range -20 °C to +40 °C, with a rapid cool program; High RCF: up to 16.504 for 50 ml tubes and 94 ml tubes /up to 24,400 for reaction tubes. · Protection against aerosols through the TÜV-tested accessories with bioseals; Compact benchtop centrifuge; Ergonomic C control panel allows complex centrifugation programs to be entered, stored and called up as required; 90 programmable memories; · 9 braking and acceleration ramps, unbraked run- down, or a cut-off speed can be selected. Ramp 9 represents the shortest braking and acceleration time; High capacity: max. 4 x 290 ml/ max. 36 x 15 ml · The accessories come separately. Therefore, it needs to be ordered with: Swing-out rotor 4 - place; Round bucket 4 - place; Lids for rotor 4 - place; Adapter for 50ml tubes 4- place; Adapter for 12ml/15ml tubes 4-place for refrigerated centrifuge. If an UPS is needed, please select the UPS 3000 VA. | TTM |  | 6060.3 | USD | 0 | General Cargo | 107 | 92 | 62 | 63 | 0.3594 | Andreas Hettich GmbH Hettich Rotina 380R | 84211920 | https://www.hettichlab.com/en/product/rotina-380-380-r/
106225 | Swing-out rotor 4 - place for refrigerated benchtop centrifuge | 1 unit | N/A | Swing-out rotor 4-place to fit in Benchtop centrifuge without carriers. · Angle: 90°; RPM max.: 5.000 min -1; RCF max.: 4.863 | TTM |  | 948.7 | USD | 0 | General Cargo | 3 | 26 | 21 | 11 | 0.006 | Andreas Hettich GmbH 1754 | 84219100 | https://profilab24.com/en/laboratory/centrifuge-with-vortex-function/hettich-swing-out-rotor-4-place-1754-for-benchtop-centrifuge
106226 | Round bucket 4 - place for refrigerated benchtop centrifuge | 1 set | N/A | Circular rotor buckets, made of aluminium. Able to accept up to 4 tubes of 50 ml with a proper adapter. This item includes: 1 set with 4 carriers. | TTM |  | 579.25 | USD | 0 | General Cargo | 2 | 29 | 21 | 10 | 0.0061 | Andreas Hettich GmbH 1752 | 84219100 | https://www.hettichlab.com/package/rotina-380-r-conical-package-1/
106227 | Lids for rotor 4 - place for refrigerated benchtop centrifuge | 1 set | N/A | Tight lid for rotor bucket made of high-performance plastic, hermetically sealed. Safety mechanisms to prevent opening during the centrifugation. This item includes: 1 set with 4 lids. | TTM |  | 174.2 | USD | 0 | General Cargo | 0.1 | 22 | 22 | 6 | 0.0029 | Andreas Hettich GmbH 1751 | 84219100 | https://www.hettichlab.com/en/package/rotina-380-blood-tube-package-1/
106228 | Adapter for 50 ml tubes 4- place for refrigerated centrifuge | 1 set | N/A | Adapter for falcon type tubes of 50 ml, able to work in refrigerated centrifuge. Made of high-performance plastic. This item includes: 1 set with 4 adapters of 4 places. | TTM |  | 327.15 | USD | 0 | General Cargo | 0.4 | 20 | 17 | 9 | 0.0031 | Andreas Hettich GmbH 1772-A | 84219100 | https://www.hettichlab.com/downloadcenter/Products/Datasheets/ZENTRIMIX380R_EN.pdf
106229 | Adapter for 12ml/15ml tubes 4-place for refrigerated centrifuge | 1 set | N/A | Adapter for tubes of 15 ml and/or 12 ml, able to work in refrigerated centrifuge. Made of high-performance plastic. This item includes: 1 set with 4 adapters of 9 places. | TTM |  | 327.15 | USD | 0 | General Cargo | 0.667 | 10 | 10 | 5 | 0.0005 | Andreas Hettich GmbH 1771-A | 84219100 | https://www.hettichlab.com/downloadcenter/Products/Datasheets/ZENTRIMIX380R_EN.pdf
106219 | Minishaker / Vortex | 1 unit | N/A | Minishaker / Vortex with strong and silent motor. Touch activation or continuous modes. · Touch activation or continuous modes; Speed control from 20 rpm to 3,000 rpm · Suction cups at base to minimize movement; Supplied with tube adapter and flask accessory · Capability to shake tubes with diameters up to at least 30 mm; Orbital diameter: 4,9 mm; · Power supply: 100-220V, 50/60 HZ This item contains: the adaptors for mixing cup, for platform head, for micro plate kit and for multi sample plate kit | TTM |  | 217.2 | USD | 0 | General Cargo | 4.7 | 27 | 20 | 25 | 0.0135 | LMS Consult GmbH KMC-1300V, V130001, V130003 and V130004 | 84798200 | http://02f5425.netsolhost.com/ussin/products/vs1300v.html
1.4.2 LPA reagents
106512 | GenoType MTBDR plus V2.0, 96 test kit 1876641 (30496A+51610) | 1 kit | 13 months | Molecular genetic assay for identification of resistance to rifampicin and/or isoniazid of M. tuberculosis complex from pulmonary specimens and/or liquid/solid culture samples. · Countries where US/EU/OSCE/UN sanctions apply are excluded. · The total transportation time from supplier warehouse to the end customer cannot exceed 19 days for the cooled transportation (2-8°C) of this item. · The Kit Component No. 2 of this product, must be stored at -20°C immediately upon arrival in the lab! The other components (under the inlay) have to be stored at +2-+8°C. It is important that upon receipt the kits components will be stored at the appropriate temperature immediately. It is the standard procedure for shipping and doesn´t have any effect on the quality or the stability of the products. This item includes: The GenoType MTBDRplus Version 2.0 (30496A), GenoLyse (51610) and the Master mix. For an order of 30+ kits of this product, excluding Genolyse, Hain Lifescience will take over the transportation cost to the respective country (CIP Incoterm). This does not include customs charges, local taxes and shipment to the end user site within the country. | Hain LifeScience GMBH |  | 750 | EURO | 0 | Cool cargo (+2-+8°C) | 1.5 | 24 | 15 | 14 | 0.00504 | 1876641 ((30496A+51610) | 38221900 | https://www.hain-lifescience.de/en/products/microbiology/mycobacteria/tuberculosis/genotype-mtbdrplus.html
106513 | GenoType MTBDRsl V2.0, 96 tests (31796A) | 1 kit | 13 months | Molecular genetic assay for identification resistance to fluoroquinolones, aminoglycosides /cyclic peptides of of M. tuberculosis complex from decontaminated sputum specimens and/or liquid/solid culture samples. · It is recommended to use MTBDRplus extracts; if not, Genolyse must be ordered separated. · Countries where US/EU/OSCE/UN sanctions apply are excluded. · The total transportation time from supplier warehouse to the end customer cannot exceed 19 days for the cooled transportation (2-8°C) of this item. · The Kit Component No. 2 of this product, must be stored at -20°C immediately upon arrival in the lab! The other components (under the inlay) have to be stored at +2-+8°C. It is important that upon receipt the kits components will be stored at the appropriate temperature immediately. It is the standard procedure for shipping and doesn´t have any effect on the quality or the stability of the products. This item includes: Master mix. For an order of 30+ kits of this product, excluding Genolyse, Hain Lifescience will take over the transportation cost to the respective country (CIP Incoterm). This does not include customs charges, local taxes and shipment to the end user site within the country. | Hain LifeScience GMBH |  | 750 | EURO | 0 | Cool cargo (+2-+8°C) | 1.5 | 24 | 15 | 14 | 0.00504 | 31796A | 38221900 | https://www.hain-lifescience.de/en/products/microbiology/mycobacteria/tuberculosis/genotype-mtbdrsl.html
106590 | GenoType Mycobacterium CM V2.0, 96 tests (29996A) | 96 tests | 13 months | Molecular genetic assay for identification and differentiation of M. tuberculosis complex and 27 clinically relevant Nontuberculous Mycobacterial species (NTM) from liquid or solid culture samples. · Validated with GenoLyse only. Genolyse must be ordered separately. · Countries where US/EU/OSCE/UN sanctions apply are excluded. · The total transportation time from supplier warehouse to the end customer cannot exceed 19 days for the cooled transportation (2-8°C) of this item. · The Kit Component No. 2 of this product, must be stored at -20°C immediately upon arrival in the lab! The other components (under the inlay) have to be stored at +2-+8°C. It is important that upon receipt the kits components will be stored at the appropriate temperature immediately.It is the standard procedure for shipping and doesn´t have any effect on the quality or the stability of the products. This item includes: Master mix. For an order of 30+ kits of this product, excluding Genolyse, Hain Lifescience will take over the transportation cost to the respective country (CIP Incoterm). This does not include customs charges, local taxes and shipment to the end user site within the country. | Hain LifeScience GMBH |  | 750 | EURO | 0 | Cool cargo (+2-+8°C) | 1.5 | 24 | 15 | 14 | 0.00504 | 29996A | 38221900 | https://www.hain-lifescience.de/en/products/microbiology/mycobacteria/ntm/genotype-mycobacterium-cm.html
106888 | GenoType Mycobacterium AS V1.0, 12 tests (298/H298) | 12 tests | 13 months | Molecular genetic Assay for identification of Nontuberculous Mycobacteria from Cultured material. · Validated with GenoLyse only. Genolyse must be ordered separately. · Countries where US/EU/OSCE/UN sanctions apply are excluded. · The total transportation time from supplier warehouse to the end customer cannot exceed 19 days for the cooled transportation (2-8°C) of this item. · The Kit Component No. 2 of this product, must be stored at -20°C immediately upon arrival in the lab! The other components (under the inlay) have to be stored at +2-+8°C. It is important that upon receipt the kits components will be stored at the appropriate temperature immediately.It is the standard procedure for shipping and doesn´t have any effect on the quality or the stability of the products. This item includes: Master mix. For orders with requests of more than 8 kits the price is 94 Euro | Hain LifeScience GMBH |  | 125 | EURO | 0 | Cool cargo (+2-+8°C) | 1.5 | 28 | 15 | 13 | 0.00546 | (298/H298) | 38221900 | https://www.hain-lifescience.de/en/products/microbiology/mycobacteria/ntm/genotype-mycobacterium-as.html
106889 | GenoType MTBC V1.X, 12 tests (301/H301) | 12 tests | 13 months | Molecular genetic assay for identification and differentiation of the M. tuberculosis complex from cultured material (liquid and/or solid culture samples). The GenoType MTBC guarantees reliable identification of the Bacillus Calmette-Guérin (BCG) strain. The BCG vaccine strain is a derivative of M. bovis attenuated in its pathogenicity and is used for immunization against TB and for immunotherapy of malignant tumours, such as bladder cancer. · Countries where US/EU/OSCE/UN sanctions apply are excluded. · The total transportation time from supplier warehouse to the end customer cannot exceed 19 days for the cooled transportation (2-8°C) of this item. · The Kit Component No. 2 of this product, must be stored at -20°C immediately upon arrival in the lab! The other components (under the inlay) have to be stored at +2-+8°C. It is important that upon receipt the kits components will be stored at the appropriate temperature immediately.It is the standard procedure for shipping and doesn´t have any effect on the quality or the stability of the products. This item includes: Master mix. For orders with requests of more than 8 kits the price is 94 Euro | Hain LifeScience GMBH |  | 125 | EURO | 0 | Cool cargo (+2-+8°C) | 0.5 | 16 | 13 | 11 | 0.002288 | (301/H301) | 38221900 | https://www.hain-lifescience.de/en/products/microbiology/mycobacteria/tuberculosis/genotype-mtbc.html
106477 | GenoLyse V 1.0, 96 tests (51610) | 96 tests | 18 months | Kit for isolation of genomic bacterial DNA from patient specimens (sputum, bronchial material, tracheal secretion, gastric juice, urine). Allows you to extract highly purified genomic bacterial DNA in a few steps. If the lab performs only Genotype MTBDRsl, need to order Genolyse separately. If Genotype MTBDRsl is performed after GenoType MTBDRplus, there is no need to order Genolyse. · Countries where US/EU/OSCE/UN sanctions apply are excluded. | Hain LifeScience GMBH |  | 52 | EURO | 0 | Cool cargo (+2-+8°C) | 0.5 | 14 | 6 | 7 | 0.000588 | 51610 | 38221900 | https://www.hain-lifescience.de/en/products/dna-isolation/genolyse.html
106890 | GenoType NTM-DR V1.0, 12 tests (29712) | 12 tests | 13 months | Molecular genetic assay for detection of resistance to macrolides and aminoglycosides in various Nontuberculous Mycobacterial species (NTM) from cultured material (bacteria grown on solid medium or in liquid medium). Allows the detection of several clinically relevant NTMs including the differentiation between M. intracellulare and M. chimaera. Mycobacteria resistance to macrolides and aminoglycosides are detected within the same step. Provides information as a prerequisite for an adequate therapy regimen. · Countries where US/EU/OSCE/UN sanctions apply are excluded. · The total transportation time from supplier warehouse to the end customer cannot exceed 19 days for the cooled transportation (2-8°C) of this item. · The Kit Component No. 2 of this product, must be stored at -20°C immediately upon arrival in the lab! The other components (under the inlay) have to be stored at +2-+8°C. It is important that upon receipt the kits components will be stored at the appropriate temperature immediately.It is the standard procedure for shipping and doesn´t have any effect on the quality or the stability of the products. This item includes: Master mix. For orders with requests of more than 8 kits the price is 94 Euro | Hain LifeScience GMBH |  | 125 | EURO | 0 | Cool cargo (+2-+8°C) | 0.5 | 16 | 13 | 11 | 0.002288 | 29712 | 38221900 | https://www.hain-lifescience.de/en/products/microbiology/mycobacteria/ntm/genotype-ntm-dr.html
106439 | GT-Blot 48 reagent set for DNA•STRIP Kits (R003) | 1 box | 13 months | Provided free or charge for customers with GT-Blot48 and upon purchase of GenoType assays (ratio: 1 Reagent Set for DNA STRIP Kits for any 4 GenoType Kits of 96 tests purchased). The reagent set for DNA includes: · Strip kits for use with test kits based on the DNA STRIP technology · Reagents that can replace or replenish the solutions provided with DNA Strip kits | Hain LifeScience GMBH |  | 0 | EURO | 0 | Cool cargo (+2-+8°C) | 1.5 | 25 | 15 | 13 | 0.004875 | R003 | 38221900 | https://www.hain-lifescience.de/en/products/microbiology/mycobacteria/tuberculosis/genotype-mtbdrplus.html
106254 | Molecular grade water, 1.7ml | 10 units | N/A | DNAase, RNAase free water - not chemically treated molecular grade water. • Molecular weight: 18,02g/mol; Boiling point: 100°C; Density: 1,0g/cm3; This item includes: 17 ml (10 x 1,7 ml) | TTM |  | 201.85 | USD | 0 | General Cargo | 0.1 | 6.3 | 6.3 | 7.5 | 0.0004 | AppliChem GmbH A8510,1017 | 28539010 | https://www.itwreagents.com/italy/en/product/water+pcr+tested%2C+dna+free%2C+for+molecular+biology/A8510
1.4.3 LPA consumables
106053 | Single channel pipette 1 - 10 µl | 1 unit | N/A | Single channel pipette variable from 1 - 10 µl, to be used with one-hand, made of PVDF and stainless steel. Provided with automatic tip-eject feature; Chemical and UV resistant, also resistant to disinfectant. Autoclavable without handle. | TTM |  | 204.25 | USD | 0 | General Cargo | 1 | 31 | 10.5 | 3.5 | 0.0011393 | Gilson S.A.S F144055M | 84131900 | https://gb.gilson.com/GBSV/catalogsearch/result/?q=F144055M
106054 | Single channel pipette 2 - 20 µl | 1 unit | N/A | Single channel pipette variable from 2 - 20 µl, to be used with one-hand, made of PVDF and stainless steel. Provided with automatic tip-eject feature; Chemical and UV resistant, also resistant to disinfectant. Autoclavable without handle. | TTM |  | 204.25 | USD | 0 | General Cargo | 1 | 31 | 10.5 | 3.5 | 0.0011393 | Gilson S.A.S F144056M | 84131900 | https://gb.gilson.com/GBSV/catalogsearch/result/?q=F144056M
106055 | Single channel pipette 10 - 100 µl | 1 unit | N/A | Single channel pipette variable from 10 - 100 µl, to be used with one-hand, made of PVDF and stainless steel. Provided with automatic tip-eject feature; Chemical and UV resistant, also resistant to disinfectant. Autoclavable without handle. | TTM |  | 204.25 | USD | 0 | General Cargo | 1 | 31 | 10.5 | 3.5 | 0.0011393 | Gilson S.A.S F144057M | 84131900 | https://gb.gilson.com/GBSV/catalogsearch/result/?q=F144057M
106056 | Single channel pipette 20 - 200 µl | 1 unit | N/A | Single channel pipette variable from 20 - 200 µl, to be used with one-hand, made of PVDF and stainless steel. Provided with automatic tip-eject feature; Chemical and UV resistant, also resistant to disinfectant. Autoclavable without handle. | TTM |  | 204.25 | USD | 0 | General Cargo | 1 | 31 | 10.5 | 3.5 | 0.0011393 | Gilson S.A.S F144058M | 84131900 | https://gb.gilson.com/GBSV/catalogsearch/result/?q=F144058M
106057 | Single channel pipette 100 - 1000 µl | 1 unit | N/A | Single channel pipette variable from 100 - 1000 µl, to be used with one-hand, made of PVDF and stainless steel. Provided with automatic tip-eject feature; Chemical and UV resistant, also resistant to disinfectant. Autoclavable without handle. | TTM |  | 204.25 | USD | 0 | General Cargo | 1 | 31 | 10.5 | 3.5 | 0.0011393 | Gilson S.A.S F144059M | 84131900 | https://gb.gilson.com/GBSV/catalogsearch/result/?q=F144059M
106059 | 8-channel pipette 2 - 20 µl | 1 unit | N/A | Laboratory 8 channel pipette, volume 2 -20 µl. Systematic error +/- 0,40; Random error =0.15. Pipette to be used with one-hand only; Provided with automatic tip-eject feature with color coded push-button for easy identification; Autoclavable at 121°C without handle. | TTM |  | 577.3 | USD | 0 | General Cargo | 0.5 | 39.5 | 15.2 | 5.5 | 0.0033022 | Gilson S.A.S FA10009 | 84131900 | https://gb.gilson.com/GBSV/catalogsearch/result/?q=FA10009
106060 | 8-channel pipette 20 - 300 µl | 1 unit | N/A | Laboratory 8 channel pipette, volume 20 -300 µl. Systematic error +/- 3,00; Random error =1.0. Pipette to be used with one-hand only; Provided with automatic tip-eject feature with color coded push-button for easy identification; Autoclavable at 121°C without handle. | TTM |  | 577.3 | USD | 0 | General Cargo | 0.5 | 39.5 | 15.2 | 5.5 | 0.0033022 | Gilson S.A.S FA10015 | 84131900 | https://gb.gilson.com/GBSV/catalogsearch/result/?q=FA10015
106218 | Stand for 8-channel pipettes | 1 unit | N/A | Universal Multichannel stand. It holds any brand of 8 or 12 channel pipettes. Made from rugged materials that are resistant to many chemicals with 3 positions. Compact design for space savings. | TTM |  | 25.95 | USD | 0 | General Cargo | 0.5 | 20 | 20 | 20 | 0.008 | Gilson S.A.S F161417 | 39269097 | https://ch.gilson.com/CHSV/universal-multichannel-stand.html
106554 | Sterile filter tips 0.1-10 µl | 10 x 96 units | 16 months | Sterile filter tips with HDPE filter barrier with volume of 0,1 - 10 µl. RNase, DNase, DNA and Endotoxin-free; Material of tips: PP, Material of filter: PE; Products are sterilized using electron beam irradiation; This item includes: 10 packs with 96 pieces | TTM |  | 115.7 | USD | 0 | General Cargo | 3 | 30 | 18 | 24 | 0.013 | Biotix/Neptune BT10XL | 39269097 | http://neptunescientific.com/wp-content/uploads/2015/11/neptune_bt10xl_2340_series_spec_sheet_110406.pdf
106555 | Sterile filter tips 0.2-20 µl | 10 x 96 units | 16 months | Sterile filter tips with HDPE filter barrier with volume of 0,2 - 20 µl. RNase, DNase, DNA and Endotoxin-free; Material of tips: PP, Material of filter: PE; Products are sterilized using electron beam irradiation; This item includes: 10 packs with 96 pieces. | TTM |  | 115.7 | USD | 0 | General Cargo | 3 | 30 | 18 | 24 | 0.013 | Biotix/Neptune BT20 | 39269097 | http://neptunescientific.com/products/pipette-tips/filter-tips/
106388 | Sterile filter tips 5-100 µl | 10 x 96 units | 16 months | Pipette Filter Tips 5-100 µl, free from DNase, RNase, Human DNA, Pyrogens, PCR Inhibitors. The tips are made of autoclavable, non-wettable, virgin polypropylene with length of 53.40 mm. ·This item includes 10 boxes with 96 units each for a total of 960 units. | TTM |  | 61 | USD | 0 | General Cargo | 1.5 | 33 | 20 | 13.5 | 0.00891 | AHN Biotechnologie GmbH 2-120-C5-0 | 39269097 | http://ahn-bio.de/wp-content/uploads/2018/11/AHN-Consumables-Catalogue.pdf
106389 | Sterile filter tips 20-200 µl | 10 x 96 units | 16 months | Pipette Filter Tips 20-200 µl, free from DNase, RNase, Human DNA, Pyrogens, PCR Inhibitors. The tips are made of autoclavable, non-wettable, virgin polypropylene with length of 53.40 mm. ·This item includes 10 boxes with 96 units each for a total of 960 units. | TTM |  | 61 | USD | 0 | General Cargo | 1.5 | 33 | 20 | 13.5 | 0.00891 | AHN Biotechnologie GmbH 2-128-C5-0 | 39269097 | http://ahn-bio.de/wp-content/uploads/2018/11/AHN-Consumables-Catalogue.pdf
106390 | Sterile filter tips 100-1000 µl | 10 x 96 units | 16 months | Pipette Filter Tips 100-1000 µl, free from DNase, RNase, Human DNA, Pyrogens, PCR Inhibitors. The tips are made of autoclavable, non-wettable, virgin polypropylene with length of 84.30 mm. ·This item includes 10 boxes with 96 units each for a total of 960 units. | TTM |  | 88.5 | USD | 0 | General Cargo | 2.2 | 49 | 26 | 11 | 0.014014 | AHN Biotechnologie GmbH 2-203-C4-0 | 39269097 | http://ahn-bio.de/wp-content/uploads/2018/11/AHN-Consumables-Catalogue.pdf
106061 | Repetitive pipette for combitips | 1 unit | N/A | Hand-dispenser pipette with volume range: from 2 µl to 5 ml. Increased chemical resistance through innovative plastic materials; Easy tip mounting; Up to 49 dispensing steps; One-handed volume adjustment with a dial. - Pipette with a button for contact-free, safe ejection of contaminated dispenser tips; This item includes: shelf/rack mount, 3 PD-Tips: 0.1 ml, 1 ml and 10 ml | TTM |  | 181.4 | USD | 0 | General Cargo | 0.3 | 24 | 17 | 6 | 0.002448 | Brand GmbH 705110 | 84798997 | https://www.sigmaaldrich.com/catalog/product/aldrich/z692867?lang=en&region=GB
106396 | Dispenser tips 5 ml | 100 units | 16 months | Dispenser tips 5 ml hat work well for infectious and volatile liquids, as the fluid never will get in touch with the instrument which avoids any contamination. Direct displacement principle; Pyrogenic, ATP, RNase and DNA-free; · Single sealed and sterilized. Product linked to the hand-dispenser pipette (106115) | TTM |  | 50 | USD | 0 | General Cargo | 0.8 | 28 | 22 | 15 | 0.00924 | Ritter GmbH 40001-0004 | 39269097 | https://www.ritter-medical.de/en/products/ritips/
106398 | Combitips 12.5 ml | 100 units | 16 months | Dispenser tips 12.5 ml hat work well for infectious and volatile liquids, as the fluid never will get in touch with the instrument which avoids any contamination. Direct displacement principle; Pyrogenic, ATP, RNase and DNA-free; · Single sealed and sterilized. Product linked to the hand-dispenser pipette (106115) | TTM |  | 50 | USD | 0 | General Cargo | 1 | 28.5 | 22 | 15 | 0.009405 | Ritter GmbH 40002-0002 | 39269097 | https://www.ritter-medical.de/en/products/ritips/
106412 | Long 1 ml tips with ART barrier | 8x 100 units | 16 months | Long 1 ml tips with ART barrier that perform applications consistently and accurately by using the ART Barrier, which seal when exposed to potential contaminants, trapping them inside the barrier. • Tips are furnished with a unique ART filter to eliminate any risk of cross-contamination between samples; • Sterile, single use; Free of RNase, DNase, DNA and pyrogens and colorless. This item includes: 8 racks with 100 pieces. Product linked to the pipette 200 - 1000 µl (106057) | TTM |  | 90 | USD | 0 | General Cargo | 2.2 | 28 | 23 | 27 | 0.0174 | Fisher Scientific UK Ltd. 2079 | 39269097 | https://www.thermofisher.com/order/catalog/product/2079#/2079
106068 | 96-well rack for PCR tubes of 0.2 ml | 10 units | N/A | 96-Well PCR-Racks for storage of 0.2 ml, and 0.5 ml PCR tubes, made of Polypropylene (PP) with transparent lids. • 8 x 12 positions, micro test plate format, autoclavable • Available in different colors for identification of different types of samples; Thi item includes: 4 x blue, 4 x red, 2 x white and 10 x lids | TTM |  | 25.55 | USD | 0 | General Cargo | 0.2 | 13 | 9 | 10 | 0.00117 | Ratiolab GmbH 4 x 6114102, 4 x 6114104, 2 x 6114100, 10 x 6114111 | 39269097 | http://ratiolab.com/en/167-96-well-pcr-racks
106446 | Quick chiller for tubes 0.2 to 1.5 ml | 1 unit | N/A | Quick chiller for tubes 0.2 to 1.5 ml unit rapidly and repeatedly chills biological samples to -20°C for use in molecular biology applications such as ethanol precipitation of DNA/RNA samples • The durable polycarbonate is pre-filled with a non-toxic insulating solution, and holds twelve 0.2 to 1.5 mL microcentrifuge tubes. • There is no direct contact between the tubes and the insulating solution so labels and markings won't come off; Tube Array is 3 x 4 and No. of Inserts included is 16; • The Quick Chill unit can be used repeatedly for up to 45 minutes, eliminating the inconvenience of dry ice/alcohol baths | TTM |  | 136.5 | USD | 0 | General Cargo | 1.3 | 12 | 16 | 14 | 0.002688 | Nalge Nunc International Corp. 2600-646 | 39269097 | https://us.vwr.com/store/product/7978910/nalgene-quick-chill-unit-thermo-scientific
106384 | Sterile pp centrifuge tubes 15 ml | 500 units | 30 months | Sterile pp centrifuge tubes 15 ml, made of polypropylene Tubes, conical bottom, natural. Free of detectable DNase, RNase, human DNA; Non-Pyrogenic, non-cytotoxic. ·Polypropylene displays high thermal, mechanical and chemical resistance. It is therefore recommended for the storage of chemical and biological samples. Diameter: 17 mm; height: 120 mm. ·With blue screw cap and with blue graduation and white writing area. Subject to availability from our supplier. Product lead-time may be longer than usual. Please contact your GDF Country Supply Officer for further information | TTM |  | 48.95 | USD | 0 | General Cargo | 3 | 30 | 40 | 30 | 0.04 | Greiner Bio-One 188271 | 39269097 | https://shop.gbo.com/en/row/products/bioscience/covid-19/covid-19-tc-tubes-15/188271.html
106340 | Sterile pp centrifuge tubes 50 ml | 500 units | 30 months | Sterile pp centrifuge tubes 50 ml, made of polypropylene Tubes, conical bottom, natural. • Free of detectable DNase, RNase, human DNA; Non-Pyrogenic, non-cytotoxic; • Polypropylene displays high thermal, mechanical and chemical resistance. It is therefore recommended for the storage of chemical and biological samples; • With blue screw cap and with blue graduation and white writing area. Subject to availability from our supplier. Product lead-time may be longer than usual. Please contact your GDF Country Supply Officer for further information | TTM |  | 45 | USD | 0 | General Cargo | 8.05 | 60 | 40 | 37 | 0.0888 | Greiner Bio-One 227261 | 3926 9097 | https://shop.gbo.com/de/austria/products/bioscience/covid-19/covid-19-zentrifugenroehrchen-50/227261.html
106741 | Rack for 1.5-2 ml tubes, 36 wells | 1 unit | N/A | Tube rack for 1.5-2.0 ml tubes made of polypropylene. Capacity of 36 wells • Rack of 3 rows of 12 wells each, numbered locations; • The racks are autoclavable (121 ° C/20 min) | TTM |  | 25.85 | USD | 0 | General Cargo | 0.02 | 21.5 | 9 | 5 | 0.0009675 | Eppendorf AG 0030 119.819 | 39269097 | https://www.eppendorf.com/product-media/doc/en/235590/Sample-Preparation_Flyer_Tube-Racks_Well-Organized.pdf
1.4.6 BD MAX equipment
106919 | BD MAX Equipment 6C Clinical (441916) | 1 unit | N/A | The BD MAX System for molecular diagnostics fully automates cell lysis, nucleic acid extraction, PCR set-up, target amplification and detection from a variety of specimen types. Processes and analyzes up to 24 specimens per run (i.e. 120 specimens per day). 5 fluorescence channels allow for direct measurement of multiple targets within the same sample. System display available in English, French, Spanish, German and Italian. System includes 1-year warranty and it starts on the date of user acceptance or 13 months after date of shipment, whichever comes first. This item contains: The BD MAX System with all hardware and software necessary to drive and monitor operations, as well as calculate, communicate and store results. Thi item should be procured together with the below items: 106921 BD MAX MDR TB Assay kit, 2x 24 tests (443878) 106922 BD MAX STR Reagent, 48 tests (443806) 106923 BD MAX Starter Kit Accessories (445284) 106925 BD MAX Windows 10 Image Key (443436) 106926 BD MAX Windows 10 License with Cylance (444819) 106927 BD MAX Rack Assay Locking Blue racks (444808) 106928 BD MAX Printer HP M255DW (257819) 106929 BD MAX Cable Printer USB 15 FT (441478) 106930 BD MAX UPS Online SYN SIII 2KVA (257687) 106931 BD MAX External Battery Module SIII 2KVA (257688) 106932 BD MAX Qualification kit, 2x 24 tests (444048) 106685 BD MAX PCR Cartridges, 24 cartridges (437519) 106933 BD MAX MDR-TB Resistance Tool, 48 tests (444522) This item includes the Installation and training only for BD Class 1 countries (please request more information to the GDF team) | BD International |  | 65750 | USD | 0 | General Cargo - large | 168 | 119 | 91 | 94 | 1.0179 | 441916 | 9027900000 | https://moleculardiagnostics.bd.com/bd-max-system/#:~:text=The%20BD%20MAX%E2%84%A2%20System,in%20less%20than%20three%20hours*.
106920 | BD MAX System Installation and Training | 1 unit | N/A | Installation and training (3 days in-country). Installation includes: - Performed on-site by certified service engineers, in accordance with the recommendations and instructions of BD MAX System User Manual - Includes calibration & qualification - Provision of an Installation and Verified Functioning report Training includes: - User training provided on-site by certified Application Specialists with laboratory / scientific backgrounds with prepared reference materials - Scope: response to technical inquires, troubleshooting, on-going support. - Result: upon completion, the user will be able to: load Sample Buffer Tubes and Unitized Reagent Strips onto the sample rack Load reagents into the Unitized Reagent Strips; load a BD MAX PCR Cartridge onto the system; program a Worklist; print and download data onto a USB thumb drive. Additionally, “Administrator” level users will be able to: set up user accounts and passwords and create and edit tests in the Test Editor. BD Max training and installation for only Class 2 and Class 3 countries (please request more information to the GDF team) | BD International |  | 2500 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
1.4.7 BD MAX reagents
106921 | BD MAX MDR TB Assay kit, 2 x 24 tests (443878) | 48 tests | 5 months | Real-time PCR assay detecting M. tuberculosis complex (MTBC) DNA as well as key mutations associated with rifampin (RIF) resistance (rpoB gene) and isoniazid (INH) resistance (katG gene and inhA promoter region) in raw or concentrated sputum sediments prepared from induced or expectorated sputa. To be used on the BD MAX Instrument. · For the initila installation and training it is needed 2 x 24 tests. | BD International |  | 412.68 | USD | 0 | General Cargo | 1.01 | 25.5 | 30.7 | 11.5 | 0.009 | 443878 | 3822190090 | https://www.bd.com/resource.aspx?IDX=35759
106922 | BD MAX STR Reagent, 48 tests (443806) | 48 tests | 8 months | Used to liquefy sputum specimens and reduce the viability of M. tuberculosis complex prior to further processing on the BD MAX Instrument. | BD International |  | 206.57 | USD | 0 | Dangerous cargo UN1824 Class 8; PG II | 1.2 | 19.5 | 15 | 15.5 | 0.00453 | 443806 | 3822190090 | https://www.bd.com/resource.aspx?IDX=35759
106978 | BD MAX ExK DNA-1 Kit, 24 tests (442818) | 24 tests | 10 months | The BD MAX ExK DNA-1 is intended to extract DNA from plasma, serum or neat urine specimens. The DNA Extraction Tube contains a Sample Processing Control sequence from Drosophila melanogaster (GenBank: AC246436; nucleotides 35779 to 35978) cloned in a pUC119 derivative (GenBank: U07650.1). The BD MAX ExK DNA-1 has not been validated for use with any specific test method. this item contains: 1, BD MAX DNA Extraction Tubes (A8) 2. BD MAX ExK DNA-1 Sample Buffer Tube (SB-1) 24(2 x 12 tubes) 3. BD MAX DNA-1 Reagent Strip (24) 4. Conical Tubes (24) 5. Septum Caps (25) | BD International |  | 250 | USD | 0 | General Cargo | 0.93 | 25.5 | 30.7 | 11.5 | 0.009 | 442818 | 8473308000 | https://www.bd.com/en-us/products-and-solutions/products/product-page.442818
1.4.8 BD MAX Consumables
106923 | BD MAX Starter Kit Accessories (445284) | 1 unit | N/A | Accessories mandatory to install and use BD Max instrument. Includes all-in-one touchscreen computer with pre-installed BD MAX Software required for BD MAX Instrument installation. BD MAX Starter Kit includes Touch Pad Keyboard, AIO computer, USB cable, USB key, and HD Barcode Scanner, Stand and BD MAX Windows Software Package | BD International |  | 1475 | USD | 0 | General Cargo | 20.5 | 66.7 | 66.7 | 66.7 | 0.29673 | 443362 | 8471607000 | https://moleculardiagnostics.bd.com/bd-max-system/#:~:text=The%20BD%20MAX%E2%84%A2%20System,in%20less%20than%20three%20hours*.
106925 | BD MAX Windows 10 Image Key (443436) | 1 unit | N/A | Required for BD MAX Instrument installation. | BD International |  | 59.11 | USD | 0 | General Cargo | 0.02 | 28 | 21 | 1 | 0.00029 | 443436 | 8523519000 | https://moleculardiagnostics.bd.com/bd-max-system/#:~:text=The%20BD%20MAX%E2%84%A2%20System,in%20less%20than%20three%20hours*.
106926 | BD MAX Windows 10 License with Cylance (444819) | 1 unit | N/A | Required for BD MAX Instrument installation. | BD International |  | 372 | USD | 0 | General Cargo | 0.02 | 21 | 28.5 | 1 | 0.00059 | 444819 | 4911990000 | https://moleculardiagnostics.bd.com/bd-max-system/#:~:text=The%20BD%20MAX%E2%84%A2%20System,in%20less%20than%20three%20hours*.
106927 | BD MAX Rack Assay Locking Blue racks (444808) | 4 units | N/A | Only racks that fit onto the BD MAX Instrument. Each sample rack holds up to 12 tests (i.e. 12 Unitized Reagent Strips and 12 Sample Buffer Tubes). Two racks can be placed onto the instrument deck for each run. Each sample rack has 3 banks, each bank is composed of 4 adjacent rows, each row accommodates a Unitized Reagent Strip and its associated Sample Buffer Tube. Samples within a bank are processed simultaneously by the 4-nozzle pipette head. This allows for a throughput of 24 tests per run. | BD International |  | 3389 | USD | 0 | General Cargo | 9.2 | 43.5 | 43.5 | 43 | 0.08136 | 444808 | 7326909890 | https://moleculardiagnostics.bd.com/bd-max-system/#:~:text=The%20BD%20MAX%E2%84%A2%20System,in%20less%20than%20three%20hours*.
106928 | BD MAX Printer HP M255DW (257819) | 1 unit | N/A | Compatible color printer to print reports from the BD MAX Instrument. | BD International |  | 722 | USD | 0 | General Cargo | 17.44 | 49 | 29.5 | 50 | 0.07227 | 257819 | 8443321000 | https://moleculardiagnostics.bd.com/bd-max-system/#:~:text=The%20BD%20MAX%E2%84%A2%20System,in%20less%20than%20three%20hours*.
106929 | BD MAX Cable Printer USB 15 FT (441478) | 1 unit | N/A | Cable to connect BD MAX Printer HP M255DW to the BD MAX Instrument. | BD International |  | 41 |  | 0 | General Cargo | 0.216 | 36.1 | 27.3 | 6.5 | 0.0064 | 8544429090 | 8544429090 | https://moleculardiagnostics.bd.com/bd-max-system/#:~:text=The%20BD%20MAX%E2%84%A2%20System,in%20less%20than%20three%20hours*.
106930 | BD MAX UPS Online SYN SIII 2KVA (257687) | 1 unit | N/A | Uninterrupted Power Supply for the BD MAX Instrument. 2 KVA. | BD International |  | 1923 | USD | 0 | General Cargo | 34 | 70 | 24 | 58 | 0.09744 | 257687 | 8504409090 | https://moleculardiagnostics.bd.com/bd-max-system/#:~:text=The%20BD%20MAX%E2%84%A2%20System,in%20less%20than%20three%20hours*.
106931 | BD MAX External Battery Module SIII 2KVA (257688) | 1 unit | N/A | With no power supply, one BD MAX Instrument keeps working 16 minutes. If 4 external batteries are added the System keeps working for 300 minutes without power. Up to 4 external batteries can be added to one UPS. | BD International |  | 895 | USD | 0 | General Cargo | 34 | 70.5 | 56.5 | 25.5 | 0.10157 | 257688 | 8473308000 | https://moleculardiagnostics.bd.com/bd-max-system/#:~:text=The%20BD%20MAX%E2%84%A2%20System,in%20less%20than%20three%20hours*.
106932 | BD MAX Qualification kit, 2 x 24 tests (444048) | 48 tests | 13 months | Required for BD MAX Instrument installation. · For the initial installation and training it is needed 2 x 24 tests. | BD International |  | 376 | USD | 0 | General Cargo | 0.94 | 31 | 25.5 | 11.5 | 0.00909 | 444048 | 3822190090 | https://moleculardiagnostics.bd.com/bd-max-system/#:~:text=The%20BD%20MAX%E2%84%A2%20System,in%20less%20than%20three%20hours*.
106685 | BD MAX PCR Cartridges, 24 cartridges (437519) | 24 tests | 16 months | Microfluidic cartridge for RT-PCR incorporates a proprietary design that allows for rapid amplification and sensitive detection. Each cartridge is build with 24 independently controlled reaction chambers. | BD International |  | 8.64 | USD | 0 | Cool cargo (+2-+8°C) | 0.2838 | 11.2 | 5.8 | 7.7 | 4.9e-05 | 437519 | 3926909790 | https://www.bd.com/en-us/products-and-solutions/products/product-page.437519
106933 | BD MAX MDR-TB Resistance Tool, 48 tests (444522) | 48 tests | 8 months | Quality Control tool to test the BD MAX Instrument. Usage is not standardized and depends on customer needs. | BD International |  | 205 | USD | 0 | General Cargo | 0.36 | 14.5 | 9.5 | 12 | 0.00165 | 444522 | 3821000000 | https://moleculardiagnostics.bd.com/bd-max-system/#:~:text=The%20BD%20MAX%E2%84%A2%20System,in%20less%20than%20three%20hours*.
1.4.9 Fluorotype equipment
106961 | FluoroCycler XT kit (first time user) (1805235) | 1 kit | N/A | The FluoroCycler XT permits the automated amplification of nucleic acids and fluorescence-based detection of target sequences using (real-time) PCR and melting curve technologies. · Capacity: 96 samples. Overvoltage category: II; Noise level: <70 dB (A). Power supply: 110 - 230 V AC +/- 10%, 50 to 60 Hz; Power: 1200 W max.; Fuses: 2x T8A 230V or 2x T12A 110V · The FluroroCycler comes with 1 year manufacturer warranty on parts with one maintenance visit. The warranty starts from documented installation date. Also includes the onsite Installation and Training by an authorized representative of Hain Lifescience to do the unpacking, physical installation, wire connections, power supply check, plug-in and start-up. For installation of UPS a local electrician is needed. The training includes up to 3-days of theoretical and practical instruction on principle use, equipment programming and basic regular maintenance. This kit contains: · The FluoroCycler XT with a Desktop Computer (keyboard, mouse, screen and FluoroSoftware XT-IVD · Barcode scanner, Heat sealer, power cable for specific country for Computer and Screen. · Power cable and additional power cable Type G and Type M for FluoroCycler XT · Installation set: Unpacking instructions, Ethernet cable blue; Toolset for setup (including side cutter, Allen key and cable ties)and relevant Documentation for installation. · 2x FrameStar 96well PCR Plate Roche Style (Box of 50 Ref code: 1876705 · 1x Clear Weld Seal Mark II (Box of 100 REf code: 1877137); Fuses (2x T8A, 2x T12A) and Ethernet cable (black). FluoroCycler XT and FluoroSoftware XT-IVD Operator manual (printed in English). Items not supplied with and needed are a Centrifuge for PCR plates and an UPS 3000 VA with battery pack | Hain LifeScience GMBH |  | 42959 | EUR | 0 | General Cargo | 125 | 200 | 140 | 180 | 1.25 | 1805235 | 90181990 | https://www.hain-lifescience.de/en/products/equipment/pcr/fluorocycler-xt.html
106962 | FluoroCycler XT kit (additional cycler) (1805236) | 1 kit | N/A | The FluoroCycler XT permits the automated amplification of nucleic acids and fluorescence-based detection of target sequences using (real-time) PCR and melting curve technologies. · Capacity: 96 samples. Overvoltage category: II; Noise level: <70 dB (A). Power supply: 110 - 230 V AC +/- 10%, 50 to 60 Hz; Power: 1200 W max.; Fuses: 2x T8A 230V or 2x T12A 110V · The FluroroCycler comes with 1 year manufacturer warranty on parts with one maintenance visit. The warranty starts from documented installation date. Also includes the onsite Installation and Training by an authorized representative of Hain Lifescience to do the unpacking, physical installation, wire connections, power supply check, plug-in and start-up. For installation of UPS a local electrician is needed. The training includes up to 3-days of theoretical and practical instruction on principle use, equipment programming and basic regular maintenance. This kit contains: · The FluoroCycler XT with a Desktop Computer (keyboard, mouse, screen and FluoroSoftware XT-IVD · Barcode scanner, power cable for specific country for Computer and Screen. · Power cable and additional power cable Type G and Type M for FluoroCycler XT · Installation set: Unpacking instructions, Ethernet cable blue; Toolset for setup (including side cutter, Allen key and cable ties)and relevant Documentation for installation. · Fuses (2x T8A, 2x T12A) and Ethernet cable (black). FluoroCycler XT and FluoroSoftware XT-IVD Operator manual (printed in English). Items not supplied with and needed are a Centrifuge for PCR plates and an UPS 3000 VA with battery pack, FrameStar 96well PCR Plate Roche Style, Clear Weld Seal Mark II and a Heat Sealer | Hain LifeScience GMBH |  | 39000 | EUR | 0 | General Cargo | 109 | 200 | 140 | 180 | 1.25 | 1805236 | 90181990 | https://www.hain-lifescience.de/en/products/equipment/pcr/fluorocycler-xt.html
106963 | GenoXtract (8.31.01) | 1 unit | N/A | The GenoXtract is a reliable instrument for nucleic acid extraction from different patient materials. Within less than an hour, DNA or RNA is available for subsequent applications. ONLY VALIDATED FOR FluoroType MTBDR Ver 2.0 (H-62896). The instrument applies an isolation method based on the Magnetic Bead Technology and permits the processing of up to 12 samples in parallel. It features maximum protection against contamination due to an innovative pump assembly and an integrated UV lamp. This item includes: · GenoXtract , GenoXtract rack , Magnet plate, Power adapter , Power cable. · Set of 12 adapters for small reaction tubes , Piercing Tool , Tool for calibration , Calibration set (3x empty cartridges, tips, pumps), 2 OQ cartridges , · Unpacking instructions, Operator’s manual, Short instructions, Checklist for installation, Decontamination protocol. | Hain LifeScience GMBH |  | 8700 | EUR | 0 | General Cargo | 40 | 59 | 58 | 72 | 0.24638 | 8.31.01 | 90181990 | https://www.hain-lifescience.de/en/products/equipment/extraction/genoxtract.html
106964 | Centrifuge for PCR plates (1876728) | 1 unit | N/A | The PCR Plate Spinner is a centrifuge supplied with a rotor to hold 2 PCR microplates, designed for quick spins of samples in PCR plates. Primarily designed for use with 96 or 384 well plates and small volumes, it will accommodate skirted, non skirted and semi skirted styles. The centrifuge for PCR plates comes with 1 year manufacturer warranty on parts. The warranty starts with documented installation date (if centrifuge is delivered together with FluroCycler XT). After the one year warranty only a replacement purchase is possible. | Hain LifeScience GMBH |  | 525 | EUR | 0 | General Cargo | 16.5 | 48 | 40 | 44 | 0.08448 | 1876728 | 84798997 | Not available
106965 | Heat Sealer (1876703) | 1 unit | N/A | The Semi-Automatic Sheet Heat Sealer is a compact semi-automatic sheet heat sealer which is compatible with a wide range of seals and plates of differing designs and heights. With variable temperature and time settings, sealing conditions are easily optimized to produce a 100% tight seal, eliminating sample loss. The Heat Sealer is included in the FluoroCycler XT – 1st time user (1805235). Therefore there is no need to order it separately. It can be requested in the case of a replacement | Hain LifeScience GMBH |  | 3600 | EUR | 0 | General Cargo | 15 | 45 | 35 | 52 | 0.0819 | 1876703 | 84798997 | Not available
1.4.10 Fluorotype reagents
106966 | FluoroLyse Ver 1.0 - 96 extractions (H-51696) | 1 kit of 96 extractions | 13,5 months | The FluoroLyse kit permits the fast and easy manual extraction of bacterial DNA. Depending on the subsequent test, various clinical specimens or cultivated samples can be used as starting material. The FluoroLyse kit permits the fast and easy manual extraction of bacterial DNA for further use with the following diagnostic assays from Hain Lifescience: FluoroType CT, FluoroType MTB, FluoroType MTBDR VER 2.0, FluoroType Mycobacteria and FluoroType NG. This kit is for manual DNA extraction of 96 samples. | Hain LifeScience GMBH |  | 52 | EUR | 0 | Cool cargo (+2-+8°C) | 4.5 | 47 | 46 | 43 | 0.09 | H-51696 | 38221900 | https://www.hain-lifescience.de/en/products/dna-isolation/fluorolyse.html
106967 | X2 Cartridge (12.02.02) | 1 kit with 96 extraction cartridges | 9,5 months | The X2 Cartridge VER 1.0 is a kit for automated extraction of DNA from various clinical specimens using dedicated extraction instruments for subsequent use in diagnostic procedures. ONLY VALIDATED FOR FluoroType MTBDR Ver 2.0 (H-62896). The X2 Cartridge VER 1.0 is based on Magnetic Bead Technology. The automated extraction of DNA comprises the following four steps: (i) coupling of bacterial cells to Magnetic Beads, (ii) washing of the Bacteria-Bead complex, (iii) bacterial cell lysis and release of DNA under alkaline conditions, and (iv) neutralization. This item is for 96 extraction cartridges. | Hain LifeScience GMBH |  | 432 | EUR | 0 | General cargo | 5 | 39 | 30 | 28 | 0.03276 | 12.02.02 | 38221900 | https://www.hain-lifescience.de/en/products/dna-isolation/gxt-extraction-kits.html MSDS: https://www.hain-lifescience.de/en/downloads/material-safety-data-sheet.html
106968 | Inactivation Set VER 1.0 (522960) | 1 kit up to 960 reactions | 13,5 months | For Inactivation of Mycobacteria in Decontaminated Patient Specimens and Cultured Material (Liquid Medium) in use with GenoXtract and X2 Cartridge. ONLY VALIDATED FOR FluoroType MTBDR Ver 2.0 (H-62896). This set is for up to 960 reactions. | Hain LifeScience GMBH |  | 99 | EUR | 0 | Dangerous cargo UN 3316, class 9; PG II | 0.5 | 16.5 | 13 | 10 | 0.002145 | 522960 | 38221900 | MSDS: https://www.hain-lifescience.de/en/downloads/material-safety-data-sheet.html
106969 | FluoroType MTBDR Ver 2.0 - 96 tests (H-62896) | 1 kit of 96 tests | 13,5 months | The FluoroType MTBDR VER 2.0 is a qualitative in vitro test for automated detection of the Mycobacterium tuberculosis (MTB) complex and its resistance to rifampicin (RMP) and/or isoniazid (INH) from sputum specimens and cultivated samples. Cycler running time approx.: 1:25 up to 2h (depending on cycler and sample material). This item is for 96 tests. | Hain LifeScience GMBH |  | 811.2 | EUR | 0 | Dangerous cargo UN 1845. Dry ice (-80 to -18°C) | 11 | 47 | 46 | 43 | 0.09 | H-62896 | 38221900 | https://www.hain-lifescience.de/en/products/microbiology/mycobacteria/tuberculosis/fluorotype-mtbdr.html
1.4.11 Fluorotype consumables
106970 | FrameStar 96well PCR Plate (1876705) | 50 units | N/A | FrameStar 96 Well Semi-Skirted PCR Plate, Roche Style. Low profile, 0.1ml polypropylene wells, polycarbonate frame, cut corner H12; working volume: <100µl, total well capacity: 200µl; designed for use on Roche LightCycler 96 and 480 (with 96 well block). | Hain LifeScience GMBH |  | 254 | EUR | 0 | General Cargo | 1.5 | 27 | 23 | 9 | 0.005589 | 1876705 | 39231010 | Not available
106971 | Clear Weld Seal Mark II (1877137) | 100 units | N/A | Clear Weld Heat Seal Mark 2; optically clear heat sealing film, sheets (125 x 78 mm). Optically clear heat sealing film, non-peelable, difficult to pierce; suitable for qPCR, optical applications and storage | Hain LifeScience GMBH |  | 105 | EUR | 0 | General Cargo | 0.1 | 15 | 10 | 3 | 0.00045 | 1877137 | 39231010 | Not available
1.4.12 Fluorotype services
106972 | LabScape Complete FluoroCycler XT 2 years for 1 unit (DAL00550) | 1 unit | N/A | 2 years on-site repair-support and 1 Maintenance visit per year (LabScape complete/full service contract) for the FluoroCycler XT (H-2017, for 1 device, including spare parts). Online support is free of charge. During online support it will be determined if on-site visit is necessary. This Support and Maintenance is limited to 5 years (after installation of device). | Hain LifeScience GMBH |  | 9600 | EUR | 0 | Service | N/A | N/A | N/A | N/A | N/A | DAL00550 | N/A | Not available
106973 | LabScape Complete GenoXtract XT 2 years for 1 unit (DAL00551) | 1 unit | N/A | 2 years on-site repair-support and 1 Maintenance visit per year (LabScape complete/full service contract) for the first GenoXtract (for 1 device, including spare parts). Online support is free of charge. During online support it will be determined if on-site visit is necessary. This Support and Maintenance is limited to 5 years (after installation of device). | Hain LifeScience GMBH |  | 4800 | EUR | 0 | Service | N/A | N/A | N/A | N/A | N/A | DAL00551 | N/A | Not available
106974 | LabScape Complete GenoXtract XT 2years for 2 units (DAL00552) | 1 unit | N/A | 2 years on-site repair-support and 1 Maintenance visit per year (LabScape complete/full service contract) for each further GenoXtract in addition to first one (for additional GenoXtract device, including spare parts). Online support is free of charge. During online support it will be determined if on-site visit is necessary. This Support and Maintenance is limited to 5 years (after installation of device). | Hain LifeScience GMBH |  | 900 | EUR | 0 | Service | N/A | N/A | N/A | N/A | N/A | DAL00552 | N/A | Not available
1.5 SPECIMEN CULTURE AND DST
1.5.1 Solid culture equipment
106843 | Incubator 245 lt | 1 unit | N/A | Laboratory incubator 245 lt, PID microprocessor control with integrated auto-diagnostic system. Calibrated 37 °C. Temperature range 5°C above ambient temperature up to +100°C. Setting accuracy 0.1°C; Housing Stainless steel; · Forced air circulation by air turbine with noise no greater than 70 db. This item includes: 1 shelf or grid and 1 additional perforated shelf. | AMEX |  | 2892.5 | USD | 0 | General Cargo | 130 | 87 | 92 | 124 | 0.99 | Pol-Eko-Aparatura CLW 240 STD INOX/G | 84198990 | https://www.pol-eko.com.pl/model/laboratory-incubator-clw/laboratory-incubator-clw-240/
106844 | Shelf for incubator 245 lt | 1 unit | N/A | Perforated stainless steel shelf, made of stainless steel. Shelf support 15kg load. | AMEX |  | 56.9 | USD | 0 | General Cargo | 2 |  |  |  | 0.012 | Pol-Eko- Aparature CL/IL/SL/SR/KK 240 | 84198990 | https://www.pol-eko.com.pl/model/laboratory-incubator-clw/laboratory-incubator-clw-240/
106247 | Incubator 749 lt | 1 unit | N/A | Laboratory incubator 749 lt, PID microprocessor control with integrated auto-diagnostic system. Calibrated 37 °C. Temperature range min. 10 above ambient temperature up to +80°C. Setting accuracy 0.1°C; Housing Stainless steel 1.4016; · Forced air circulation by 2 quiet air turbines, adjustable in 10 % steps; approx. 55 dB(A); This item includes: 3 stainless steel shelves. | TTM |  | 6399.5 | USD | 0 | General Cargo Large | 290 | 105 | 133 | 191 | 2.6673 | Memmert GmbH + Co. KG IF750 | 8419899890 | https://www.memmert.com/products/incubators/incubator/IF750/
106248 | Shelf for incubator 749 lt | 1 unit | N/A | Perforated stainless steel shelf, made of stainless steel 1.4301. shelfs support a max. load of 20kg. | TTM |  | 183 | USD | 0 | General Cargo | 15 | 141 | 69 | 21 | 0.20431 | Memmert GmbH + Co. KG B00328 | 8419908509 | https://www.memmert.com/products/incubators/incubator/IF750/
1.5.2 Solid culture reagents
106038 | BBL Lowenstein-Jensen Medium + PACT, tubes (220502) | 100 units | 8 months | Lowenstein-Jensen Medium + PACT is an egg-based medium containing antimicrobials for the cultivation of Mycobacterium tuberculosis and other mycobacterial species. This item contains 100 x BBL Lowenstein-Jensen Medium + PACT Size A tubes. This item has a very short shelf life. Need to order according to consumption. | BD International |  | 184.54 | USD | 0 | Cool cargo (+2-+8°C) | 4 | 35 | 20 | 16 | 0.0112 | 220502 | 3821000000 | https://legacy.bd.com/resource.aspx?IDX=20666
106039 | BBL Lowenstein-Jensen Medium Slants, tubes (220909) | 100 units | 12 months | Egg-based medium used for the isolation and cultivation of M. tuberculosis and other mycobacterial species. Lowenstein-Jensen medium, tubed as deeps is used for the semi-quantitative catalase test as an aid to the classification of mycobacteria. This item contains 100 x BBL Lowenstein-Jensen Medium Slants Size A tubes. | BD International |  | 157.51 | USD | 0 | Cool cargo (+2-+8°C) | 4.1 | 35.2 | 19.8 | 16 | 0.011151 | 220909 | 3821000000 | https://legacy.bd.com/ds/technicalCenter/inserts/L007464(11).pdf
106559 | BBL Middlebrook OADC Enrichment, 100 ml (212240) | 6 units | 8 months | Used to supplement Difco Middlebrook 7H10 Agar (262710) in qualitative procedures for the isolation and cultivation of mycobacteria. This item includes: 6 bottles of 100ml. This item has a very short shelf life. Need to order according to consumption. | BD International |  | 241.17 | USD | 0 | Cool cargo (+2-+8°C) | 1.36 | 20.2 | 14 | 14.8 | 0.004185 | 212240 | 3821000000 | https://legacy.bd.com/europe/regulatory/Assets/IFU/Difco_BBL/212240.pdf
106614 | BBL Seven H11 Agar, 500 g (212203) | 1 unit | 32 months | Used in qualitative procedures for isolation and cultivation of mycobacteria, especially M. tuberculosis, from clinical and nonclinical specimens. This item contains 1 bottle of dehydrated agar of 500g. | BD International |  | 96.13 | USD | 0 | General Cargo | 0.6 | 12 | 7 | 20 | 0.00168 | 212203 | 3821000000 | https://legacy.bd.com/europe/regulatory/Assets/IFU/Difco_BBL/283810.pdf
106615 | Difco Middlebrook 7H10 Agar, 500 g (262710) | 1 unit | 40 months | Used for the isolation, cultivation and susceptibility testing of mycobacteria. Used with Middlebrook OADC Enrichment (cat. no. 212240) in qualitative procedures for the isolation and cultivation of mycobacteria. This item contains: 1 bottle of dehydrated agar of 500g. | BD International |  | 86.39 | USD | 0 | General Cargo | 0.59 | 12 | 6.5 | 19.5 | 0.00152 | 262710 | 3821000000 | https://legacy.bd.com/europe/regulatory/Assets/IFU/Difco_BBL/212351.pdf
106620 | BBL Lowenstein-Jensen Medium Slants, tubes (220908) | 10 units | 12 months | Egg-based medium used for the isolation and cultivation of M. tuberculosis and other mycobacterial species. Lowenstein-Jensen medium, tubed as deeps is used for the semi-quantitative catalase test as an aid to the classification of mycobacteria. This item contains 10 x BBL Lowenstein-Jensen Medium Slants Size A tubes. | BD International |  | 17.17 | USD | 0 | Cool cargo (+2-+8°C) | 0.28 | 14 | 6.6 | 15.6 | 0.001441 | 220908 | 3821000000 | https://legacy.bd.com/ds/technicalCenter/inserts/L007464(11).pdf
106259 | TB medium LJ base, 500 g | 1 unit | 16 months | Lowenstein-Jensen Medium is used for the cultivation of Mycobacterium tuberculosis and other mycobacterial species. Color: light to medium blue-green; · Approximate Formula Per 600 mL Purified Water: Monopotassium Phosphate: 2.5 g; Potato Flour: 30.0 g; Magnesium Sulphate: 0.24 g ; Malachite Green: .0.4 g; Sodium Citrate;: 0.6 g; Glycerol: 12.0 mL; L-Asparagine: 3.6 g; Whole Egg: 1000.0 m | TTM |  | 69.55 | USD | 0 | General Cargo | 0.6 | 12.5 | 7 | 19.5 | 0.0017063 | BD 244420 | 38210000 | https://legacy.bd.com/europe/regulatory/Assets/IFU/Difco_BBL/244420.pdf
106268 | Glycerol 99.5, 1 lt | 1 lt | 16 months | Glycerol for analyses, Cas No. 56-81-5. · Reagent grade/purum ≥99,5%, comply, p.a., anhydrous | TTM |  | 10.55 | USD | 0 | General Cargo | 1.2 | 20 | 20 | 25 | 0.01 | Carl Roth GmbH 3783.1 | 29054500 | https://www.carlroth.com/de/en/cryo-conservation/glycerol/p/3783.1
106265 | L-Asparagine-monohydrate, 100 g | 1 unit | 16 months | L-Asparagin-Monohydrate for biochemical usage, CAS No 5794-13-8. · Reagent grade/purum ≥ 99%; Subject to availability from our supplier. Product lead-time may be longer than usual. Please contact your GDF Country Supply Officer for further information | TTM |  | 21.6 | USD | 0 | General Cargo | 0.3 | 10 | 10 | 12 | 0.0012 | AppliChem GmbH A3721,0100 | 29241200 | https://itwreagents.com/germany/de/product/l-asparagin-monohydrat-biochemica/A3721
106267 | Malachite green oxalate, 100 g | 1 unit | 30 months | Malachite green oxalate for microscopy. CAS No 2437-29-8. · Reagent grade/purum ≥90% | TTM |  | 112.95 | USD | 0 | Dangerous cargo UN 2811, Class 6.1, PG III | 6 | 21 | 28 | 32 | 0.018816 | Merck KGaA 1.15942.0100 | 32041300 | https://www.merckmillipore.com/CH/de/product/Malachite-green-oxalate-C.I.-42000,MDA_CHEM-115942?ReferrerURL=https%3A%2F%2Fwww.google.com%2F
106279 | Di-Sodium hydrogen phosphate dodecahydrate, 1 kg | 1 unit | 31 months | Di-Sodium hydrogen phosphate dodecahydrate for analyses, CAS number 10039-32-4; · Reagent grade/purum ≥98%, Ph.Eur., USP | TTM |  | 17.15 | USD | 0 | General Cargo | 1.2 | 20 | 20 | 25 | 0.01 | Carl Roth GmbH N350.1 | 28352200 | https://www.carlroth.com/de/de/von-a-bis-z/di-natriumhydrogenphosphat-dodecahydrat/p/n350.1
106303 | Skim milk powder, 500 g | 1 unit | 12 months | Skim milk powder for microbiology. CAS No 999999-99-4. · Assay protein ≥35%; Lactose (enymatically) 40 - 52 %; · Bulk density appr. 500 kg/m³; pH 6-8 at 20 g/l 20°C. | TTM |  | 45.65 | USD | 0 | General Cargo | 0.7 | 20 | 20 | 10 | 0.004 | Merck KGaA 1.153.630.500 | 4021011 | https://www.sigmaaldrich.com/catalog/product/mm/115363?lang=en&region=GB&cm_sp=Insite-_-rvRecBlock_recentlyViewed_userHistory-_-recentlyViewed5-4
1.5.3 Solid culture consumables
106235 | Inoculation loop 3 mm | 1 unit | N/A | Inoculating loop, Ø: 3 mm, length: 60mm, made of platinum-iridium wire · They are heat resistant and can be flamed to sterilise the loop. | TTM |  | 73.55 | USD | 0 | General Cargo | 0.005 | 4 | 0.3 | 0.3 | 3.6e-07 | Witeg Labortechnik GmbH 7414172 | 71101910 | https://www.witeg.de/en/products/consumables/accessories/inoculation-tools/inoculating-loop-oe-3-mm
106234 | Inoculation loop 4 mm | 1 unit | N/A | Inoculating loop, Ø: 3 mm, length: 60mm, made of platinum. · They are heat resistant and can be flamed to sterilise the loop. | TTM |  | 73.55 | USD | 0 | General Cargo | 0.01 | 6 | 1.5 | 1.5 | 1.35e-05 | WITEG 9318504 | 71101910 | https://www.witeg.de/en/detail/index/sArticle/11260
106714 | Disposable inoculation loops 10 µl | 600 units | 16 months | Disposable inoculation loops 10 µl, flexible, made of polystyrene (PS). Length: 200 mm; Single use. Free of lubricants, oils and electrostatic charges and individually wrapped. | TTM |  | 16.8 | USD | 0 | General Cargo | 2.9 | 31 | 27 | 33 | 0.02125 | Greiner Bio-One GmbH 731171 | 39269097 | https://shop.gbo.com/en/england/products/bioscience/microbiology-bacteriology/disposable-inoculation-loops-needles/731171.html
106098 | Lime glass culture tubes 12 ml | 100 units | N/A | Culture tubes 12 ml with screw caps for cultivation, storage and transport of microorganisms. Disposable Culture Tube, with TPE seal, GL 18, Ø 16 x 100 mm and PP screw cap from Soda-lime Glass. · Round bottom; Thread csp GL 18; Wall Thickness 1 mm; External Ø 16 mm/height 100 mm; · High chemical resistance; Heat resistant to 140 °C | TTM |  | 43.6 | USD | 0 | General Cargo | 1.6 | 34 | 11 | 14 | 0.005236 | DWK Life Sciences GmbH 231751459 | 70172000 | https://www.dwk.com/disposable-culture-tube-with-tpe-seal-gl-18-16-x-100-mm-231751459
106739 | Borosilicate glass culture tubes 12 ml with screw caps | 100 units | N/A | Borosilicate glass culture tubes 12 ml with screw caps. Capacity of 20 ml, round bottom. Non sterile with High chemical and heat resistance. High tightness due to PTFE-coated silicone seal; · With DIN thread and screw cap from PBT for airtight sealing (also available without the screw cap); · Cultivating, storage and transport of microorganisms; This item includes: 2 bags of 50 units each, for a total of 100 units. | TTM |  | 325.6 | USD | 0 | General Cargo | 6.5 | 46 | 23 | 18 | 0.019044 | DWK Life Sciences GmbH 261352254 | 70172000 | https://www.dwk.com/duran-culture-tube-with-screw-cap-gl-18-16-x-150-mm-261352254
106339 | Glass bottle 30 ml | 1 unit | N/A | Universal bottle for samples storage. Capacity 30 ml with screw caps made of PP with rubber seals. Autoclavable at 121°C. · Bottles with screw thread and wide neck; clear glass. · Made of transparent soda lime glass, hydrolytic class 2. Chemical and heat resistant; · The outer diameter is 36 mm/height: 68 mm | TTM |  | 0.6 | USD | 0 | General Cargo | 0.05 | 3 | 3 | 8 | 7.2e-05 | Witeg Labortechnik GmbH 5894030 | 70172000 | https://www.witeg.de/en/products/laboratory-glassware/bottles/reagent-bottles/reagent-bottles-with-screw-cap-wide-neck-clear-glass?number=5894030
106849 | Glass bottle 50 ml | 10 units | N/A | Universal bottle for samples storage. Capacity 50 ml with screw caps made PP of aluminium with rubber seals. Autoclavable at 121°C. · Bottles with screw thread and wide neck; clear glass. · Made of transparent soda lime glass, hydrolytic class 2. Chemical and heat resistant. | TTM |  | 10.7 | USD | 0 | General Cargo | 0.05 | 4 | 4 | 8 | 0.000128 | Witeg Labortechnik GmbH 5894050 | 70172000 | https://www.witeg.de/en/products/laboratory-glassware/bottles/reagent-bottles/reagent-bottles-with-screw-cap-wide-neck-clear-glass?number=5894050
106432 | Slanted rack for 16 mm tubes | 1 unit | N/A | Test tube racks, inclined for 16 mm tubes. The unique end plate design allows 5° or 20° slant when the rack is set down on either side to allow media to harden at an angle for incubating liquid cultures. · Heat resistant up to +135 °C; Capacity 4x10 holes; Autoclavable; Rack in white/transparent; · Alphanumeric labelling of rack positions for easy identification of samples. Subject to availability from our supplier. Product lead-time may be longer than usual. Please contact your GDF Country Supply Officer for further information | TTM |  | 44.45 | USD | 0 | General Cargo | 0.4 | 20 | 10 | 8 | 0.0016 | Nalge Nunc International Corp. 5935-0016 | 39269097 | https://www.thermofisher.com/order/catalog/product/5935-0016#/5935-0016
106496 | Slanted rack for 20 mm tubes | 1 unit | N/A | Test tube racks, inclined for 20 mm tubes. The unique end plate design allows 5° or 20° slant when the rack is set down on either side to allow media to harden at an angle for incubating liquid cultures. · Heat resistant up to +135 °C; Capacity 4x10 holes; Autoclavable; Rack in white/transparent; · Alphanumeric labelling of rack positions for easy identification of samples. Subject to availability from our supplier. Product lead-time may be longer than usual. Please contact your GDF Country Supply Officer for further information | TTM |  | 44.45 | USD | 0 | General Cargo | 0.4 | 20 | 10 | 8 | 0.0016 | Nalge Nunc International Corp. 5935-0020 | 39269097 | https://www.thermofisher.com/order/catalog/product/5935-0016#/5935-0016
1.5.4 Liquid culture equipment
106013 | BACTEC MGIT 960 System (445870 including 440107 + 445941) | 1 unit | N/A | Non-radiometric, fully automated mycobacterial detection system for the rapid detection of mycobacteria in clinical specimens, other than blood, as well as DST of anti-tuberculosis drugs known as S.I.R.E., I.R., and PZA. Designed for medium- and high-volume laboratories: holds 320 tubes (3 drawers) for an annual capacity of approximatively 2,700 specimens per year. This item includes: • One Starter Kit (440107) including: Users Manual; Quick Ref Guide; Label Barcode BBL MGIT Tube; Station Plug Detail; Temperature QC Device; Kit Calibrator MGIT 17 Pk; Cord 17255 Belden; Cordset 3 Cond Detachable; Cordset UK 6A/250V Black; Filter Air 5.5in X 5.5in; Filter Intake; Software Bactec MGIT USB Packaged; Carton 14.25in x 11.12 x 7.5in; Bubble Bag 4in x 7-1/2in; Label Drawer Needs Service; Label All Instrument (Label MGIT Tube Brkg Warn, Label MGIT Tube Brkg Jap, Label MGIT Tube Brkg Span, Label MGIT Tube Brkg Ital, Label MGIT Tube Brkg Germ, Label MGIT Tube Brkg French) • One BACTEC MGIT AST Starter Kit (445941) including: 'User Manual' and 'Quick reference guide'; 2 x Transport rack; 16 x 5-tube AST Set carrier; 16 x 2-tube AST Set carrier; 3 x 3-tube AST Set carrier; 3 x 4-tube AST Set carrier and 3 x 8-tube AST Set carrier. · Includes 1-year warranty. Please add to your request the BACTEC printer, UPS 3000 VA with its battery pack and installation and training. | BD International |  | 42391.47 | USD | 0 | General Cargo Large | 421.84 | 100.5 | 160 | 157.5 | 2.5326 | 445870 including 440107 + 445941 | 9027900000 | https://www.bd.com/resource.aspx?IDX=15213
106596 | BACTEC MGIT 320 System - 441743 (including 440107 + 445941) | 1 unit | N/A | Non-radiometric, fully automated mycobacterial detection system for the rapid detection of mycobacteria in clinical specimens, other than blood, as well as DST of anti-tuberculosis drugs known as S.I.R.E., I.R., and PZA. Designed for medium- and high-volume laboratories: holds 320 tubes (3 drawers) for an annual capacity of approximatively 2,700 specimens per year. This item includes: • One Starter Kit (440107) including Users Manual; Quick Ref Guide; Label Barcode BBL MGIT Tube; Station Plug Detail; Temperature QC Device; Kit Calibrator MGIT960 51 Pk; Cord 17255 Belden; Cordset 3 Cond Detachable; Cordset UK 6A/250V Black; Filter Air 5.5in X 5.5in; Filter Intake; Software Bactec MGIT USB Packaged; Carton 14.25in x 11.12 x 7.5in; Bubble Bag 4in x 7-1/2in; Label Drawer Needs Service; Label All Instrument (Label MGIT960 Tube Brkg Warn, Label MGIT960 Tube Brkg Jap, Label MGIT960 Tube Brkg Span, Label MGIT960 Tube Brkg Ital, Label MGIT960 Tube Brkg Germ, Label MGIT960 Tube Brkg French) • One BACTEC MGIT AST Starter Kit (445941) including: User Manual and Quick reference guide; 2 x Transport rack; 16 x 5-tube AST Set carrier; 16 x 2-tube AST Set carrier; 3 x 3-tube AST Set carrier; 3 x 4-tube AST Set carrier and 3 x 8-tube AST Set carrier · Includes 1-year warranty. Please add to your request the BACTEC printer, UPS 3000 VA with its battery pack and installation and training. | BD International |  | 22798.61 | USD | 0 | General Cargo | 180 | 101 | 93 | 87 | 0.8084 | 441743 including 440107 + 445941 | 9027899000/9027900000 | https://www.bd.com/resource.aspx?IDX=22628
106408 | BACTEC MGIT System Installation and Training (01SM0001) | 1 unit | N/A | Installation includes: · Includes: unpacking, assembly and interconnection of all components and accessories, power up, verification of instrument operation through the running of all diagnostics tests and quality checks. Provision of documentation (i.e. 'Pre-installation procedure' and 'System Test'). Instrument introduction and provision of a Service Report · Performed in-country by staff from BD or BD certified partner Training includes: · Duration: 3 days (in-country) · Scope: application of procedures to work with instruments and application of MGIT technology (i.e. preparation of material and detection procedures, sample preparation, operation of samples on BACTEC MGIT equipment, 1st line DST procedure and 2nd line procedure). At the end of the training, the participant will be able to: set up and analyze assays, perform the maintenance requirements, follow Best Practices for troubleshooting and use reference material (e.g. User Manual, Pack Inserts) · Unlimited number of participants | BD International |  | 4200 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | 01SM0001 | N/A | Not available
106016 | BACTEC Printer Brother L5000D (257710) | 1 unit | N/A | Monochrome laser printer with a Parallel interface for wide application compatibility. Printing speed up to 42ppm with high-yield replacement toner cartridge. Automatic two-sided printing. Built-in parallel and Hi-Speed USB 2.0 interfaces. 250-sheet capacity paper tray for letter or legal-size documents. Print envelopes and alternative media types from the 50-sheet capacity multi-purpose tray. Paper capacity can be expanded to 1,340 sheets with optional add-on trays. • Print Resolution: maximum 1,200 x 1,200 dpi • Emulations: PCL6, BR-Script3, IBM Proprinter, Epson FX, PDF version 1.7, XPS Version 1.0 • 1 year limited warranty with toll-free technical support for the life of your machine | BD International |  | 753.96 | USD | 0 | General Cargo | 13.5 | 45 | 48 | 38 | 0.08208 | Brother L5000D - 257710 | 8443321000 | https://legacy.bd.com/ds/technicalCenter/promotionalFlyers/ss-BACTEC_MGIT_Specifications.pdf
106451 | Densitometer set | 1 unit | 10 months | Densitometer for measurement of cell suspension's turbidity. Measurement range at least from 0.5 to 12 McF; Accuracy at least +/- 3% or better. Wavelength in the range 550 to 600 nm; This item includes: Complete set with standards 0.5, 1.0, 2.0, 3.0, 4.0; Tube sizes: 15 x 100-105mm and a comparison cart. | TTM |  | 699.05 | USD | 0 | General Cargo | 1 | 24.5 | 24.5 | 15.5 | 0.0093039 | Biosan / Remel Biosan Densitometer BS-050102 + A-16 Adapter and Remel 20421 | 90275000 | https://biosan.lv/products/den-1/
106475 | Slide warmer | 1 unit | N/A | Slide warmer for stretching and drying of tissue samples in histology, pathology and bacteriology laboratories. The slide warmer provides constant temperature by a 6 mm (0.24″) thick aluminium working plate with black plastic coating. · The temperature is reached fast and constantly maintained. Temperature range: 30 to 99 °C · An integrated safety cut-out prevents the system from being overheated | TTM |  | 1056.75 | USD | 0 | General Cargo | 3.2 | 20 | 32 | 30 | 0.0192 | MEDITE OTS 40-3040 | 84798200 | https://www.medite.de/en/portfolio/stretching-table-ots-40-3040/
106085 | Pipette filler 1 to 200 ml - battery | 1 unit | N/A | Battery powered pipette filler for all pipetting from 1 ml to 200 ml with rechargeable battery, up to 50 hour use, with charge indicator. Use with volumetric, serological and plastic pipettes, volumes 0.1 to 200 mL. • Display of selected pipetting mode, EX or blow-out;One button operation; • Indication of the selected pipetting speed; Charges automatically just by inserting into stand; • Buzzer signal sounds when battery capacity reaches below 10%; • Adjustable suction speed, depending upon volume and viscosity of medium; | TTM |  | 429.3 | USD | 0 | General Cargo | 1 | 24 | 23 | 8 | 0.0044 | Hirschmann Laborgeräte GmbH 9907200 | 84223000 | https://www.hirschmann-laborgeraete.de/en/artikelgruppe/99072
106115 | Dispenser pipette 10 ml | 1 unit | N/A | The Dispenser pipette 10 ml for precise dispensing of all types of liquid, with no risk of contamination by aresols. Battery powered with digital display to show the dispensed volume; · Battery autonomy at least 1 hour of continuous use; Dispensing range at least: 1µl to 10 ml · Possibility to dispense 20 different volumes with one size Combitip; One button tip ejector with volume selector dial; Accuracy ±0,5%; Precision of at least 0.6 to 0.25 µl; This item includes: charging adapter, autoswitching from 100 to 240V, rechargeable battery (4980215003). | TTM |  | 429.15 | USD | 0 | General Cargo | 1.2 | 43 | 17 | 15 | 0.010965 | Eppendorf AG 4982000012 | 84798997 | https://online-shop.eppendorf.com/OC-en/Manual-Liquid-Handling-44563/Manual-Pipetting--Dispensing-44564/Multipette-M4-PF-44619.html
106180 | Electronic pipette dispenser 50 ml | 1 unit | N/A | Electronic dispenser with digital display, dispensing range at least 1µl to 50 ml; One button tip ejector. · Line or rechargeable battery powered; Accuracy ±0,6% to ±2,5% ; Precision < 0,3% to 0,8%; · With dial to select the desired operation mode: pipetting mode, dispensing mode or automatic mode; This item includes: Battery charger | TTM |  | 761.1 | USD | 0 | General Cargo | 2 | 20 | 20 | 35 | 0.014 | Eppendorf AG 4987000010 | 84798997 | https://online-shop.eppendorf.com/OC-en/Manual-Liquid-Handling-44563/Manual-Pipetting--Dispensing-44564/MultipetteE3-E3x-PF-135444.html
1.5.5 Liquid culture reagents
106273 | BBL Middlebrook 7H9 Broth with glycerol, tubes, 5 ml (221832) | 10 units | 8 months | Non-selective culture medium used in qualitative procedures for isolation and cultivation of mycobacteria, when supplemented with Middlebrook ADC Enrichment (212352). This item contains 10 x BBL Middlebrook 7H9 Broth w/ glycerol Tube 5 ml This item has a very short shelf life. Need to order according to consumption. | BD International |  | 29.77 | USD | 0 | Cool cargo (+2-+8°C) | 0.2 | 12 | 5 | 11 | 0.00066 | 221832 | 3821000000 | https://legacy.bd.com/ds/technicalCenter/inserts/L007467(12).pdf
106558 | Difco Middlebrook 7H9 Broth, 500g (271310) | 1 unit | 40 months | Non-selective culture medium used in qualitative procedures for isolation and cultivation of mycobacteria, when supplemented with either glycerol or polysorbate 80 (Tween 80) and Middlebrook ADC Enrichment (212352). This item contains 1 bottle of dehydrated agar of 500g. | BD International |  | 43.3 | USD | 0 | General Cargo | 0.6 | 12.1 | 7 | 19.3 | 0.0016347 | 271310 | 3821000000 | https://legacy.bd.com/europe/regulatory/Assets/IFU/Difco_BBL/212352.pdf
1.5.6 Liquid culture consumables
106272 | BACTEC MGIT Calibrators Kit for 1 drawer (445871) | 17 vials | 16 months | BACTEC MGIT Calibrators Kit for 1 drawer. This item contains 17 vials. 17 vials are sufficient for 1 drawer. | BD International |  | 175 | USD | 0 | General Cargo | 0.36 | 15 | 10.5 | 11.7 | 0.0018427 | 445871 | 9027900000 | https://www.bd.com/resource.aspx?IDX=27408
106918 | BACTEC MGIT Calibrators Kit for 3 drawers (445999) | 51 vials | 16 months | BACTEC MGIT Calibrators Kit for 3 drawers. This item contains 51 vials. 51 vials are sufficient for 3 drawers. | BD International |  | 525 | USD | 0 | General Cargo | 0.94 | 21 | 17.4 | 11.8 | 0.00431 | 445999 | 9027900000 | Not available
106026 | BACTEC MGIT Tubes 7ml (245122) | 100 units | 12 months | The MGIT Tubes are intended for the detection and recovery of mycobacteria using the BACTEC MGIT System. Each plastic tube has a screw-top cap and contains 7 ml of modified Middlebrook 7H9 Broth. Acceptable specimen types are digested and decontaminated clinical specimens (except urine) and sterile body fluids (except blood). Can also be used for DST of mycobacteria, including SIRE and PZA testing. This item contains 100 x BD BBL MGIT Mycobacteria Growth Indicator Tube 7 mL | BD International |  | 200.85 | USD | 0 | General Cargo | 1.9 | 30 | 17 | 12 | 0.00612 | 245122 | 3822190090 | https://legacy.bd.com/ds/technicalCenter/inserts/L000180JAA(04).pdf
106027 | MGIT Supplement Kit, PANTA and OADC combined vials (245124) | 1 kit | 12 months | The BD BACTEC MGIT Growth Supplement contains Middlebrook OADC enrichment (Oleic acid, Bovine albumin, Dextrose, Catalase and Polyoxyethylene stearate [POES]). The BBL™ MGIT™ PANTA™ Antibiotic Mixture contains Polymyxin B, Amphotericin B, Nalidixic acid, Trimethoprim and Azlocillin. This item contains 6 vials of 15 mL BD BACTEC MGIT Growth Supplement and 6 vials of lyophilized BD BBL MGIT PANTA Antibiotic Mixture. Sufficient for 100 tests | BD International |  | 74.34 | USD | 0 | Cool cargo (+2-+8°C) | 0.54 | 23.2 | 16.4 | 7 | 0.0026634 | 245124 | 3822190090 | https://www.bd.com/resource.aspx?IDX=27408
106028 | BACTEC MGIT SIRE kit (245123) | 1 kit | 12 months | Rapid qualitative procedure for DST of M. tuberculosis using the BACTEC MGIT System (445870 or 441743) to streptomycin (STR), isoniazid (INH), rifampin (RIF) and ethambutol (EMB). This item contains 4 lyophilized drugs vials and 8 SIRE Supplements, for approximately 40 tests per drug per kit. For each box of SIRE kits, you need 2 boxes of MGIT Tubes (106026) | BD International |  | 79.97 | USD | 0 | Cool cargo (+2-+8°C) | 0.74 | 22 | 15 | 7 | 0.00231 | 245123 | 3822190090 | https://www.bd.com/resource.aspx?IDX=28198
106623 | MGIT OADC Enrichment, vials, 15 ml (245116) | 6 units | 12 months | Lyophilized MGIT OADC for DST on liquid culture. Provides substances essential for the rapid growth of mycobacteria in MGIT Tubes 4 ml (245113). OADC stands for oleic acid, albumin, dextrose and catalase. This item contains 6 x BD BBL MGIT OADC vial 15 mL, each vial sufficient for 25 MGIT tubes. | BD International |  | 58.28 | USD | 0 | Cool cargo (+2-+8°C) | 0.06 | 18 | 11 | 8 | 0.001584 | 245116 | 3821000000 | https://www.humeau.com/media/blfa_files/TC_MGIT-OADC_41624511602_EN_030820.pdf
106582 | BBL Middlebrook OADC Enrichment, 20 ml (211886) | 10 units | 8 months | Used to supplement culture media for the isolation and cultivation of mycobacteria for Middlebrook 7H10 Agar and Mycobacteria 7H11 Agar. This item contains 10 tubes (size A) of 20 ml each. This item has a very short shelf life. Need to order according to consumption. | BD International |  | 67.77 | USD | 0 | Cool cargo (+2-+8°C) | 0.5 | 14 | 6.6 | 15.6 | 0.001441 | 211886 | 3821000000 | https://www.bd.com/resource.aspx?IDX=27408
106033 | BACTEC MGIT PZA Kit (245128) | 1 kit | 12 months | The BD BACTEC MGIT 960 PZA Kit is used for susceptibility testing of Mycobacterium tuberculosis in culture to pyrazinamide (PZA) in a qualitative test lasting 4-17 days. This item contains 2 vials of lyophilized antimicrobic and 6 PZA Supplements, sufficient for 50 tests. PZA must be reconstituted with 2.5 ml of sterile/deionized water before addition to BBL MGIT tubes. | BD International |  | 95.13 | USD | 0 | Cool cargo (+2-+8°C) | 0.46 | 17.6 | 10.6 | 8.2 | 0.0015298 | 245128 | 3822190090 | https://www.bd.com/resource.aspx?IDX=27408
106036 | BACTEC MGIT PZA Tubes (245115) | 25 units | 12 months | The BACTEC MGIT 960 PZA Medium is a tube containing modified Middlebrook 7H9 Broth with a reduced pH of 5.9 and is to be used for susceptibility testing of Mycobacterium tuberculosis. This item contains 25 x BD BACTEC MGIT 960 PZA Medium tube | BD International |  | 56.23 | USD | 0 | General Cargo | 0.42 | 12 | 11 | 13 | 0.001716 | 245115 | 3822190090 | https://www.bd.com/resource.aspx?IDX=27408
106037 | BBL MycoPrep Kit, 10x150 ml (240863) | 1 kit | 7 months | Used for the digestion and decontamination of clinical specimens suspected to contain mycobacteria, especially M. tuberculosis for minimizing contamination in solid and liquid media. Contains reagents necessary for the preparation of respiratory and non respiratory (blood excluded) specimens for identification of mycobacteria. Includes bottles of NaOH-citrate solution (containing a plastic ampoule with NALC) and powdered phosphate buffer. After activation of the NALC ampoule, NALC-NaOH mixture is stable for 24 hours. This item includes: 10 bottles of 150 ml of MycoPrep and 10 packages of powdered phosphate buffer (pH 6.8). This item has a very short shelf life. Need to order according to consumption. | BD International |  | 85.58 | USD | 0 | Dangerous cargo UN 1824, Class 8, PG III | 2.35 | 21 | 21 | 18 | 0.007938 | 240863 | 3822190090 | https://www.bd.com/en-uk/offerings/capabilities/microbiology-solutions/clinical-microbiology/mycobacteria-testing/bd-bbl-mycoprep-kit
106021 | BACTEC MGIT AST Carrier Set 2-tube (445946) | 3 units | N/A | BACTEC MGIT AST Carrier Set (2-tube). This item contains 3 carriers. | BD International |  | 37.8 | USD | 0 | General Cargo | 0.0314 | 30.4 | 18.4 | 4.5 | 0.0025171 | 445946 | 3923900000 | https://www.bd.com/resource.aspx?IDX=27408
106020 | BACTEC MGIT AST Carrier Set 3-tube (445945) | 3 units | N/A | BACTEC MGIT AST Carrier Set (3-tube). This item contains 3 carriers. | BD International |  | 37.8 | USD | 0 | General Cargo | 0.0187 | 16.7 | 10.1 | 3 | 0.00050601 | 445945 | 3923900000 | https://www.bd.com/resource.aspx?IDX=27408
106019 | BACTEC MGIT AST Carrier Set 4-tube (445944) | 3 units | N/A | BACTEC MGIT AST Carrier Set (4-tube). This item contains 3 carriers. | BD International |  | 37.8 | USD | 0 | General Cargo | 0.0261 | 16.7 | 10.1 | 3 | 0.00050601 | 445944 | 3923900000 | https://www.bd.com/resource.aspx?IDX=27408
106018 | BACTEC MGIT AST Carrier Set 5-tube (445943) | 3 units | N/A | BACTEC MGIT AST Carrier Set (5-tube). This item contains 3 carriers. | BD International |  | 37.8 | USD | 0 | General Cargo | 0.0375 | 14.6 | 3.8 | 4.4 | 0.00024411 | 445943 | 3923900000 | https://www.bd.com/resource.aspx?IDX=27408
106022 | BACTEC MGIT AST Carrier Set 8-tube (445993) | 3 units | N/A | BACTEC MGIT AST Carrier Set (8-tube). This item contains 3 carriers. | BD International |  | 37.8 | USD | 0 | General Cargo | 0.0575 | 25.4 | 15.2 | 3.1 | 0.0011968 | 445993 | 3923900000 | https://www.bd.com/resource.aspx?IDX=27408
106616 | BACTEC MGIT QC Temperature Tube (445872) | 1 unit | N/A | BACTEC MGIT Temperature QC Tube. This item contains 1 QC Temperature Tube | BD International |  | 122.49 | USD | 0 | General Cargo | 0.1545 | 20 | 13.8 | 8.6 | 0.0023736 | 445872 | 9025118090 | https://www.bd.com/resource.aspx?IDX=27408
106617 | BACTEC MGIT spare bar code set carrier (445959) | 1 unit | N/A | BACTEC MGIT AST Carrier Set Spare Barcodes. This item contains 1 carrier Set, spare barcode | BD International |  | 59.95 | USD | 0 | General Cargo | 0.0111 | 19.6 | 13.3 | 0.5 | 0.00013 | 445959 | 4911990000 | https://www.bd.com/resource.aspx?IDX=27408
106618 | BACTEC MGIT Plug for Bad Station (445873) | 10 units | N/A | BACTEC MGIT Plug for Bad Station. This item contains 10 plug for Bad Station | BD International |  | 659.2 | USD | 0 | General Cargo | 0.016 | 1.5 | 12.5 | 19 | 0.00035625 | 445873 | 8536909599 | https://www.bd.com/resource.aspx?IDX=27408
106473 | BACTEC Air Filters, square (444374) | 2 units | N/A | BACTEC MGIT Air Filters, square. This item contains 2 air filters | BD International |  | 59.77 | USD | 0 | General Cargo | 0.13 | 30.4 | 18.4 | 2.1 | 0.0011747 | 444374 | 9027900000 | http://www.solone.in/docs/blood-culture.pdf
106474 | BACTEC Air Filters, rectangular (445888) | 1 unit | N/A | BACTEC MGIT Air Filters, rectangular. This item contains 2 air filters | BD International |  | 72.54 | USD | 0 | General Cargo | 0.1 | 30.4 | 18.4 | 5 | 0.0027968 | 445888 | 8421392590 | https://www.bd.com/resource.aspx?IDX=27408
106452 | BACTEC MGIT 960 AST transport rack - 445942 | 1 unit | N/A | BACTEC MGIT AST transport rack. This item contains 1 AST transport rack | BD International |  | 169.07 | USD | 0 | General Cargo | 0.945 | 46.8 | 22.6 | 12.8 | 0.01354 | 445942 | 9027900000 | https://www.bd.com/resource.aspx?IDX=27408
106468 | Mc Farland standard set | 1 set | 10 months | Turbidity Standard Set McFarland Equivalence Turbidity Standard Tube Format: 1 each of 0.5, 1.0, 2.0, 3.0, 4.0.Tube size is 15 x 103mm and fits in most spectrophotometers. Packaged in plastic case. · Polystyrene technology enables longer shelf life. This item includes: a visual comparison card. | TTM |  | 139.25 | USD | 0 | General Cargo | 0.2 | 15 | 10 | 5 | 0.00075 | Remel Inc. R20421 | 38221900 | https://www.thermofisher.com/order/catalog/product/R20421#/R20421
106181 | Sterile dispenser tips 10ml | 100 units | 16 months | Sterile dispenser-tips Universal 10ml sterile - DNA- /RNase- /ATP- and endotoxin free. · 100 % contamination prevention · Cylinder made of PP, individually wrapped. | TTM |  | 196.65 | USD | 0 | General Cargo | 1.2 | 30 | 25 | 25 | 0.0188 | Eppendorf 0030089677 | 39269097 | https://online-shop.eppendorf.com/OC-en/Manual-Liquid-Handling-44563/Pipette-Tips-44569/Combitips-advanced-PF-18250.html
106209 | Sterile serological pipette 5 ml | 200 units | 16 months | Sterile serological pipette - 5 ml, its paper/plastic wrap allows easy opening by either pop-through or peel-apart techniques. It is also ideal for sterile tissue culture applications, especially while wearing latex gloves, as the wrap reduces static cling.Color-coding on paper allows easy size identification. Paper/plastic wrap allows full view; Unique color-coded magnifying stripe for easier meniscus viewing. Sterile, individually wrapped and nonpyrogenic and DNase-/RNase-free Accuracy within ± 2% at full volume; Lot numbers printed on each pipet wrapper Calibrated to deliver (TD-EX 20ºC) their contents with blowout. | TTM |  | 60.95 | USD | 0 | General Cargo | 2.3 | 40 | 21 | 19 | 0.01596 | Corning Incorporated Corning 4487 | 39269097 | https://ecatalog.corning.com/life-sciences/b2c/CH/en/Liquid-Handling/Pipets/Serological-Pipets/Stripette%E2%84%A2-Serological-Pipets/p/4487
106550 | Sterile serological pipette 10 ml | 200 units | 16 months | Sterile serological pipette - 10 ml, its paper/plastic wrap allows easy opening by either pop-through or peel-apart techniques. It is also ideal for sterile tissue culture applications, especially while wearing latex gloves, as the wrap reduces static cling.Color-coding on paper allows easy size identification. Paper/plastic wrap allows full view; Unique color-coded magnifying stripe for easier meniscus viewing. Sterile, individually wrapped and nonpyrogenic and DNase-/RNase-free Accuracy within ± 2% at full volume; Lot numbers printed on each pipet wrapper Calibrated to deliver (TD-EX 20ºC) their contents with blowout. | TTM |  | 62.1 | USD | 0 | General Cargo | 9 | 35 | 25 | 20 | 0.0175 | Corning Incorporated Corning 4488 | 39269097 | https://www.corning.com/catalog/cls/documents/product-information-sheets/CLS-LH-367_DL.pdf
106241 | Measuring pipette for BSC, 1ml | 1 unit | N/A | Graduated pipettes 1 ml, AR glass, short type, especially for tissue culture and working on cleanroom benches, with durable amber stain graduation, main point ring graduation, zero at the bottom, standard cotton plug end (o.d. 8 mm), EX (TD) calibrated. Double scale graduated 0,1 ml · Marks and inscriptions in durable amber stain; Individually wrapped. | TTM |  | 5 | USD | 0 | General Cargo | 0.2 | 26 | 8 | 7 | 0.001456 | Hirschmann Laborgeräte GmbH 1170135 | 70179000 | https://www.hirschmannlab.com/en/artikelgruppe/11702?parent=%7B031F8BD4-1D4C-44DA-8BB8-3D1379F43010%7D
106715 | Sterile syringes 20 ml | 80 units | 16 months | Sterile syringes 20 ml, sterile, single-use, transparent barrel for good visualization of syringe content. · Smooth plunger movement reduced sliding force; leak-tight. Accurate graduation suited to user’s needs. · Retaining ring to prevent accidental plunger rod withdrawal · Design of finger grip flanges ensures stability and comfort during injection · Ergonomically designed plunger for single-handed operation | TTM |  | 7.15 | USD | 0 | General Cargo | 0.8 | 43 | 17 | 11 | 0.008041 | Becton Dickinson Discardit II 300296 | 90183110 | https://www.bd.com/en-eu/offerings/capabilities/syringes-and-needles/injection-syringes/bd-discardit-ii
106822 | Glass petri dishes | 20 units | N/A | Glass petri dish absolutely plain, heat resisting glass. Made from DURAN borosilicate glass 3.3, in a special production process, microscopically flat and free from bubbles and streaks. Dishes enable an even distribution of culture media and have a distortion-free transparency. Autoclavable. · Diameter: 100 mm; Base height: 15 mm; Lid height: 17mm; This item includes: 2 bags with 10 units each for a total of 20 units. Each one with a lid. | TTM |  | 279.9 | USD | 0 | General Cargo | 4.3 | 32 | 50 | 40 | 0.064 | DWK Lifesciences 2x 217554607 | 70179000 | https://www.fishersci.be/shop/products/duroplan-borosilicate-petri-dishes-2/11748001
106823 | Plastic petri dishes | 20 units | 30 months | Plastic petri dish, sterile and made of PS with vents. High chemical resistance. Max. volume: 80 ml, working volume: 15 - 16 ml. Diameter: 94 mm; Base height: 16 mm; This item includes 20 units, each one with a lid. | TTM |  | 1.95 | USD | 0 | General Cargo | 0.3 | 35 | 11 | 11 | 0.00424 | Greiner Bio-One GmbH 633181 | 39269097 | https://shop.gbo.com/en/switzerland/products/bioscience/microbiology-bacteriology/dishes/petri-dishes/633181.html
1.5.7 Solid and liquid culture DST: FLD pure substances
106308 | Isoniazid 5 g - for laboratory use | 1 unit | 16 months | Isoniazid powder presentation; needs to be prepared manually. CAS No 54-85-3 · Grade: analytical standard; assay: ≥99% (TLC) Subject to availability from our supplier. Product lead-time may be longer than usual. Please contact your GDF Country Supply Officer for further information | TTM |  | 40.5 | USD | 0 | General Cargo | 0.3 | 6 | 6 | 12 | 0.000432 | Merck KGaA I3377-5G | 32041300 | https://www.sigmaaldrich.com/catalog/product/sial/i3377?lang=en&region=GB
106309 | Rifampicin 1 g - for laboratory use | 1 unit | 16 months | Rifampicin powder presentation; needs to be prepared manually. CAS No 13292-46-1. Rifampicin; About 1000 test done with 1 gr. · Assay: ≥97% (HPLC) Subject to availability from our supplier. Product lead-time may be longer than usual. Please contact your GDF Country Supply Officer for further information | TTM |  | 131.6 | USD | 0 | Cool cargo (+2-+8°C) | 0.25 | 6 | 6 | 12 | 0.000432 | Merck KGaA R3501-1G | 29419000 | https://www.sigmaaldrich.com/catalog/product/sigma/r3501?lang=en&region=GB
106310 | Ethambutol 25 g - for laboratory use | 1 unit | 16 months | Ethambutol powder presentation; needs to be prepared manually. CAS No 1070-11-7 · Purity (TLC) > 99 % Subject to availability from our supplier. Product lead-time may be longer than usual. Please contact your GDF Country Supply Officer for further information | TTM |  | 140.1 | USD | 0 | General Cargo | 0.25 | 6 | 6 | 12 | 0.000432 | Merck KGaA E4630-25G | 29221900 | https://www.sigmaaldrich.com/catalog/product/sigma/e4630?lang=en&region=GB
1.5.8 Solid and liquid culture DST: SLD pure substances
106584 | BACTEC MGIT Kanamycin lyophilized drug for lab use (215348) | 6 units | 8 months | BACTEC MGIT Kanamycin lyophilized drug vials for laboratory use to perform liquid culture. To be used on the BACTEC MGIT 960 and 320 System. This item contains 6 vials with 830 µg/vial to run 240 tests in total. This item has a very short shelf life. Need to order according to consumption | BD International |  | 94.5 | USD | 0 | Cool cargo (+2-+8°C) | 0.16 | 9.8 | 6.6 | 6.6 | 0.00043 | 215348 | 3822190090 | Not available
106586 | BACTEC MGIT Amikacin lyophilized drug for lab use (215350) | 6 units | 8 months | BACTEC MGIT Amikacin lyophilized drug vials for laboratory use to perform liquid culture. To be used on the BACTEC MGIT 960 and 320 System. This item contains 6 vials with 332 µg/vial to run 240 tests in total. This item has a very short shelf life. Need to order according to consumption | BD International |  | 94.5 | USD | 0 | Cool cargo (+2-+8°C) | 0.16 | 9.8 | 6.6 | 6.6 | 0.00043 | 215350 | 3822190090 | Not available
106587 | BACTEC MGIT Capreomycin lyophilized drug for lab use (215351) | 6 units | 8 months | BACTEC MGIT Capreomycin lyophilized drug vials for laboratory use to perform liquid culture. To be used on the BACTEC MGIT 960 and 320 System. This item contains 6 vials with 830 µg/vial to run 240 tests in total. This item has a very short shelf life. Need to order according to consumption | BD International |  | 94.5 | USD | 0 | Cool cargo (+2-+8°C) | 0.16 | 9.8 | 6.6 | 6.6 | 0.00043 | 215351 | 3822190090 | Not available
106588 | BACTEC MGIT Ofloxacin lyophilized drug for lab use (215352) | 6 units | 8 months | BACTEC MGIT Ofloxacin lyophilized drug vials for laboratory use to perform liquid culture. To be used on the BACTEC MGIT 960 and 320 System. This item contains 6 vials with 664 µg/vial to run 240 tests in total. This item has a very short shelf life. Need to order according to consumption | BD International |  | 94.5 | USD | 0 | Cool cargo (+2-+8°C) | 0.16 | 9.8 | 6.6 | 6.6 | 0.00043 | 215352 | 3822190090 | Not available
106975 | BACTEC MGIT Moxifloxacin lyophilized drug for lab use (215404) | 6 units | 8 months | BACTEC MGIT Moxifloxacin lyophilized drug vials for laboratory use to perform liquid culture. To be used on the BACTEC MGIT 960 and 320 System. Approximate formula per vial: 249µg (active); Flcoll (Inactive). This item contains 6 vials of 249µg to run 180 tests in total. This item has a very short shelf life. Need to order according to consumption | BD International |  | 94.5 | USD | 0 | Cool cargo (+2-+8°C) | 0.115 | 6.5 | 9.7 | 6.5 | 0.00043 | 215404 | 3822190090 | Not available
106976 | BACTEC MGIT Bedaquiline lyophilized drug for lab use (215449) | 2 units | 12 months | BACTEC MGIT Bedaquiline lyophilized drug vials for laboratory use to perform liquid culture. To be used on the BACTEC MGIT 960 and 320 System. Approximate formula per vial: 170µg (active); Flcoll (Inactive). Contents of vial soluble in 2mL dimethyl sulfoxide (DMSO). This item contains 2 vials with 170 µg/vial to run 40 tests in total. An equal quantity of 8x BACTEC MGIT Bedaquiline lyophilized drug for lab use should be ordered for 1x of BACTEC MGIT Dimethyl sulfoxide (DMSO) Diluent | BD International |  | 53.6 | USD | 0 | Cool cargo (+2-+8°C) | 0.03 | 3.1 | 6.5 | 6.4 | 0.00012 | 215449 | 3822190090 | Not available
106977 | BACTEC MGIT Dimethyl sulfoxide (DMSO) Diluent (246017) | 2 units | 20 months | BACTEC MGIT Dimethyl sulfoxide (DMSO) Diluent for laboratory use to perform liquid culture. To be used on the BACTEC MGIT 960 and 320 System. Contents: DimethylSulfoxide (DMSO). Purity: ≥ 99.0%, 16 mL. This item contains 2 diluents of 16 ml each 1 box to run 320 tests in total. | BD International |  | 29.6 | USD | 0 | General Cargo | 0.0992 | 6.5 | 3.5 | 6.5 | 0.00014 | 246017 | 2930909899 | Not available
106560 | Levofloxacin 1 g - for laboratory use | 1 unit | 16 months | Levofloxacin for laboratory use, powder presentation. CAS No 100986-85-4. Assay (HPLC) 98.0 - 102.0 %. · 1 g is needed for 100 tests to be done in solid or liquid culture. | TTM |  | 43.75 | USD | 0 | General Cargo | 0.3 | 6 | 6 | 12 | 0.000432 | Merck KGaA 28266 | 29419000 | https://www.sigmaaldrich.com/catalog/substance/levofloxacin3613710098685411?lang=en&region=GB
106653 | Linezolid 25 mg - for laboratory use | 1 unit | 16 months | Temporarily not available | TTM |  | 516.95 | USD | 0 | General Cargo | 0.4 | 6 | 6 | 12 | 0.000432 | Merck KGaA 0014 | 29419000 | https://www.sigmaaldrich.com/catalog/product/sigma/pz0014?lang=en&region=GB
106654 | Clofazimine 1 g - for laboratory use | 1 unit | 16 months | Clofazimine for laboratory use, powder presentation; CAS No 2030-63-9. · 1 g is needed for 100 tests to be done in solid or liquid culture. | TTM |  | 111.3 | USD | 0 | General cargo | 0.3 | 6 | 6 | 12 | 0.000432 | Merck KGaA C8895-1G | 29419000 | https://www.sigmaaldrich.com/catalog/product/sigma/c8895?lang=en&region=GB
106318 | Amikacin 5 g - for laboratory use | 1 unit | 16 months | Amikacin for laboratory use, powder presentation; CAS No.1257517-67-1 · 1 g is needed for 100 tests to be done in solid or liquid culture. | TTM |  | 651.8 | USD | 0 | Cool cargo (+2-+8°C) | 0.4 | 6 | 6 | 12 | 0.000432 | Merck KGaA A3650-5G | 29419000 | https://www.sigmaaldrich.com/catalog/product/sigma/a3650?lang=en&region=GB
106313 | Ofloxacin 10 g - for laboratory use | 1 unit | 16 months | Ofloxacin for laboratory use, powder presentation; CAS No 82419-36-1. Assay ≥99% (HPLC). · 1 g is needed for 100 tests to be done in solid or liquid culture. | TTM |  | 295.15 | USD | 0 | Cool cargo (+2-+8°C) | 0.4 | 6 | 6 | 12 | 0.000432 | Merck KGaA O8757-10G | 29419000 | https://www.sigmaaldrich.com/catalog/product/sigma/o8757?lang=en&region=GB
106316 | Ethionamide 5 g - for laboratory use | 1 unit | 16 months | Ethionamide for laboratory use, powder presentation; CAS No 536-33-4 Assay ≥98.0%. · 1 g is needed for 100 tests to be done in solid or liquid culture. | TTM |  | 135.2 | USD | 0 | Cool cargo (+2-+8°C) | 0.4 | 6 | 6 | 12 | 0.000432 | Merck KGaA E6005-5G | 29333999 | https://www.sigmaaldrich.com/catalog/product/sigma/e6005?lang=en&region=GB
106315 | Protionamide 2 g - for laboratory use | 1 unit | 16 months | Protionamide for laboratory use, powder presentation; CAS No 14222-60-7. Assay ≥99% (HPLC). · 1 g is needed for 100 tests to be done in solid or liquid culture. | TTM |  | 165.9 | USD | 0 | Cool cargo (+2-+8°C) | 0.4 | 6 | 6 | 12 | 0.000432 | Sigma Aldrich SMB00387 | 29333999 | https://www.sigmaaldrich.com/catalog/product/sigma/smb00387?lang=en&region=GB
106317 | Kanamycin 5 g - for laboratory use | 1 unit | 16 months | Kanamycin for laboratory use, powder presentation; CAS No 25389-94-0. Potency≥750 μg/ mg (dry basis) · 1 g is needed for 100 tests to be done in solid or liquid culture. | TTM |  | 171.4 | USD | 0 | Cool cargo (+2-+8°C) | 0.25 | 6 | 6 | 12 | 0.000432 | Sigma Aldrich K1377-5G | 29419000 | https://www.sigmaaldrich.com/catalog/product/sigma/k1377?lang=en&region=GB
106319 | Capreomycin 1 g - for laboratory use | 1 unit | 16 months | Capreomycin for laboratory use, powder presentation; CAS No 1405-37-4. · 1 g is needed for 100 tests to be done in solid or liquid culture. | TTM |  | 274.6 | USD | 0 | Cool cargo (+2-+8°C) | 0.3 | 6 | 6 | 12 | 0.000432 | Merck KGaA C4142-1G | 29419000 | https://www.sigmaaldrich.com/catalog/product/sigma/c4142?lang=en&region=GB
1.6 MTB IDENTIFICATION
1.6.1 MTB Identification Rapid test
106611 | Capilia TB-Neo assay, MPB64 Ag, 100 tests (CATB0870) | 100 tests | 18 months | Rapid test, immunochromatography assay, for identification of M. Tuberculosis complex: Detecting MPB64 protein (using a monoclonal antibody that recognizes the MPB64 antigen), specifically produced by M. Tuberculosis complex. Rapid test used to qualitatively detect mycobacterium antigens in a culture grown in a medium for acid-fast bacteria (AFB) inoculated with properly pre-treated clinical specimens, such as body fluid, tissue, and bronchoalveolar lavage fluid collected from patients with suspected tuberculosis. Fast test results with one-step operation: Reading time is 15 min. No special equipment required. Both liquid medium and solid medium can be used as sample. For solid medium, Capilia TB-Neo assay, MPB64 Ag, extraction buffer, 20 ml (procured separately)) is required. | TAUNS Laboratories |  | 100 | USD | 0 | General Cargo | 0.78 | 17.4 | 13 | 13.4 | 0.003 | CATB0870 | 38221900 | https://www.tauns.co.jp/en/products/capilia-tb-neo-e/
106953 | Capilia TB-Neo assay, MPB64 Ag, 10 tests (CATB0871) | 10 tests | 18 months | Rapid test, immunochromatography assay, for identification of M. Tuberculosis complex: Detecting MPB64 protein (using a monoclonal antibody that recognizes the MPB64 antigen), specifically produced by M. Tuberculosis complex. Rapid test used to qualitatively detect mycobacterium antigens in a culture grown in a medium for acid-fast bacteria (AFB) inoculated with properly pre-treated clinical specimens, such as body fluid, tissue, and bronchoalveolar lavage fluid collected from patients with suspected tuberculosis. Fast test results with one-step operation: Reading time is 15 min. No special equipment required. Both liquid medium and solid medium can be used as sample. For solid medium, Capilia TB-Neo assay, MPB64 Ag, extraction buffer, 20 ml (procured separately)) is required. | TAUNS Laboratories |  | 12 | USD | 0 | General Cargo | 0.1 | 12.5 | 7 | 4.5 | 0.00039375 | CATB0871 | 38221900 | https://www.tauns.co.jp/en/products/capilia-tb-neo-e/
106612 | Capilia TB-Neo assay, MPB64 Ag, extraction buffer, 20 ml | 1 unit | 20 months | This product is an extraction buffer using Capilia TB-Neo to prepare a sample that is employed for the identification test of M. Tuberculosis when using a solid medium for acid-fast bacteria (AFB). Extraction buffer composed of: Phosphate buffer, detergent, sodium azide (0.09%). Extraction buffer is required for solid medium. This item contains: a bottle of 20 ml. Product required when procuring the Capilia TB-Neo assay, 100 tests or 10 tests. | TAUNS Laboratories |  | 8 | USD | 0 | General Cargo | 0.08 | 8 | 4 | 10.3 | 0.0003 | CATB0877 | 38221900 | https://www.tauns.co.jp/en/products/capilia-tb-neo-e/
2. DETECTION OF LATENT TB
2.1 IGRA
2.1.1 IGRA Equipment
106841 | Incubator 74 lt | 1 unit | N/A | Laboratory incubator 74 lt, PID microprocessor control with integrated auto-diagnostic system. Calibrated 37 °C. Temperature range min. 10 above ambient temperature up to +80°C. Setting accuracy 0.1°C; Housing Stainless steel; · Digital backwards counter with target time setting, adjustable from 1 minute to 99 days This item includes: 1 shelf or grid and 1 additional perforated shelf. | AMEX |  | 1463.9 | USD | 0 | General Cargo | 85 | 67 | 73 | 113 | 0.55 | Memmert IN 75 | 84198990 | https://www.memmert.com/products/incubators/incubator/IN75/pdf/
106842 | Shelf for incubator 74 lt | 1 unit | N/A | Perforated stainless steel shelf, made of stainless steel. Shelf support 15kg load. | AMEX |  | 55.5 | USD | 0 | General Cargo | 2 |  |  |  | 0.012 | Memmert IN 75 shelf B03916 | 84198990 | https://www.memmert.com/products/incubators/incubator/IN75/pdf/
106845 | Laboratory compact centrifuge | 1 unit | N/A | Laboratory compact centrifuge with RPM: 500 - 16,000 min -1; Adjustable in increments of 10 and max. RCF: 24,900. Max. capacity: 4 x 200 ml / 6 x 94 ml. Centrifuge with a noise level of 48 dB(A) with angle rotor 1611 and impulse button for short centrifugation. Easy operation with keypad and control knob. · Centrifuge with 9 program memories and 9 individual acceleration and 10 deceleration stages · Adjustable timer range not smaller than 0 to 59 minutes; This item includes: 4-pace rotor, 4 rotor buckets and 4 bucket lids. | AMEX |  | 2972.2 | USD | 0 | General Cargo | 45 | 45 | 58 | 40 | 0.1044 | Hettich 1401 Universal 320 (4-pace rotor: Hettich 1624, 4 rotor buckets: Hettich 1481 + 1329 and 4 bucket lids: Hettich 1492) | 84211999 | https://www.hettichlab.com/downloadcenter/Products/Datasheets/UNIVERSAL320-320R_EN.pdf
106857 | Sterile D-PBS solution, 500 ml | 1 unit | 16 months | Phosphate Buffered Saline (PBS), 10X, Sterile, Ultra Pure Grade. 1X Solution Composition:137 mM NaCl, 2.7 mM KCl, 9.5 mM Phosphate buffer. · Dilute 1:10 in sterile water for making 1X solution; 10X solution is supplied 0.2 um filter sterilized · pH at 25°C (1X PBS): 6.6 – 7.2; DNase, RNase,Protease none detected. Bioburden Pass (USP) · Suitable for cell culture; Does not contain magnesium or calcium | TTM |  | 92.75 | USD | 0 | General Cargo | 0.6 | 20 | 15 | 15 | 0.0045 | VWR Chemicals K813 | 90189084 | https://us.vwr.com/store/product/7676424/phosphate-buffered-saline-pbs-10x-sterile-ultra-pure-grade
2.1.2 IGRA Testing
106936 | QIAreach e-Hub | 1 unit | N/A | Temporarily not available | QIAGEN GmbH |  | 1200 | USD | 0 | General Cargo | 1 | 5.1 | 30.4 | 11.2 | 0.005 | 9002969 | 8517620000 | https://www.qiagen.com/ca/applications/tb-management/products/qiareach-quantiferon-tb-test
106937 | QIAreach QuantiFERON-TB, 50 test | 1 unit | 10 months | Temporarily not available | QIAGEN GmbH |  | 600 | USD | 0 | Cool cargo (+2-+8°C) | 1.5 | 29.1 | 19.7 | 17.1 | 0.01 | 622724 | 3002150000 | https://www.qiagen.com/ca/applications/tb-management/products/qiareach-quantiferon-tb-test
106671 | QuantiFERON-TB Gold Plus ELISA, 2x96 tests (622120) | 2 boxes | 24 months | The QuantiFERON-TB Gold Plus (QFT-Plus) assay is an in vitro diagnostic test using a peptide cocktail simulating ESAT-6 and CFP-10 proteins to stimulate cells in heparinized whole blood. Detection of interferon-γ (IFN-γ) b enzyme-linked immunosorbent assay (ELISA) is used to identify in vitro responses to those peptide antigens that are associated with Mycobacterium tuberculosis infection. QFT-Plus is an indirect test for M. tuberculosis infection (including disease) and is intended for use in conjunction with risk assessment, radiography, and other medical and diagnostic evaluations. This item includes: Microplate Strips 2 x (12 x 8 wells); IFN- Standard, lyophilized; Green Diluent; Conjugate 100x Concentrate, lyophilized; Wash Buffer 20x Concentrate; Enzyme Substrate Solution; Enzyme Stopping Solution; QFT-Plus ELISA Package Insert. | QIAGEN GmbH |  | 279.4 | USD | 0 | Dangerous and Cool cargo (+2 -+8°C) UN 3264 Class: 8, Packing group: III - Corrosive liquid, acidic, inorganic, n.o.s. (sulfuric acid) | 0.6 | 19.8 | 11 | 13.5 | 0.003 | 622120 | 30021500 | https://www.qiagen.com/ch/products/diagnostics-and-clinical-research/tb-management/quantiferon-tb-gold-plus/quantiferon-tb-gold-plus/#orderinginformation
106672 | QuantiFERON-TB Gold Plus ELISA, 20x96 tests (Ref lab) (622822) | 20 boxes | 24 months | The QuantiFERON-TB Gold Plus (QFT-Plus) assay is an in vitro diagnostic test using a peptide cocktail simulating ESAT-6 and CFP-10 proteins to stimulate cells in heparinized whole blood. Detection of interferon by enzyme-linked immunosorbent assay (ELISA) is used to identify in vitro responses to those peptide antigens that are associated with Mycobacterium tuberculosis infection. QFT-Plus is an indirect test for M. tuberculosis infection (including disease) and is intended for use in conjunction with risk assessment, radiography, and other medical and diagnostic evaluations. This item includes: Microplate Strips 20 x (12 x 8 wells); IFN Standard, lyophilized; Green Diluent; Conjugate 100x Concentrate, lyophilized; Wash Buffer 20x Concentrate; Enzyme Substrate Solution; Enzyme Stopping Solution; QFT-Plus ELISA Package Insert. | QIAGEN GmbH |  | 2798.4 | USD | 0 | Dangerous and Cool cargo (+2 -+8°C) UN 3264 Class: 8, Packing group: III - Corrosive liquid, acidic, inorganic, n.o.s. (sulfuric acid) | 0.6 | 19.8 | 11 | 13.5 | 0.003 | 622822 | 30021500 | https://www.qiagen.com/ch/products/diagnostics-and-clinical-research/tb-management/quantiferon-tb-gold-plus/quantiferon-tb-gold-plus/#orderinginformation
106673 | QFT-Plus Blood Collection 200 Tubes (622526) | 200 units | 10 months | The blood collection volume of a tube is affected by altitude. Standard QFT-Plus Blood Collection Tubes should be used between sea level and 810 m (2650 ft). This item includes 200 Tubes (50 sets of 4 tubes): 1.QuantiFERON Nil Tube (grey cap, white ring): 50 tubes; 2.QuantiFERON TB1 Tube (green cap, white ring): 50 tubes; 3.QuantiFERON TB2 Tube (yellow cap, white ring): 50 tubes; 4.QuantiFERON Mitogen Tube (purple cap, white ring): 50 tubes. | QIAGEN GmbH |  | 477 | USD | 0 | CRT cargo (+15-+25°C) | 1.6 | 19.5 | 19.5 | 17 | 0.006 | 622526 | 30021500 | https://www.qiagen.com/ch/products/diagnostics-and-clinical-research/tb-management/quantiferon-tb-gold-plus/quantiferon-tb-gold-plus/#orderinginformation
106674 | QFT-Plus Blood Collection 200 Tubes - High Altitude (623526) | 200 units | 10 months | The blood collection volume of a tube is affected by altitude. High Altitude (HA) tubes should be used at altitudes between 1020 m (3350 ft) and 1875 m (6150 ft). This item includes 200 Tubes (50 sets of 4 tubes): 1. QuantiFERON Nil Tube (grey cap, Yellow ring): 50 tubes; 2. QuantiFERON TB1 Tube (green cap, Yellow ring): 50 tubes; 3. QuantiFERON TB2 Tube (yellow cap, Yellow ring): 50 tubes; 4. QuantiFERON Mitogen Tube (purple cap, Yellow ring): 50 tubes. | QIAGEN GmbH |  | 477 | USD | 0 | CRT cargo (+15-+25°C) | 1.6 | 19.5 | 19.5 | 17 | 0.006 | 623526 | 30021500 | https://www.qiagen.com/ch/products/diagnostics-and-clinical-research/tb-management/quantiferon-tb-gold-plus/quantiferon-tb-gold-plus/#orderinginformation
106675 | QFT-Plus Blood Collect 40 Tubes Single Patient (622222) | 40 units | 10 months | For use between sea-level and 810 m (2650 ft). The blood collection tubes supplied are for use only with the QFT-Plus assay. Antigens have been dried onto the inner wall of the blood collection tubes, so it is essential that the contents of the tubes be thoroughly mixed with the blood. This item contains 40 Tubes Single Patient Pack (10 sets of 4 tubes): 1. QuantiFERON Nil Tube (grey cap, white ring): 10 tubes; 2. QuantiFERON TB1 Tube (green cap, white ring): 10 tubes; 3. QuantiFERON TB2 Tube (yellow cap, white ring): 10 tubes 4. QuantiFERON Mitogen Tube (purple cap, white ring): 10 tubes and 5 Cannulas and brackets for each patient (10) | QIAGEN GmbH |  | 95.4 | USD | 0 | CRT cargo (+15-+25°C) | 0.58 | 9.5 | 33 | 10 | 0.003 | 622222 | 30021500 | https://www.qiagen.com/ch/products/diagnostics-and-clinical-research/tb-management/quantiferon-tb-gold-plus/quantiferon-tb-gold-plus/#orderinginformation
106676 | QFT-Plus Blood Collect 40 Tubes Single Patient High Alt (623222) | 40 units | 10 months | For use between 1020 m (3350 ft) and 1875 m (6150 ft). The blood collection tubes supplied are for use only with the QFT-Plus assay. Antigens have been dried onto the inner wall of the blood collection tubes, so it is essential that the contents of the tubes be thoroughly mixed with the blood. This item includes 40 Tubes Single Patient Pack (10 sets of 4 tubes): 1. QuantiFERON Nil Tube (grey cap, Yellow ring): 10 tubes; 2. QuantiFERON TB1 Tube (green cap, Yellow ring): 10 tubes; 3. QuantiFERON TB2 Tube (yellow cap, Yellow ring): 10 tubes; 4. QuantiFERON Mitogen Tube (purple cap, Yellow ring): 10 tube; 5. Cannulas and bracket for each patient (10) | QIAGEN GmbH |  | 95.4 | USD | 0 | CRT cargo (+15-+25°C) | 0.58 | 9.5 | 33 | 10 | 0.003 | 623222 | 30021500 | https://www.qiagen.com/ch/products/diagnostics-and-clinical-research/tb-management/quantiferon-tb-gold-plus/quantiferon-tb-gold-plus/#orderinginformation
106665 | T-SPOT.TB Test, 24 tests | 1 box | 14 months | The T-SPOT.TB test is an in vitro diagnostic test for the detection of effector T cells that respond to stimulation by Mycobacterium tuberculosis antigens and is intended for use as an aid in the diagnosis of tuberculosis infection. The T-SPOT.TB test is a simplified enzyme-linked immunospot (ELISPOT) method which enumerates individual TB-specific activated effector T cells. Process peripheral blood mononuclear cells (PBMCs) within 8 hours, or if T-Cell Xtend is used, within 30 hours This item includes: a box for 24 tests. | Oxford Immunotec |  | 168 | USD | 0 | Cool cargo (+2-+8°C) | 0.3 | 15 | 12 | 9 | 0.00162 | TB.300 | 38220000 | http://www.oxforddiagnosticlaboratories.eu/products-services/t-spot-tb-test/
106667 | T-Cell Select, 144 tests | 1 kit | 12 months | The T-Cell Select kit is intended for the isolation of mononuclear immune cells from whole blood, using positive selection via a magnetic bead-based cell separation system, for use in cell-mediated immune assays such as T-SPOT.TB. This item includes: 1 kit that will run 144 blood samples (with bead reagent, buffer concentrate, Antibody 1 and Antibody 2) | Oxford Immunotec |  | 693 | USD | 0 | Cool cargo (+2-+8°C) | 0.245 | 15 | 12 | 9 | 0.00162 | TSK.910 | 38220000 | http://www.oxfordimmunotec.com/international/products-services/t-cell-select/
106668 | Leucosep Tubes for the T-SPOT.TB | 50 units | 18 months | Leucosep tubes are intended for use in the collection and separation of peripheral blood mononuclear cells (PBMCs) from whole blood. The tubes facilitate the separation of PBMCs from whole blood using a polysucrose-based gradient separation solution, prior to the use of these cells in ELISPOT assays and can be used as a substitute to the Ficoll-Paque Plus method. This item includes 50 leucosep tubes pre-filled with a polysucrose-based gradient separation solution. | Oxford Immunotec |  | 89 | USD | 0 | Cool cargo (+2-+8°C) | 0.63 | 28 | 10 | 13 | 0.003822 | LTK.615 | 38220000 | http://www.oxfordimmunotec.com/international/wp-content/uploads/sites/3/LF-TB-UK-V3.pdf
106669 | AIM-V Medium for the T-SPOT.TB, 50 ml | 1 unit | 12 months | AIM-V Medium (50 mL) is an accessory intended for use in the preparation of peripheral blood mononuclear cells (PBMCs) from whole blood and for use as a culture medium for in vitro testing using the T-SPOT.TB ELISPOT method. This item includes 1 bottle of 50 ml of AIM-V cell culture medium | Oxford Immunotec |  | 23 | USD | 0 | Cool cargo (+2-+8°C) | 0.075 | 4 | 8 | 3 | 0.000128 | A18398SA / AV.200/50 | 38210000 | http://www.oxfordimmunotec.com/international/wp-content/uploads/sites/3/PI-AIMV-UK-V2.pdf
106670 | AIM-V Medium for the T-SPOT.TB, 500 ml | 1 unit | 12 months | AIM-V Medium (500 mL) is an accessory intended for use in the preparation of peripheral blood mononuclear cells (PBMCs) from whole blood and for use as a culture medium for in vitro testing using the T-SPOT.TB ELISPOT method. This item includes 1 bottle of 500 ml of AIM-V cell culture medium | Oxford Immunotec |  | 37 | USD | 0 | Cool cargo (+2-+8°C) | 0.59 | 7 | 14.5 | 10 | 0.001015 | A18398DJ / AV.200/500 | 38210000 | http://www.oxfordimmunotec.com/international/wp-content/uploads/sites/3/PI-AIMV-UK-V2.pdf
3. DETECTION OF COVID-19
3.1 COVID-19
3.1.1 COVID-19 testing
106684 | BD MAX SARS-CoV-2 Reagents, 24 tests (445003-01) | 24 test | 4 months | Real-time RT-PCR test intended for the qualitative detection of nucleic acid from the SARS-CoV-2 in nasopharyngeal, anterior nasal, mid-turbinate, and oropharyngeal swab specimens, nasopharyngeal wash/aspirate or nasal aspirates. To be used on the BD MAX Instrument This item includes: BD SARS-CoV-2 ExK TNA 3 (LUO) BD SARS-CoV-2 Reagents and BD ExK TNA 3 (LUO) for 24 tests. This item includes: 1 kit with BD SARS-CoV-2 Reagents for BD MAX System and BD ExK TNA 3 (LUO), 24 test/kit This product should be procured together with item 106685 BD MAX PCR Cartridges, 24 cartridges (437519) | BD International |  | 408 | USD | 0 | Keep cool 2-8°C | 0.94 | 30.5 | 25.5 | 11.5 | 0.00894 | 445003 - 01 | 3822190010 | https://www.fda.gov/media/136816/download
106685 | BD MAX PCR Cartridges, 24 cartridges (437519) | 24 tests | 16 months | Microfluidic cartridge for RT-PCR incorporates a proprietary design that allows for rapid amplification and sensitive detection. Each cartridge is build with 24 independently controlled reaction chambers. | BD International |  | 8.64 | USD | 0 | General Cargo | 0.2838 | 11.2 | 5.8 | 7.7 | 4.9e-05 | 437519 | 3926909790 | https://www.bd.com/en-us/products-and-solutions/products/product-page.437519
XP3SARS-COV2-10 | Xpert Xpress SARS-CoV-2 Plus, 10 Test | 10 tests | 8 months | The Xpert Xpress CoV-2 plus test is a real-time RT-PCR test intended for the qualitative detection of nucleic acid from the SARS-CoV-2 in nasopharyngeal swab or anterior nasal swab specimen obtained from individuals meeting COVID-19 clinical and/or epidemiological criteria, as well as individuals without symptoms or other reasons to suspect COVID-19 infection. Results are for the identification of SARS-CoV-2 RNA. Positive results are indicative of the presence of SARS-CoV-2 RNA; clinical correlation with patient history and other diagnostic information is necessary to determine patient infection status. Positive results do not rule out bacterial infection or co-infection with other viruses. The agent detected may not be the definite cause of disease. Negative results do not preclude SARS-CoV-2 infection and should not be used as the sole basis for treatment or other patient management decisions. Negative results must be combined with clinical observations, patient history, and epidemiological information. This item includes: The Xpert Xpress SARS-CoV-2 Plus kit contains sufficient reagents to process 10 patients/specimens. | Cepheid HBDC |  | 149 | USD | 0 | Cool cargo (+2-+8°C) UN 2920 – CLASS 8(3)) in excepted quantity (EQ) | 2.68 | 31 | 27 | 19 | 0.015903 | SARS-CoV-2 | 3822190010 | https://www.cepheid.com/coronavirus-product-resources
4. SAMPLE AND MEDIA PREPARATION
4.1 SPECIMEN TRANPORTATION AND PROCESSING
4.1.1 Specimen transport consumables
106736 | Specimen transport cooler | 1 unit | N/A | Specimen transport cooler for the safe transport of biological samples. Transport periods: between 24 h and 8 days. • Storage capacity: 0.9 > 20 L, • Designed for tropical temperatures: +5°C to +43°C • The clasps can be sealed or equipped with locks and are therefore protected against unauthorized access during transport • Very easy to handle thanks to its adjustable shoulder strap and their light weight • 1,05 cm Polyurethane insulation the polyurethane foam injected into the double walls of these transport systems is free of CFC and HCFC and ensures optimum insulation and protection of quality of the transported goods, especially with longer transport times. This item includes: • 10 x Cooling elements of 0,6 l; 2 x Cooling element of 0,3 l; Shoulder strap - Carrying Strap (adjustable, detachable); External bags/labels for documentation and identification and Biohazard labels. | TTM |  | 455 | USD | 0 | General Cargo | 2 | 43 | 25 | 28 | 0.0301 | B Medical Systems RCW 8 | 42029219 | https://www.bmedicalsystems.com/en/product/rcw8/
106244 | Sturdy container, 2 lt | 1 unit | N/A | White plastic container 2 liters, with quarter-turn screw lid closure. This jar protects your valuable products from moisture, from contamination, and against tampering, ensuring your hazardous solids arrive safely at their destination. Wide mouth container. · Round opening of 166 cm diameter; Chemical resistance; Made of HDPE; Leak-resistant closure; HDPE has a max. temperature of +120 °C (short term) and perating temperature: -50 °C - +80 °C. | TTM |  | 7.3 | USD | 0 | General Cargo | 0.1 | 30 | 13 | 21.5 | 0.008385 | Curtec International B.V. 4320 | 39231090 | https://www.curtec.com/en/products/jars/packo/detail/packo-2000-ml
4.1.2 Specimen processing equipment
106220 | Overhead stirrer | 1 unit | N/A | Overhead stirrer for test tubes for simultaneous mixing of up to 20 vessels with diameters of 10-18 mm. Equipped with a solid metal frame. Accepts load capacities up to 1 kg; self-resetting; · Speed continuously adjustable from 20 to 100 rpm; Rotation speed range 20 - 100 rpm. This item includes: - Universal Adapter: allows for the use of various vessels of 50 - 160-mm height; - Stainless steel Adapter for 20 test tubes: allows for simultaneously mixing 20 test tubes Ø 10-18 mm with just one shaker. | TTM |  | 1891.55 | USD | 0 | General Cargo | 9.4 | 67 | 49 | 31 | 0.10177 | Heidolph Instruments GmbH & Co. KG REAX 2 | 84798200 | https://heidolph-instruments.com/en/products/Shakers-Mixers/Reax-2~p1868
106222 | Tubing for safety gas burner | 1 unit | N/A | Safety gas tube, length: 1 m, diameter: inner diameter 10 mm. Specific for use with gas; Withstand temperatures of -45 to +100°C | TTM |  | 5.95 | USD | 0 | General Cargo | 0.1 | 23 | 23 | 5 | 0.002645 | Deutsch & Neumann GmbH 3700001 | 40091100 | https://www.deutsch-neumann.de/en/gas-hose
4.1.3 Specimen processing reagents
106283 | N-acetyl-L-cysteine, 100 g | 1 unit | 16 months | N-Acetyl-L-cysteine for biochemistry. Reagent grade/purum ≥98%; CAS No 616-91-1. · pH 1.5 - 2.5 (100 g/l, H₂O, 20 °C); Density: 1.48 g/cm3 (20 °C); Melting Point:110.4 °C Subject to availability from our supplier. Product lead-time may be longer than usual. Please contact your GDF Country Supply Officer for further information | TTM |  | 180.5 | USD | 0 | General Cargo | 0.15 | 5 | 5 | 11 | 0.000275 | Merck KGaA 1.12422.0100 | 29309016 | https://www.merckmillipore.com/CH/en/product/N-Acetyl-L-cysteine,MDA_CHEM-112422
106284 | Tri-Sodium citrate dihydrate, 1 kg | 1 unit | 31 months | Tri-Sodium citrate dihydrate, for analyses. Reagent grade/purum ≥99%· 2 H2O. CAS No 6132-04-3. · Molar mass (M) 294,1 g/mol; Melting point (mp) 150 °C; Density 1,76 g/cm³. | TTM |  | 55.35 | USD | 0 | General Cargo | 1.1 | 9 | 9 | 21 | 0.001701 | Carl Roth GmbH HN12.3 | 28352930 | https://www.carlroth.com/com/en/tri-sodium-citrate-dihydrate/p/hn12.1
106287 | Sodium chloride, 1 kg | 1 unit | 31 months | Sodium chloride for analyses Reagent grade/purum >99,5%;. CAS No [7647-14-5. · Molar mass: 58,44 g/mol: Density: 2,17 g/cm³; Boiling point: >1450 °C; Melting point: 801 °C | TTM |  | 6.25 | USD | 0 | General Cargo | 1.1 | 9 | 9 | 18 | 0.0015 | Carl Roth GmbH P029.2 | 25010091 | https://www.carlroth.com/de/en/general-reagents/sodium-chloride/p/p029.2
4.1.4 Specimen processing consumables
106237 | PP test tube rack for 16 mm tubes | 1 unit | N/A | Test tube rack, for 16 mm diameter tubes, made of Polypropylene. Autoclavable at least at 121°C; Stackable; Total of 60 places; Rows with alphanumeric labelling; | TTM |  | 6.95 | USD | 0 | General Cargo | 0.1 | 24 | 12 | 8 | 0.002304 | Heinz Herenz 1131956 | 39269097 | https://www.herenz.de/Kataloge/Englisch_2021.pdf
106238 | PP test tube rack for 18 mm tubes | 1 unit | N/A | Test tube rack, for 18 mm diameter tubes, made of Polypropylene. Autoclavable at least at 121°C; Stackable; Total of 40 places; Rows with alphanumeric labelling; | TTM |  | 6.65 | USD | 0 | General Cargo | 0.1 | 26 | 12 | 8 | 0.002496 | Heinz Herenz 1131957 | 39269097 | https://www.herenz.de/Kataloge/Englisch_2021.pdf
106239 | PP test tube rack for 30 mm tubes | 1 unit | N/A | Test tube rack, for 30 mm diameter tubes, made of Polypropylene. Autoclavable at least at 121°C; Stackable; Total of 24 places; Rows with alphanumeric labelling. | TTM |  | 8.65 | USD | 0 | General Cargo | 0.2 | 20 | 20 | 10 | 0.004 | Heinz Herenz Medizinalbedarf GmbH 1131959 | 39269097 | https://www.herenz.de/Kataloge/Englisch_2021.pdf
106230 | Stainless steel test tube rack for 18 mm tubes | 1 unit | N/A | Test tube rack for 18 mm tubes, made of stainless steel; 3 x 12 places (= 36 places); Autoclavable at 121 °C. | TTM |  | 26.85 | USD | 0 | General Cargo | 0.2 | 25 | 8 | 7 | 0.0014 | Hupfer Metallwerke GmbH 7500628 | 39269097 | https://ie.vwr.com/store/product/568444/racks-for-test-tubes
106231 | Stainless steel test tube rack for 35 mm tubes | 1 unit | N/A | Test tube rack for 18 mm tubes, made of stainless steel, 1.4301, polished; for tube with a ø of 35 mm; Autoclavable at 121 °C. | TTM |  | 74.3 | USD | 0 | General Cargo | 0.75 | 27 | 13 | 12 | 0.004212 | Württembergische Drahtwarenfabrik Funke GmbH Special Production 106231 | 39269097
106509 | pH strips | 100 units | 16 months | pH- indication strips. Measuring range at least (pH): 6,5 – 10; · With color coded scale; with different color codes within each step · pH gradation including 6.5 - 6.8 - 7.1 - 7.4 - 7.7 - 7.9 - 8.1 - 8.3 - 8.5 - 8.7 - 9.0 - 9.5 - 10.0 pH | TTM |  | 23.6 | USD | 0 | General Cargo | 0.3 | 10 | 9 | 3 | 0.00027 | Merck KGaA 1.09543.0001 | 28530090 | https://www.merckmillipore.com/CH/en/product/pH-indicator-strips-pH-6.5-10.0,MDA_CHEM-109543?ReferrerURL=https%3A%2F%2Fwww.google.com%2F
106355 | Permanent marker | 1 unit | N/A | Permanent marker for marking and labelling almost all materials, e.g. paper, cardboard, metal, plastic and glass. The round nib has a stroke width of 1.5-3 mm. Waterproof and wear-resistant on almost all surfaces. | TTM |  | 2.05 | USD | 0 | General Cargo | 0.1 | 15 | 8 | 4 | 0.0005 | Edding International GmbH edding 3000 | 96082000 | https://www.edding.com/products/edding-3000-permanent-marker/
106738 | Permanent markers with 1 refill | 4 units | N/A | Permanent marker pen for PCR tubes and laboratory use. ACMI Certified AP product, nontoxic ink; . Dry-safe type acc. to ISO 554. Tip 0,6 mm; . Resistant against alcohol and water. . Includes a set of 4 units and 1 refill station | TTM |  | 12.7 | USD | 0 | Dangerous cargo UN 1263-Paint, Class 3, PG II | 0.1 | 15 | 8 | 10 | 0.0012 | Staedler Noris GmbH Lumocolor 318 + 487 17 | 96082000 | https://www.staedtler.com/de/de/produkte/marker/folienstifte/lumocolor-permanent-pen-318-permanent-universalstift-f-318-wp4/
106356 | Adhesive labels 50 x 19 mm | 486 units | N/A | White multipurpose Labels adhesive. Writeable by hand, for fast and easy use · No printer needed, labels can be written on by hand; Bleached without chlorine. This item includes: 486 labels/ 27 sheets | TTM |  | 5.9 | USD | 0 | General Cargo | 0.1 | 18.5 | 13.5 | 1 | 0.0002 | Avery Dennison Materials GmbH Avery Zweckform 3327 | 48219010 | https://www.avery.ae/product/multipurpose-general-use-labels-3327
4.2 MEDIA PREPARATION
4.2.1 Media preparation equipment
106257 | Laboratory oven 32 lt | 1 unit | N/A | Benchtop heating oven for laboratory with forced air, made of stainless steel with a temperature range of 20 °C (68 °F) to 300 °C (572 °F). Capacity of 32 l (8.5 gal). Setting accuracy up to 99.9°C: 0.1°C / from 100°C: 0.5°C; Digital backwards counter with target time setting, adjustable from 1 minute to 99 days. · PID microprocessor control with integrated auto-diagnostic system; One high-resolution TFT colour display; touch-sensitive buttons for selection of functions, amongst others for temperature and timer · Vent connection with restrictor flap; mechanical temperature limiter TB, protection class 1 according to DIN 12880 to switch off the heating approx. 20°C above nominal temperature. This item includes: 3 stainless steel shelves. | TTM |  | 1306.55 | USD | 0 | General Cargo | 43.5 | 65 | 66 | 89 | 0.38181 | Memmert GmbH + Co. KG UF30 | 84198998 | https://www.memmert.com/products/heating-drying-ovens/universal-oven/UF30/
106099 | Laboratory oven 256 lt | 1 unit | N/A | Heating oven for laboratory with forced air, made of stainless steel with a temperature range of 20 °C (68 °F) to 300 °C (572 °F). Capacity of 256 lt. Setting accuracy up to 99.9°C: 0.1°C / from 100°C: 0.5°C; Digital backwards counter with target time setting, adjustable from 1 minute to 99 days. · PID microprocessor control with integrated auto-diagnostic system; One high-resolution TFT colour display; touch-sensitive buttons for selection of functions, amongst others for temperature and timer · Vent connection with restrictor flap; mechanical temperature limiter TB, protection class 1 according to DIN 12880 to switch off the heating approx. 20°C above nominal temperature | TTM |  | 3530.65 | USD | 0 | General Cargo | 166 | 93 | 93 | 138 | 1.1936 | Member GmbH + Co. KG UF260 | 84198998 | https://www.memmert.com/products/heating-drying-ovens/universal-oven/UF260/
106100 | Shelves for laboratory oven 256 lt | 1 unit | N/A | Perforated stainless steel shelf for Laboratory Hot-air oven 256 lt. Shelfs support a max. load of 20kg. | TTM |  | 111 | USD | 0 | General Cargo | 2 | 82 | 58 | 24 | 0.11414 | Memmert GmbH + Co. KG B29725 | 84199085 | https://www.memmert.com/de/produkte/optionen-zubehoer/brutschrank-im/IN260mplus/
106136 | Automated culture media dispenser | 1 unit | N/A | Automatic culture media dispenser, with a flow rate of 2.0 - 4,056 ml / min, when operated with single-channel pump heads, flows from feed-batch and chemostat modes of operation can be conveyed precisely. This pump drive is not suitable for multi-channel operation. · Flow rates of 2.0 - 4,056 ml/min with single-channel pump heads; High speed range from 24 - 600 rpm · Control speed (rpm), direction and On/Off function via analog interface for 0 – 10 V or 4 – 20 mA or digital RS 232 interface. This item includes: A - Foot pedal; B - Aspiration/dispensing tubes; C - Tubing set of at least six (6) different diameters in the range of 1 to 8 mm; D - Preferably tubing holder for free-hand operation. | TTM |  | 3394.2 | USD | 0 | General Cargo | 9 | 46 | 26 | 32 | 0.038272 | Heidolph Instruments GmbH & Co. KG 523-52060-00 | 84136080 | https://heidolph-instruments.com/en/products/Peristaltic-Pumps/Hei-FLOW-Precision-06~p1008
106127 | Water distiller | 1 unit | N/A | Fully automatic water distiller with electronic control system, made of stainless steel. Capacity 4 l/h water supply connected through a built-in tank of 8 lt. Thermostatic low water cut-off. · Evaporator, condenser (cooling coil), heating element and storage tank made of stainless steel; · Storage tank for distillate accepts double the hourly capacity of the unit; · Water supply through an integrated solenoid valve with connection for pressure hose ½ inch* (inner Ø 12.7 mm), water pressure required: > 3 bar to max. 7 bar; Distillate withdrawal at the front; · Cooling water outlet with hose connection ¾ inch* (inner Ø 19 mm). | TTM |  | 1914.45 | USD | 0 | General Cargo | 24 | 69 | 43 | 59 | 0.17505 | GFL Gesellschaft für Labortechnik mbH 2004 + 2941 | 84194000 | https://www.irmeco.de/VGKL16/pdf/en/KG16/05/04/KG16_05_04_0640.pdf
106137 | Glassware stand set | 1 unit | N/A | Stand plate: Steel, varnished: Firm, non-slip stand. This item includes: A· The base that is made of steel, powder coated, Thickness 8 mm, dimensions: 315x200 mm 4x M 10 threaded rubber feet; Support rod mad of zinc plated steel, removable, ø: 12 mm, length: 1000 mm; Built-in rod holder with threaded hole to support the rod; B· A clamp of 10 mm, length of extension rod:1000 mm; clamp with wing screw for easy open and close of the jaws; cork lined jaws; adjustable from 20 to 40 mm; made of powder coated zinc diecasting; C· A bosshead with hand adjustable screws (made of nickel plated brass), allows 90° angulation, for rods up to 16 mm ø, powder coated zinc diecasting. | TTM |  | 48.35 | USD | 0 | General Cargo | 4 | 32 | 32 | 20 | 0.02048 | Carl Friedrich Usbeck KG 2005 + 2144 + 2802 + 2201 | 39269097 | https://www.usbeck.eu/en/index.php?artikel&gruppe=2800
106594 | Cryo-gloves size L for laboratory deep freezer | 1 unit | N/A | Reusable cryo gloves; Waterproof; Large size: 345-390 mmh, Mid-Arm; Withstand temperatures of -196°C; Breathable; Static-free; Resistant to chemicals and solvents used in laboratories; Latex- free, powder free; Lightweight; Non-sterile. This item includes: 1 pair of gloves. | TTM |  | 143.4 | USD | 0 | General Cargo | 0.198 | 48.5 | 6 | 15 | 0.004365 | Tempshield MAL WP | 39262000 | https://tempshield.com/products/waterproof-cryo-gloves?variant=12515021389890
106595 | Voltage stabilizer 2000W for laboratory deep freezer | 1 unit | N/A | Automatic voltage stabilizer for deep-freezer 556 lt; Voltage stabilizer with AC Input and Output plugs conforming to IEC60320 type. Efficiency of 97% able to work in high or low line voltage. Voltage stabilizer with led indicator for high, low and normal input line voltage. · The Voltage stabilizer helps against voltage fluctuations, the UPS protects also against power failure; · Protection against voltage peaks and short circuits with 1 regulated and stabilized output sockets. | TTM |  | 255.95 | USD | 0 | General Cargo | 6.5 | 57 | 40 | 19 | 0.04332 | Sollatek UK SVS16(15)-22-CE | 85363010 | https://www.sollatek.com/wp-content/uploads/2020/04/SVS-brochure-2020-April-2020.pdf
106130 | Digital benchtop pH meter | 1 unit | N/A | Digital benchtop pH meter with automatic microprocessor controlled equipment. pH meter with 1 to 3 point calibration. Measurement max. 10 sec; measures PH, Temperature and mV. · Automatic buffer recognition; automatic calibration. Temperature measurement range -10 – 110°C. · LCD 3-color display, with icon support all calibration editing function with alarm for electrode diagnostic; USB-port, RS-232 (data memory of 500 data) · pH measurement range for 2-16 pH with a resolution of 0.01 /0.1, accuracy of 0,01; This item includes: pH electrode; base holder with electrode arm; temperature probe; ph calibration kit and consumables for 500 test. | TTM |  | 458.25 | USD | 0 | General Cargo | 1.1 | 26 | 20 | 10 | 0.0052 | Phoenix Instrument/ LAT Labor und Analysen-Technik EC-35 pH | 90278930 | http://www.phoenix-instrument.de/e_EC-35-ph.html
106562 | Consumable kit for digital benchtop pH meter | 1 unit | 15 months | Consumable kit for digital benchtop pH meter. Ready to use solutions. This item includes: A. Buffer solutions for electrode calibration: 20x25 ml buffer solution for 4,01 pH; 20x25 ml buffer solution for 6,86 pH; 20x25 ml buffer solution for 7,0 pH; 20x25 ml buffer solution for 9,1879,21 pH and 20x25 ml buffer solution for 10,0 or 10,01 pH; B. Storage solution for electrode, to minimize microbial growth and keep the pH electrode moist: 1 bottle of 500 ml solution and Light blocking bottle, FDA compliant; C. Cleaning solutions to eliminate impurities and residues on electrode surface: 500 ml general cleaning solution and 500 ml cleaning solution for proteins, to remove proteins from a pH electrode. | TTM |  | 171.4 | USD | 0 | General Cargo | 3 | 25 | 30 | 20 | 0.015 | Phoenix Instrument/ LAT Labor und Analysen-Technik consumable kit for Phoenix pH meter | 90278011 | http://www.phoenix-instrument.de/PDF/Electrochemistry.pdf
106132 | Magnetic hot plate stirrer | 1 unit | N/A | Magnetic stirrer with heating plate with integrated temperature control. Aluminium heating plate for optimum heat transfer. Speed range of 50 - 1500 rpm and temperature Setting Range of 0-310°C · Hot plate of Ø 135 mm; Stirring quantity max. per stirring position (H2O) 20 lt; · Digital display for visualization of speed and temperature settings; Heat control by turning knob; · Equipment with over-temperature protection; This item includes: all the accessories: Thermometer, supporting rod and clamp and stir bar. | TTM |  | 1224.55 | USD | 0 | General Cargo | 3.7 | 46 | 18 | 14 | 0.011592 | IKA Werke GmbH IKA RCT basic | 84798200 | https://www.ika.com/en/Products-Lab-Eq/Magnetic-Stirrers-Hot-Plate-Lab-Mixer-Stirrer-Blender-csp-188/RCT-basic-cpdt-3810000/
106720 | Magnetic stirring bars set of 18 | 1 unit | N/A | White magnetic bars set; PTFE covered, 18 cylindrical magnetic stirring bars kept in a transparent box for storage. · Magnetic bars with low friction; Autoclavable at 121 °C; Storage box made of PS; This item includes: an encapsulated magnet and 18 magnetic bars with the following dimensions: 10× 6, 15× 4,5, 20× 6, 25× 6, 30× 6, 40× 8, 50× 8, 60× 10 and 70× 10 mm length | TTM |  | 86.35 | USD | 0 | General Cargo | 0.8 | 10 | 10 | 4 | 0.0004 | VVWR International GmbH 442-0519 | 85051990 | https://uk.vwr.com/store/product/10908873/magnetic-stirring-bars-boxed-sets
4.2.2 Media preparation plasticware
106157 | Swan-neck bottles set of 250ml, 500ml and 1,000ml | 1 set | N/A | Set of 3 swan neck bottles for safe storage of sensitive solutions. Bottle made of Polyethylene with a swan neck, lid and dispenser tube; With dispenser tube; Secure leakage proof closing mechanism. · Squeeze type dispensing bottles for staining solutions; Resistant to staining solutions, and most common cleaning solutions. This item includes: 1 bottle of 250 ml, 1 bottle of 500 ml and 1 bottle of 1000ml. | TTM |  | 4 | USD | 0 | General Cargo | 0.1 | 18 | 13 | 36 | 0.008424 | Witeg Labortechnik GmbH 7.169.003, 7.169.004, 7.169.005 | 39239000 | https://www.witeg.de/en/detail/index/sArticle/8256
106158 | Laboratory tray 12 lt | 1 unit | N/A | Laboratory tray of 12 litres; autoclavable tray for a variety of uses. Made of PP/PVC transparent, multi usage for storage and laboratory utensils. Tray with flat bottom and rounded corners; Very good chemical and thermal resistance · It can hold 18x500 ml or 12x1 lt bottles; Tray for holding sample bottles with base diameter not greater than 22 mm. Ideal for 7 ml (1/4oz) Bijou bottles or specimen tubes. | TTM |  | 8.1 | USD | 0 | General Cargo | 0.3 | 52 | 35 | 12 | 0.02184 | Kartell SPA 548 | 39239000 | https://www.kartelllabware.com/en/products/plastilab/general-purpose-labware/input-trays-1/
106201 | Laboratory drying rack for glassware | 1 unit | N/A | Drying rack for laboratory glassware made of PE-coated wire with 24 rods (= small pegs) and 20 arches (= large pegs) with fixed rods/arches · The rods/arches are suitable for test tubes and for flasks/beakers; free-standing or wall-mounted and with drip pan; The feet are protected by plastic caps; Very good chemical resistance. This item includes: all accessories needed for hanging on wood, concrete and plasterboard walls. | TTM |  | 118.45 | USD | 0 | General Cargo | 2 | 50 | 20 | 100 | 0.1 | Württembergische Drahtwarenfabrik Funke GmbH 5281-01-11 | 73262000
106202 | PP funnel 81 mm | 1 unit | N/A | Analytical funnel, made of PP. Designed specifically for analytical chemistry. Each size takes a standard size filter paper. The body of funnel is exact 60° angle. Provided with external and internal ribs. Autoclavable. Very good chemical resistance. | TTM |  | 0.75 | USD | 0 | General Cargo | 0.8 | 10 | 10 | 15 | 0.0015 | Kartell SPA 150 | 39269097 | https://www.kartelllabware.com/en/products/plastilab/filtration-and-vacuum-pumps/analytical-funnels/
106204 | PP powder funnel 100 mm | 1 unit | N/A | Powder funnel translucent, made of transparent PP, 60° angle, autoclavable. Large bore avoids bridging of powders during transfer. Ribbed outside to avoid pressure lock. Can be used with liquids. Excellent chemical resistance and high temperature resistance. | TTM |  | 2 | USD | 0 | General Cargo | 0.2 | 12 | 12 | 15 | 0.00216 | Kartell SPA 168 | 39269097 | https://www.kartelllabware.com/en/products/plastilab/general-purpose-labware/powder-funnels/
106171 | PMP measuring cylinder 1000 ml | 1 unit | N/A | Measuring cylinder, 1000 ml, made of PMP with excellent chemical resistance and chemically non absorbent. Cylinder tall form with non wetting surface. The measuring cylinder has a pentagonal base. · It can be used with liquids up to +170 °C; autoclavable at 121 °C; · Measuring cylinder graduated (10,0 ml), molded blue graduations. | TTM |  | 19.55 | USD | 0 | General Cargo | 0.3 | 12 | 12 | 15 | 0.00216 | Kartell SPA 2576 | 39223010 | https://www.kartelllabware.com/en/products/plastilab/graduated-and-volumetric-plasticware/blue-graduated-measuring-cylinders-tall-form-class-b-1/
106172 | PMP measuring cylinder 500 ml | 1 unit | N/A | Measuring cylinder, 500 ml, made of PMP with excellent chemical resistance and chemically non absorbent. Cylinder tall form with non wetting surface. The measuring cylinder has a pentagonal base. · It can be used with liquids up to +170 °C; autoclavable at 121 °C; · Measuring cylinder graduated (5,0 ml), molded blue graduations. | TTM |  | 12.05 | USD | 0 | General Cargo | 0.3 | 12 | 12 | 15 | 0.00216 | Kartell SPA 2575 | 39223010 | https://www.kartelllabware.com/en/products/plastilab/graduated-and-volumetric-plasticware/blue-graduated-measuring-cylinders-tall-form-class-b-1/
106173 | PMP measuring cylinder 250 ml | 1 unit | N/A | Measuring cylinder, 250 ml, made of PMP with excellent chemical resistance and chemically non absorbent. Cylinder tall form with non wetting surface. The measuring cylinder has a pentagonal base. · It can be used with liquids up to +170 °C; autoclavable at 121 °C; · Measuring cylinder graduated (2,0 ml), molded blue graduations. | TTM |  | 9.5 | USD | 0 | General Cargo | 0.3 | 12 | 12 | 15 | 0.00216 | Kartell SPA 2574 | 39223010 | https://www.kartelllabware.com/en/products/plastilab/graduated-and-volumetric-plasticware/blue-graduated-measuring-cylinders-tall-form-class-b-1/
106174 | PMP measuring cylinder 100 ml | 1 unit | N/A | Measuring cylinder, 100 ml, made of PMP with excellent chemical resistance and chemically non absorbent. Cylinder tall form with non wetting surface. The measuring cylinder has a pentagonal base. · It can be used with liquids up to +170 °C; autoclavable at 121 °C; · Measuring cylinder graduated (1,0 ml), molded blue graduations. | TTM |  | 5.25 | USD | 0 | General Cargo | 0.3 | 12 | 12 | 15 | 0.00216 | Kartell SPA 2573 | 39223010 | https://www.kartelllabware.com/en/products/plastilab/graduated-and-volumetric-plasticware/blue-graduated-measuring-cylinders-tall-form-class-b-1/
106175 | PMP measuring cylinder 50 ml | 1 unit | N/A | Measuring cylinder, 50 ml, made of PMP with excellent chemical resistance and chemically non absorbent. Cylinder tall form with non wetting surface. The measuring cylinder has a pentagonal base. · It can be used with liquids up to +170 °C; autoclavable at 121 °C; · Measuring cylinder graduated (1,0 ml), molded blue graduations. | TTM |  | 4.65 | USD | 0 | General Cargo | 0.3 | 12 | 12 | 15 | 0.00216 | Kartell SPA 2572 | 39233010 | https://www.kartelllabware.com/en/products/plastilab/graduated-and-volumetric-plasticware/blue-graduated-measuring-cylinders-tall-form-class-b-1/
106402 | Plastic bottle 500 ml | 1 unit | N/A | Narrow mouth bottle of 500 ml, made of transparent LDPE. Graduated of 100 ml. Flexible, leakproof, washable bottle; BPA-Free; High chemical resistance; Withstand low and high temperatures of 121°C; Bottle with threaded neck standard DIN GL25 and Standard thread DIN GL 25 screw cap; Height of 182 and Neck diameter of 18 mm. | TTM |  | 8.45 | USD | 0 | General Cargo | 0.05 | 8 | 8 | 18 | 0.001152 | Carl Roth GmbH A752.1 | 39233010 | https://www.carlroth.com/com/en/plastic-bottles/narrow-mouth-bottle/p/a752.1
4.2.3 Media preparation glassware
106134 | Glass dropping funnel 1000 ml | 1 unit | N/A | Glass dropping funnel of culture media 1000 ml, 20 ml step graduation. Funnel with built-in pressure equalizing tube and stopper, made of borosilicate glass 3.3 . Cylindrical with glass stopper. · Standard NS 29/32 inner and outer joints with high chemical resistance; · Built-in stopcock bore 6,3 mm; Stopcock made of PTFE. | TTM |  | 143.8 | USD | 0 | General Cargo | 1.5 | 15 | 10 | 35 | 0.00525 | Witeg Labortechnik GmbH 1895108 | 39269097 | https://www.witeg.de/en/detail/index/sArticle/3024
106135 | Glass dropping funnel 250 ml | 1 unit | N/A | Glass dropping funnel of culture media 250 ml, 5 ml step graduation. Funnel with built-in pressure equalizing tube and stopper, made of borosilicate glass 3.3 . Cylindrical with glass stopper. · Standard NS 29/32 inner and outer joints with high chemical resistance; · Built-in stopcock bore 4 mm; Stopcock made of PTFE. | TTM |  | 83.3 | USD | 0 | General Cargo | 0.4 | 42 | 13 | 10 | 0.00546 | Witeg Labortechnik GmbH 1895106 | 39269097 | https://www.witeg.de/en/detail/index/sArticle/3022
106203 | Glass funnel 120 mm | 1 unit | N/A | Plain glass funnel with top diameter of 120 mm. Made by soda lime glass with short and angled-ground stern. Body Angle 60°. · Resistant to the chemicals commonly used in laboratory and high temperature resistance; · Short and angled ground stem and wall thickness of at least 2 mm. | TTM |  | 7.3 | USD | 0 | General Cargo | 1 | 15 | 15 | 20 | 0.0045 | Witeg Labortechnik GmbH 5770011 | 70172000 | https://www.witeg.de/en/detail/index/sArticle/5583
106152 | Woulff bottle 1000 ml kit | 1 kit | N/A | Woulff’s bottle 1000 ml with graduation of 100 ml, Borosilicate glass 3.3. Woulff’s bottle complete with its kit, used to saturate a liquid with a gas or a gas with another. Bottle with threaded neck standard GL45; Top insert according to Drechsel-head without filter; Adjustable insert height with two (2) threaded connections; Autoclavable with high chemical and thermal resistance. This item includes: · 2 straight plastic hose connections (olive shape) of at least 8mm diameter; · 2 lateral standard GL 14 screw caps with silicone seal for hose connections; · 1 central screw cap GL45 with silicone or PTFE seal. | TTM |  | 57.85 | USD | 0 | General Cargo | 1 | 20 | 20 | 25 | 0.01 | DWK Life Sciences GmbH 218105403, 247130205, 292550603 (2 pcs.), 292270508 (2 pcs.), 292271007 and 292282501 | 70172000 | https://www.dwk.com/duran-pressure-plus-gl-45-laboratory-bottle-clear-without-screw-cap-and-pouring-ring-1000-ml-218105403
106146 | Glass measuring pipette 1 ml | 10 units | N/A | Measuring pipette of 1ml, made of AR-glass with high chemicals resistance. Pipette with ring mark; · Durably printed in blue graduation; graduation of 0,01ml. Class AS, DIN EN ISO 835; · Tolerance of at least ±0,007ml; Length oft 350 mm; Calibrate to deliver. | TTM |  | 37.5 | USD | 0 | General Cargo | 0.5 | 4 | 4 | 35 | 0.0006 | Brand GmbH 27817 | 70179000 | https://shop.brand.de/en/p347
106147 | Glass measuring pipette 2 ml | 10 units | N/A | Measuring pipette of 2ml, made of AR-glass with high chemicals resistance. Pipette with ring mark; · Durably printed in blue graduation; graduation of 0,02ml. Class AS, DIN EN ISO 835; · Tolerance of at least ±0,010ml; Length oft 350 mm; Calibrate to deliver. | TTM |  | 37.5 | USD | 0 | General Cargo | 0.5 | 4 | 4 | 35 | 0.0006 | Brand GmbH 27819 | 70179000 | https://shop.brand.de/en/p347
106148 | Glass measuring pipette 5 ml | 10 units | N/A | Measuring pipette of 5ml, made of AR-glass with high chemicals resistance. Pipette with ring mark; · Durably printed in blue graduation; graduation of 0,05ml. Class AS, DIN EN ISO 835; · Tolerance of at least ±0,030ml; Length oft 350 mm; Calibrate to deliver. | TTM |  | 42.75 | USD | 0 | General Cargo | 0.5 | 4 | 4 | 35 | 0.0006 | Brand GmbH 27821 | 70179000 | https://shop.brand.de/en/p347
106149 | Glass measuring pipette 10 ml | 10 units | N/A | Measuring pipette of 10ml, made of AR-glass with high chemicals resistance. Pipette with ring mark; · Durably printed in blue graduation; graduation of 0,1ml. Class AS, DIN EN ISO 835; · Tolerance of at least ±0,050ml; Length oft 350 mm; Calibrate to deliver. | TTM |  | 42.75 | USD | 0 | General Cargo | 0.3 | 38 | 6 | 5 | 0.00114 | Brand GmbH 27823 | 70179000 | https://shop.brand.de/en/p347
106159 | Erlenmeyer flask 2000 ml | 10 units | N/A | Erlenmeyer flask for general laboratory purpose 2000 ml, made of borosilicate glass 3.3. High chemicals resistance. Flasks withstand temperature changes; Autoclavable at least at 121°C; Wall thickness of at least 2,5 mm with wide neck. Durably printed graduation | TTM |  | 85.45 | USD | 0 | General Cargo | 5.77 | 80 | 35 | 29 | 0.0371 | Heinz Herenz Medizinalbedarf GmbH 1070238 | 70172000 | https://www.herenz.de/Kataloge/Englisch_2021.pdf
106160 | Erlenmeyer flask 1000 ml | 10 units | N/A | Erlenmeyer flask for general laboratory purpose 1000 ml, made of borosilicate glass 3.3. High chemicals resistance. Flasks withstand temperature changes; Autoclavable at least at 121°C; Wall thickness of at least 2,5 mm with wide neck. Durably printed graduation | TTM |  | 44.45 | USD | 0 | General Cargo | 2.9 | 70 | 28 | 24 | 0.047 | Heinz Herenz Medizinalbedarf GmbH 1070237 | 70172000 | https://www.herenz.de/Kataloge/Englisch_2021.pdf
106161 | Erlenmeyer flask 500 ml | 10 units | N/A | Erlenmeyer flask for general laboratory purpose 500 ml, made of borosilicate glass 3.3. High chemicals resistance. Flasks withstand temperature changes; Autoclavable at least at 121°C; Wall thickness of at least 2,5 mm with wide neck. Durably printed graduation | TTM |  | 29.2 | USD | 0 | General Cargo | 2.4 | 56 | 23 | 20 | 0.0258 | Heinz Herenz Medizinalbedarf GmbH 1070236 | 70172000 | https://www.herenz.de/Kataloge/Englisch_2021.pdf
106162 | Erlenmeyer flask 250 ml | 10 units | N/A | Erlenmeyer flask for general laboratory purpose 250 ml, made of borosilicate glass 3.3. High chemicals resistance. Flasks withstand temperature changes; Autoclavable at least at 121°C; Wall thickness of at least 2,5 mm with wide neck. Durably printed graduation | TTM |  | 20.75 | USD | 0 | General Cargo | 1.8 | 47 | 19 | 17 | 0.0152 | Heinz Herenz Medizinalbedarf GmbH 1070234 | 70172000 | https://www.herenz.de/Kataloge/Englisch_2021.pdf
106163 | Erlenmeyer flask 100 ml | 10 units | N/A | Erlenmeyer flask for general laboratory purpose 100 ml, made of borosilicate glass 3.3. High chemicals resistance. Flasks withstand temperature changes; Autoclavable at least at 121°C; Wall thickness of at least 2,5 mm with wide neck. Durably printed graduation | TTM |  | 19.05 | USD | 0 | General Cargo | 1 | 38 | 15 | 13 | 0.00741 | Heinz Herenz Medizinalbedarf GmbH 1070232 | 70172000 | https://www.herenz.de/Kataloge/Englisch_2021.pdf
106164 | Erlenmeyer flask 50 ml | 10 units | N/A | Erlenmeyer flask for general laboratory purpose 50 ml, made of borosilicate glass 3.3. High chemicals resistance. Flasks withstand temperature changes; Autoclavable at least at 121°C; Wall thickness of at least 2,5 mm with wide neck. Durably printed graduation | TTM |  | 54.3 | USD | 0 | General Cargo | 0.5 | 27 | 11 | 10 | 0.00297 | Heinz Herenz Medizinalbedarf GmbH 1070231 | 70172000 | https://www.herenz.de/Kataloge/Englisch_2021.pdf
106165 | Glass measuring cylinder 1000 ml | 1 unit | N/A | Measuring cylinder 1000 ml, class A with blue printed graduation in 10 ml. Made of DURAN borosilicate glass 3.3. Wall thickness of 2,5 mm and tall form with spout (max. height 470 mm). Autoclavable. · Hexagonal plastic base and protective collar with spout; Acid. alkali and chemical resistant. | TTM |  | 40.9 | USD | 0 | General Cargo | 0.6 | 47 | 15 | 14 | 0.0099 | Hirschmann Laborgeräte GmbH 2270193 | 70179000 | https://www.hirschmannlab.com/artikelgruppe/22701
106166 | Glass measuring cylinder 500 ml | 1 unit | N/A | Measuring cylinder 500 ml, class A with blue printed graduation in 5 ml. Made of DURAN borosilicate glass 3.3. Wall thickness of 2,5 mm and tall form with spout (max. height 470 mm). Autoclavable. · Hexagonal plastic base and protective collar with spout; Acid. alkali and chemical resistant. | TTM |  | 31.3 | USD | 0 | General Cargo | 0.7 | 13 | 13 | 36 | 0.0061 | Hirschmann Laborgeräte GmbH 2270190 | 70179000 | https://www.hirschmannlab.com/artikelgruppe/22701
106167 | Glass measuring cylinder 250 ml | 1 unit | N/A | Measuring cylinder 250 ml, class A with blue printed graduation in 2,5 ml. Made of DURAN borosilicate glass 3.3. Wall thickness of 2,5 mm and tall form with spout (max. height 470 mm). Autoclavable. · Hexagonal plastic base and protective collar with spout; Acid. alkali and chemical resistant. | TTM |  | 19.8 | USD | 0 | General Cargo | 0.5 | 42 | 12 | 10 | 0.00504 | Hirschmann Laborgeräte GmbH 2270186 | 70179000 | https://www.hirschmannlab.com/artikelgruppe/22701
106168 | Glass measuring cylinder 100 ml | 1 unit | N/A | Measuring cylinder 100 ml, class A with blue printed graduation in 1 ml. Made of DURAN borosilicate glass 3.3. Wall thickness of 2,5 mm and tall form with spout (max. height 470 mm). Autoclavable. · Hexagonal plastic base and protective collar with spout; Acid. alkali and chemical resistant. | TTM |  | 11.25 | USD | 0 | General Cargo | 0.06 | 8 | 8 | 26 | 0.0017 | Hirschmann Laborgeräte GmbH 2270180 | 70179000 | https://www.hirschmannlab.com/artikelgruppe/22701
106169 | Glass measuring cylinder 50 ml | 1 unit | N/A | Measuring cylinder 50 ml, class A with blue printed graduation in 1 ml. Made of DURAN borosilicate glass 3.3. Wall thickness of 2,5 mm and tall form with spout (max. height 470 mm). Autoclavable. · Hexagonal plastic base and protective collar with spout; Acid. alkali and chemical resistant. | TTM |  | 10.4 | USD | 0 | General Cargo | 0.1 | 25 | 8 | 7 | 0.0014 | Hirschmann Laborgeräte GmbH 2270175 | 70179000 | https://www.hirschmannlab.com/artikelgruppe/22701
106170 | Glass measuring cylinder 25 ml | 1 unit | N/A | Measuring cylinder 25 ml, class A with blue printed graduation in 0,5 ml. Made of DURAN borosilicate glass 3.3. Wall thickness of 2,5 mm and tall form with spout (max. height 470 mm). Autoclavable. · Hexagonal plastic base and protective collar with spout; Acid. alkali and chemical resistant. | TTM |  | 8.7 | USD | 0 | General Cargo | 0.05 | 6 | 6 | 17 | 0.000612 | Hirschmann Laborgeräte GmbH 2270170 | 70179000 | https://www.hirschmannlab.com/artikelgruppe/22701
106721 | Glass flask 5000 ml | 1 unit | N/A | Laboratory glass flask bottle 5000 ml with narrow neck, with screw cap for laboratory general purpose with DIN thread GL 45, pouring ring and blue screw cap with durably printed graduation; · Made of borosilicate glass 3.3, chemical resistant; High chemicals resistance; Autoclavable at 121°C. | TTM |  | 61.8 | USD | 0 | General Cargo | 1 | 25 | 25 | 40 | 0.025 | VWR International GmbH Kavalier 632414321956 | 70172000 | https://ch.vwr.com/store/product/it/9873355/laboratory-bottles-narrow-neck-with-screw-cap
106177 | Glass flask 2000 ml | 1 unit | N/A | Laboratory glass flask bottle 2000 ml with narrow neck, with screw cap for laboratory general purpose with DIN thread GL 45, pouring ring and blue screw cap with durably printed graduation; · Made of borosilicate glass 3.3, chemical resistant; High chemicals resistance; Autoclavable at 121°C. | TTM |  | 9.05 | USD | 0 | General Cargo | 1 | 25 | 25 | 40 | 0.025 | VWR International GmbH Kavalier 632414321950 | 70172000 | https://ch.vwr.com/store/product/it/9873355/laboratory-bottles-narrow-neck-with-screw-cap
106178 | Glass flask 1000 ml | 1 unit | N/A | Laboratory glass flask bottle 1000 ml with narrow neck, with screw cap for laboratory general purpose with DIN thread GL 45, pouring ring and blue screw cap with durably printed graduation; · Made of borosilicate glass 3.3, chemical resistant; High chemicals resistance; Autoclavable at 121°C. | TTM |  | 4.75 | USD | 0 | General Cargo | 0.6 | 9 | 9 | 24 | 0.001944 | VWR International GmbH Kavalier 632414321940 | 70172000 | https://ch.vwr.com/store/product/it/9873355/laboratory-bottles-narrow-neck-with-screw-cap
106179 | Glass flask 500 ml | 1 unit | N/A | Laboratory glass flask bottle 500 ml with narrow neck, with screw cap for laboratory general purpose with DIN thread GL 45, pouring ring and blue screw cap with durably printed graduation; · Made of borosilicate glass 3.3, chemical resistant; High chemicals resistance; Autoclavable at 121°C. | TTM |  | 2.45 | USD | 0 | General Cargo | 3.9 | 46 | 19 | 21 | 0.018354 | VWR International GmbH Kavalier 632414321500 | 70172000 | https://ch.vwr.com/store/product/it/9873355/laboratory-bottles-narrow-neck-with-screw-cap
106183 | Glass laboratory beaker 2000 ml | 1 unit | N/A | Laboratory beaker 2000 ml for general laboratory purpose, made of borosilicate glass 3.3, according to DIN 12331, ISO 3819 with low form with spout; Acid, alkali resistant; Autoclavable at 121°C · Double scale graduated ml/fl.oz; Durably white printed graduation. | TTM |  | 9.25 | USD | 0 | General Cargo | 0.3 | 16 | 16 | 19 | 0.004864 | Witeg Labortechnik GmbH 5500002S | 70172000 | https://www.witeg.de/en/products/laboratory-glassware/beakers/low-form/low-form-with-spout-neutral?number=5500002S
106184 | Glass laboratory beaker 1000 ml | 1 unit | N/A | Laboratory beaker 1000 ml for general laboratory purpose, made of borosilicate glass 3.3, according to DIN 12331, ISO 3819 with low form with spout; Acid, alkali resistant; Autoclavable at 121°C · Double scale graduated ml/fl.oz; Durably white printed graduation. | TTM |  | 2.4 | USD | 0 | General Cargo | 0.3 | 13 | 13 | 15 | 0.002535 | Witeg Labortechnik GmbH 5500001S | 70172000 | https://www.witeg.de/en/products/laboratory-glassware/beakers/low-form/low-form-with-spout-neutral?number=5500002S
106185 | Glass laboratory beaker 600 ml | 1 unit | N/A | Laboratory beaker 600 ml for general laboratory purpose, made of borosilicate glass 3.3, according to DIN 12331, ISO 3819 with low form with spout; Acid, alkali resistant; Autoclavable at 121°C · Double scale graduated ml/fl.oz; Durably white printed graduation. | TTM |  | 3 | USD | 0 | General Cargo | 0.2 | 11 | 11 | 13 | 0.001573 | Witeg Labortechnik GmbH 5500600S | 70172000 | https://www.witeg.de/en/products/laboratory-glassware/beakers/low-form/low-form-with-spout-neutral?number=5500002S
106186 | Glass laboratory beaker 250 ml | 1 unit | N/A | Laboratory beaker 250 ml for general laboratory purpose, made of borosilicate glass 3.3, according to DIN 12331, ISO 3819 with low form with spout; Acid, alkali resistant; Autoclavable at 121°C · Double scale graduated ml/fl.oz; Durably white printed graduation. | TTM |  | 2.4 | USD | 0 | General Cargo | 0.2 | 9 | 9 | 10 | 0.00081 | Witeg Labortechnik GmbH 5500250S | 70172000 | https://www.witeg.de/en/products/laboratory-glassware/beakers/low-form/low-form-with-spout-neutral?number=5500002S
106187 | Glass laboratory beaker 100 ml | 1 unit | N/A | Laboratory beaker 100 ml for general laboratory purpose, made of borosilicate glass 3.3, according to DIN 12331, ISO 3819 with low form with spout; Acid, alkali resistant; Autoclavable at 121°C · Double scale graduated ml/fl.oz; Durably white printed graduation. | TTM |  | 2.35 | USD | 0 | General Cargo | 0.2 | 7 | 7 | 8 | 0.000392 | Witeg Labortechnik GmbH 5500100S | 70172000 | https://www.witeg.de/en/products/laboratory-glassware/beakers/low-form/low-form-with-spout-neutral?number=5500002S
4.2.4 Media preparation consumables
106112 | Pipette washer kit | 1 unit | N/A | Pipette washer, automatic siphon system, consisting of one cylindric washing container with siphon outlet connecting tube. Made of Polyethylene (PE) and HDPEF. Diameter 170mm x H 734 mm. For washing pipettes up to 600 mm; Water inlet and outlet separated to prevent cross contamination. This item includes: · Pipette washer; Two baskets with handle; Pipette jar and Inlet and outlet tubing. | TTM |  | 352.15 | USD | 0 | General Cargo | 35 | 17 | 17 | 99 | 0.028611 | Kartell SPA 217,2x219, 2x1219,218 plus 1 inlet and outlet tubing | 39269097 | https://www.kartelllabware.com/en/products/plastilab/pipettes-and-accessories/pipette-burette-rinsing-set-automatic/
106363 | Cleaning brush for test tubes | 5 units | N/A | Laboratory test tube brush to clean test tubes of 25 mm diameter and 300 mm length; High chemical resistance. Made of steel wire; Brush with cotton tip that allows to clean the bottom of the tubes. Suitable for round and flat bottom tubes; Clean without scratching the tubes. | TTM |  | 8.95 | USD | 0 | General Cargo | 0.1 | 32 | 7 | 2 | 0.000448 | Glaswarenfabrik Karl Hecht GmbH 42760030 | 96039099 | https://www.assistent.eu/en/product/test-tube-brushes-with-cotton-end/
106150 | Desiccator 5,8 lt set | 1 unit | N/A | Desiccator set for storage of dried samples in a dry atmosphere, 5,8 litres; Vacuum tight with a diameter of 200 mm; For use with maximum technically possible vacuum (0 mbar); Suitable for use with reusable or disposable desiccant; Vacuum and High chemical resistant; Autoclavable at least at 121°C. · Lid, base and stopcock made of DURAN; · Stopcock of 124 mm length, olive Ø 8 mm with PTFE spindle; · Porcelain plate of 185 mm diameter and withstand high temperatures up to 500°C; This item includes: Bottom part, lid with stopcock, porcelain plate, join grease and 10 desiccant cartridges. | TTM |  | 342.35 | USD | 0 | General Cargo | 3 | 30 | 30 | 40 | 0.036 | DWK Lifesciences ETP5.1 | 70172000 | https://www.carlroth.com/com/en/desiccators/desiccator-set-duran-novus-with-tube-%28ns-24-29%29-stopcock-in-the-lid/p/etp5.1
106746 | Laboratory multipurpose spatulas | 6 units | N/A | Laboratory multipurpose spatulas made of stainless steel 1.4301. Doubled ended spatulas: spatula/spatula, spatula/spoon, spatula/scoop to collect dry and liquid material. Autoclavable 121°C; Flexible flat blades with rounded edges; blade length of approx. 5 mm wide. This item includes 6 spatulas, with the below sizes: 1 x 80.034.150 Micro double spatula; 1 x 80.036.150 Micro double spatula; 1 x 80.625.150 Micro-spoon powder spatula; 1 x 80.037.150 Powder spatula; 1 x 80.627.150 Micro-spoon spatula; 1 x 80.025.150 Micro-spoon spatula. | TTM |  | 58.65 | USD | 0 | General Cargo | 0.6 | 20 | 20 | 10 | 0.004 | RSG Rostfrei-Schneidewerkzeuge GmbH 80.034.150, 80.036.150, 80.625.150, 80.037.150, 80.627.150, 80.025.150 & 80.066.906 | 90189010 | https://3mkknu2r000h11bipd1ydumd-wpengine.netdna-ssl.com/wp-content/uploads/2018/08/RSG-Lab.pdf
106415 | Square filter paper 580 mm | 100 units | N/A | Filter paper made of Cellulose; Sheets of 580 mm x 580 mm. Filtration: medium fast. · Basis weight: 85 g/m²; Average retention capacity: 4 - 12 µm; Filtration speed: 27 s; Thickness: 0,2 mm. | TTM |  | 111.4 | USD | 0 | General Cargo | 3.2 | 60 | 60 | 3 | 0.0108 | Macherey-Nagel GmbH MN 616 md (133010A) | 48054000 | https://www.mn-net.com/ch/filter-paper-sheets-mn-616-md-qualitative-medium-to-slow-55-s-smooth-133010a
106353 | Round filter paper 185 mm | 100 units | N/A | Filter paper made with a minimum alpha cellulose content of 98%; Filters of 185 mm. Filtration: medium fast. · Basis weight: 85 g/m²; Particle retention 7 µm; Filtration speed: 27 s; Thickness: 0,18 mm. | TTM |  | 34.5 | USD | 0 | General Cargo | 0.3 | 19 | 19 | 3 | 0.0011 | Cytiva (formerly GE) 10311814 | 48232000 | https://www.cytivalifesciences.com/en/us/shop/whatman-laboratory-filtration/speciality-cellulose-papers/seed-germination-testing-papers/grade-597-qualitative-filter-papers-p-00212
106354 | Round filter paper 150 mm | 100 units | N/A | Filter paper made with a minimum alpha cellulose content of 98%; Filters of 150 mm. Filtration: medium fast. · Basis weight: 85 g/m²; Particle retention 7 µm; Filtration speed: 27 s; Thickness: 0,18 mm. | TTM |  | 24 | USD | 0 | General Cargo | 0.2 | 16 | 16 | 3 | 0.0008 | Cytiva (formerly GE) 10311812 | 48232000 | https://www.cytivalifesciences.com/en/us/shop/whatman-laboratory-filtration/speciality-cellulose-papers/seed-germination-testing-papers/grade-597-qualitative-filter-papers-p-00212
106494 | Calibrated thermometer | 1 unit | N/A | High precision handheld thermometer; Power supply: 9 V-block battery (standard). Temperature switchable between °C and °F. Microprocessor controlled thermometer functions with single channel thermometer with integrated calibration and with built-in sensor holder. · With measurement input through PT 100 class B temperature probe; Water proof probe/sensor; · Pt100 sensor probe and thermometer measurement range of -50 to 350°C; PT100 input: 4-wire pin; · Resolution: 0,1°C; Accuracy ± 0,1°C; Memory for record measured values; · Battery life of at least 20 hours of continuous use; Able to import and export data results by USB; · With digital LCD display which shows temperature values and battery level. This item includes: 1 battery and 1 spare battery. | TTM |  | 328.8 | USD | 0 | General Cargo | 0.5 | 7 | 4 | 19 | 0.000532 | Carl Roth GmbH / Dostmann P700 | 90251900 | https://www.carlroth.com/de/de/thermometer-handmessgeraete/temperaturmessgeraet-p700-serie-p700/p/x717.1
106063 | Digital thermometer | 1 unit | N/A | Digital thermometer with 3 digits display with temperature modes of °C/°F and resolution of 0.1°C. · Resolution: up to 999.9: 0.1°C / 0.1°F, above 1°C / 1°F · Able to accept DIN connector probes; thermometer with a 4 level battery status indicator; · Able to storage last max and min temperature reads; Able to work with standard rechargeable batteries; · Measurement range G 1200: -65 to + 1.200 °C; Measurment range GTF 300: -65 to + 300 °C. This item includes: Flexible wire probe for thermometer made of NiCR-Ni type K; Cable length of 1000mm; Two (2) standard rechargeable batteries sets: one (1) for functioning + one (1) spare and One battery charger: power supply: input voltage range at least 100 to 240 V 50/60 Hz autoswitching. | TTM |  | 273.7 | USD | 0 | General Cargo | 0.2 | 13 | 9 | 4 | 0.000468 | GHM Messtechnik G1200 | 90251900 | https://www.greisinger.de/p/neuheiten/g-1200-ge/482458/
106084 | Rubber pipetting bulb | 1 unit | N/A | High suction safety ball/bulb pipetting aid, made of natural rubber; Inside valve easily removable for cleaning; Support all universal pipettes types; High chemical resistance; The rubber pipetting bulb has three outlets for: evacuation, suction of liquid, controlled dispensing. | TTM |  | 2.8 | USD | 0 | General Cargo | 0.8 | 20 | 20 | 20 | 0.008 | Deutsch & Neumann GmbH 4700002 | 40419000 | https://www.deutsch-neumann.de/en/pipette-filler
106086 | Pipette fillers set | 1 set | N/A | Set of mechanical pipette fillers, made of Polyethylene (PE). Syringe form; High chemical resistance; Color coded; Allows to be operated with one hand and also; Allows to be disassembled for cleaning; This item includes: 1 set include the following 3 pieces in different colors: 1 piece: size 0 - 0,02 ml; 1 piece: size 0 – 2 ml and 1 piece: size 0 - 10 ml. | TTM |  | 34.75 | USD | 0 | General Cargo | 0.5 | 18 | 6 | 8 | 0.0009 | Heinz Herenz Medizinalbedarf GmbH 1132005, 1132006, 1132007. | 39269097 | https://labteamet.com/wp-content/uploads/labteamet/Herenz/Produktkatalog_forbrukning_lab_medicin_HERENZ.pdf
5. INFECTION PREVENTION PRODUCTS
5.1 PERSONAL PROTECTIVE EQUIPMENT (PPE)
5.1.1 PPE consumables
106915 | Disposable respirators FFP2 | 20 units | N/A | Healthcare Respirator, FFP2, Type IIR with protection against hazardous airborne contaminates in medical environments. Single use with low resistance filter technology helps you breathe easily. Sculpted upper panel conforms to the nose and contours of the face with a soft foam for an effective seal · Approved to the European Medical Device Regulation (EU) 2016/425 and EN14683:2005 Type IIR for use as a surgical mask. · Chin tab designed for easy positioning and adjustment ·Outer cover provides resistance to fluid splashes | TTM |  | 40.1 | USD | 0 | General Cargo | 1 | 21 | 12 | 8.5 | 0.002142 | 3M Aura Health Care Particulate Respirators 1862+ | 63079093 | https://www.3m.co.uk/3M/en_GB/p/d/v000125637/ https://www.cavitydentalstaff.co.uk/app/uploads/2021/08/3m-aura-1863-technical-data-sheet.pdf
106818 | Latex gloves size S | 100 units | 30 months | Examination latex gloves, non sterile size S, are suited for protection against pathogenic germs in the laboratory. Powder-free; Color: white to light yellow; thickness: 0,1 mm; Length: 240 mm. | TTM |  | 19.65 | USD | 0 | General Cargo | 0.7 | 25 | 13 | 8 | 0.0026 | Nobamed Paul Danz AG 905451 | 40151200 | https://www.nobamed.com/gb/disposable-gloves/175-4239-nobaglove-latex.html
106819 | Latex gloves size M | 100 units | 30 months | Examination latex gloves, non sterile size M, are suited for protection against pathogenic germs in the laboratory. Powder-free; Color: white to light yellow; thickness: 0,1 mm; Length: 240 mm. | TTM |  | 19.65 | USD | 0 | General Cargo | 0.7 | 25 | 13 | 8 | 0.0026 | Nobamed Paul Danz AG 905452 | 40151200 | https://www.nobamed.com/gb/disposable-gloves/175-4239-nobaglove-latex.html
106820 | Latex gloves size L | 100 units | 30 months | Examination latex gloves, non sterile size L, are suited for protection against pathogenic germs in the laboratory. Powder-free; Color: white to light yellow; thickness: 0,1 mm; Length: 240 mm. | TTM |  | 19.65 | USD | 0 | General Cargo | 0.7 | 25 | 25 | 25 | 0.025625 | Nobamed Paul Danz AG 905453 | 40151200 | https://www.nobamed.com/gb/disposable-gloves/175-4239-nobaglove-latex.html
106723 | Disposable laboratory coat size S | 60 units | N/A | Disposable nonsterile laboratory white coats to be used in laboratories, size S.Provides an inherent barrier to particles, microorganisms, and non-hazardous light liquid splash. Bound seams are covered with garment fabric to reinforce the seam and to reduce the potential for particle penetration. · Standard collar type - with adjustable snap at neck for better fit. Knee length - Length: 99 - 114 cm long sleeves with elastic openings for tighter fit at wrist, with front snap closures for easy donning and doffing. · Includes 2 bags of 30 units each, for a total of 60 units | TTM |  | 501.25 | USD | 0 | General Cargo | 5 | 33 | 28 | 40 | 0.03696 | DuPont de Nemours GmbH IC 270B plus size | 62101098 | https://www.dupont.co.uk/products/tyvek-isoclean-ic0270bwhms.html
106724 | Disposable laboratory coat size M | 60 units | N/A | Disposable nonsterile laboratory white coats to be used in laboratories, size M. Provides an inherent barrier to particles, microorganisms, and non-hazardous light liquid splash. Bound seams are covered with garment fabric to reinforce the seam and to reduce the potential for particle penetration. · Standard collar type - with adjustable snap at neck for better fit. Knee length - Length: 99 - 114 cm long sleeves with elastic openings for tighter fit at wrist, with front snap closures for easy donning and doffing. · Includes 2 bags of 30 units each, for a total of 60 units | TTM |  | 501.25 | USD | 0 | General Cargo | 5 | 37 | 28 | 35 | 0.03626 | DuPont de Nemours GmbH IC 270B plus size | 62101098 | https://www.dupont.co.uk/products/tyvek-isoclean-ic0270bwhms.html
106725 | Disposable laboratory coat size L | 60 units | N/A | Disposable nonsterile laboratory white coats to be used in laboratories, size L. Provides an inherent barrier to particles, microorganisms, and non-hazardous light liquid splash. Bound seams are covered with garment fabric to reinforce the seam and to reduce the potential for particle penetration. · Standard collar type - with adjustable snap at neck for better fit. Knee length - Length: 99 - 114 cm long sleeves with elastic openings for tighter fit at wrist, with front snap closures for easy donning and doffing. · Includes 2 bags of 30 units each, for a total of 60 units | TTM |  | 501.25 | USD | 0 | General Cargo | 5 | 37 | 28 | 34 | 0.035224 | DuPont de Nemours GmbH IC 270B plus size | 62101098 | https://www.dupont.co.uk/products/tyvek-isoclean-ic0270bwhms.html
106435 | Sterile disposable surgical gowns size S | 20 units | 30 months | Disposable surgical gown with excellent wearing comfort. Material fibrous web, sterile, single use. Color: blue packed individually. Latex-free; Long sleeves with cuffs. Chinese supplier: Size: M (size M is comparable to size S) | TTM |  | 72.1 | USD | 0 | General Cargo | 5.36 | 39 | 30 | 37 | 0.04329 | Nobamed Paul Danz AG 621120 | 48185000 | https://www.nobamed.com/gb/surgical-gowns/268-3544-nobadress-comfort.html
106436 | Sterile disposable surgical gowns size M | 20 units | 30 months | Disposable surgical gown with excellent wearing comfort. Material fibrous web, sterile, single use. Color: blue packed individually. Latex-free; Long sleeves with cuffs. Chinese supplier: Size: L (size L is comparable to size M) | TTM |  | 78.8 | USD | 0 | General Cargo | 4.86 | 39 | 30 | 37 | 0.04329 | Nobamed Paul Danz AG 621130 | 48185000 | https://www.nobamed.com/gb/surgical-gowns/268-3544-nobadress-comfort.html
106437 | Sterile disposable surgical gowns size L | 20 units | 30 months | Disposable surgical gown with excellent wearing comfort. Material fibrous web, sterile, single use. Color: blue packed individually. Latex-free; Long sleeves with cuffs. Chinese supplier: Size: XL (size XL is comparable to size L) | TTM |  | 82.5 | USD | 0 | General Cargo | 5.5 | 39 | 30 | 37 | 0.04329 | Nobamed Paul Danz AG 621150 | 48185000 | https://www.nobamed.com/gb/surgical-gowns/268-3544-nobadress-comfort.html
106726 | Laboratory shoe covers | 2000 units | N/A | Blue shoe covers made of Polyethylene (PE); impermeable to liquids. · Latex -free; single use and non-Sterile: the shoe covers are worn over the shoes when entering the intensive or clean room area for a short time. · Anti-skid sole with strong rubber band. Includes 20 bags of 100 units each, for a total of 2000 units. | TTM |  | 213.25 | USD | 0 | General Cargo | 6.3 | 39 | 28 | 39 | 0.042588 | Nobamed Paul Danz AG 675101 | 39262000 | https://www.nobamed.com/gb/surgical-masks-caps-and-socks/274-3586-uberschuhe.html
106727 | Laboratory hair covers | 1000 units | N/A | White laboratory hair caps with elastic band made of polypropylene, elastomer thread. · Latex-free; single use and non-sterile; · With a diameter of 450 mm (52mm ·The caps are made of a spunbonded non-woven fabric which is air-permeable and therefore breathable. In this way a heat accumulation is prevented and the wearing comfort is increased | TTM |  | 61.45 | USD | 0 | General Cargo | 4.6 | 34 | 28 | 50 | 0.0476 | Nobamed Paul Danz AG 674101 | 39262000 | https://www.nobamed.com/gb/bibs-overshoes-care-cloth/387-4164-haube.html
5.2 BIOSAFETY PRODUCTS
5.2.1 Biosafety consumables
106079 | Thermic anemometer with telescope probe | 1 unit | N/A | Thermo anemometer with telescopic probe that measures velocity, temperature, humidity, differential pressure, barometric pressure. Specially used inside the biosafety cabinets. · Airflow velocity measurement range: 0 to 50 m/s; Humidity measurement range 0 to 95 % HR; · Accuracy for air flow measurement of ± 3%; Velocity units: m/s, km/h; With 1 temperature channel; · With telescopic probe up to 101,6 cm; Temperature measurement range of -10°C to +60°C; · Accuracy for temperature measurement of ± 0.3°C; Humidity measurement accuracy 3% HR; This item includes: Straight probes for temperature, humidity and air flow measurements: · 1 Telescope type probe;· 1 Telescope length of at least 100mm; 4 x non-rechargeable batteries and · 1 Probe ranges and accuracies not below the equipment ranges and accuracies; | TTM |  | 1866.35 | USD | 0 | General Cargo | 4 | 55 | 46 | 22 | 0.05566 | Driesen+Kern GmbH (TSI INC.) TSI9565-X + TSI964 | 90268020 | https://www.driesen-kern.de/produkte/luftgeschwindigkeit-volumenstrom/handmessgeraete/multifunktions-klimamessgeraet-tsi9565-x.php
106711 | Germicidal UV-C for disinfection, portable | 1 unit | N/A | Germicidal UV-C device for disinfection by direct irradiation, portable on wheels. Selective UV-C lamp (at 253.7 nm.) with high efficiency lighting with remote control (on/off), motion sensor (for more safety), timer, emergency stop button and visual control lamp (built in an external box), without microprocessor. · All materials are tested for resistance to intense UV-C Rays. UV-C destroys airborne microbes, bacteria, mold and virus which are exposed by the light; Provide a minimum UV-C irradiance of 40 µW/cm2 ; · UV-C lamp working up to 18000 hours; The unit is on wheels. · Power supply: 230 V 50/60Hz (110 V on request) | TTM |  | 2696.8 | USD | 0 | General Cargo | 1.8 | 47 | 15 | 12 | 0.00846 | Light Progress S.r.L. UV-STICK-E2/75H-ST-NX-NT (Light progress) + external switch-box (TTM) | 85394910 | https://health.lightprogress.it/en/prodotti/uv-stick-nx-al-uv-stick-nx-st-1_45.html
106745 | Airflow test device set | 1 unit | N/A | Airflow test device check air, current indicator produces a harmless cloud of smoke that floats freely and easily, because it has the same density as ambient air. Consequently, slight air currents become visible to easily identify leaks and test the integrity of air-tight systems. · Airflow test device generates non-corrosive non-toxic white smoke to visualize airflow movements; · Odorless smoke; oil-free smoke; · Refill cartridges made of clear plastic; each cartridge generates smoke up to approx. 3 minutes. This item includes: · 3 cartridges (no sealing caps needed); 1 Ni-Cd rechargeable battery (7,2 V, 1,4 Ah) and 1 carrying case. | TTM |  | 651.45 | USD | 0 | General Cargo | 0.5 | 6.5 | 14.5 | 0.5 | 4.7125e-05 | Dräger Safety AG 6400761, 8316993 | 84799070 | https://www.draeger.com/Products/Content/flow-check-pi-9046525-en-gb.pdf
106722 | Refill for Airflow test device | 12 units | 16 months | Refill cartridges made of clear plastic. Each cartridge generates smoke up to approx. 3 minutes (continuous use). · White smoke generated by the cartridges to visualize airflow is non-corrosive and non-toxic This item includes: · 4 bags with 3 refill cartridges in each bag | TTM |  | 148.75 | USD | 0 | General Cargo | 0.24 | 28 | 29 | 32 | 0.025984 | Dräger Safety AG 6400812 | 84799070 | https://www.draeger.com/Products/Content/flow-check-pi-9046525-en-gb.pdf
106094 | First aid kit for laboratories | 1 unit | N/A | First Aid Kit for laboratories, with additional eye rinsing bottle, wall mounted, lockable cupboard, door with sealing gasket. This item includes: First Aid Case, designed for mobile and stationary use and 141 items: adhesive tapes, adhesive plasters, finger tip adhesive plasters; wound dressing; dressing sachets in different sizes; surgical sheets, gauze bandages in different sizes; fixation bandages; compress; eye pads or eye compress; instant cold packs; rescue covers or thermal emergency blankets metallized polyester foil; triangular bandages; scissors; disposable gloves and PE-bags. · Plum Emergency eyewash wall station DUO: Dust prof wall station containing: 1 bottle ph Neutral DUO (500ml) and 1 bottle eye wash DUO with pictogram and mirror. Suitable for work areas where there is a risk of acids, alkali or foreign matter getting into the eye. Easy to open in an emergency. | TTM |  | 248.85 | USD | 0 | General Cargo | 2 | 27 | 12 | 7 | 0.002268 | SÖHNGEN DIN 13169 (Filling: Industry norm); 0301155, 0303008 and 4862 | 30059099 | https://en.soehngen.com/products/first-aid/first-aid-equipment-with-din-13169/
106349 | Tripod stand for plastic bags of 2 lt | 1 unit | N/A | Tri-pod stand for laboratory waste bags of 2 liters. Made of epoxy-coated steel wire. Sturdy tripod-design, very stable with rubber feet. | TTM |  | 8.1 | USD | 0 | General Cargo | 0.3 | 20 | 20 | 25 | 0.01 | Ratiolab GmbH 7001400 | 39269097 | http://ratiolab.com/en/autoclavable-bags-dispenser-box/787-table-stand-height-25-cm-for-waste-disposal-bags-200-x-300-mm-73269098.html
106728 | Biohazard disposal bags 2.5 lt | 100 units | N/A | Red biohazard disposal bags with warning label/sterilization indicator printed. Also printed with precautionary procedures in four languages: English, French, German and Spanish. Made of Polypropylene. Capacity: approx. 2,5 lt. Extra heavy-duty with thickness of 2 ml · Liquid tight seal with heat resistant up to135°C; Supplied in roll with perforated lip for easy tear off · Extra thick, Extra Strong High Temperature Bags Reduce Autoclave Time. Autoclavable to 135°C. Feature a sterilization indicator patch that darkens when exposed to steam sterilization for 20 minutes | TTM |  | 97.3 | USD | 0 | General Cargo | 0.5 | 23 | 12 | 9 | 0.002484 | H-B Instrument / SP Scienceware (BelArt) F13165-1419 | 39239000 | https://www.belart.com/bel-art-f13165-1419-polypropylene-super-strength-red-biohazard-disposal-bags-with-warning-label-and-sterilization-indicator-patch-2-4-gallon-14-width-x-19-height-x-2-0-mil-thick-pack-of-200.html
106445 | Stand for plastic bags of 30 lt | 1 unit | N/A | Stand for plastic bags of 30 lt. It keeps bags open and securely supported on these lightweight, stainless steel wire frames. Polypropylene bottom serves as a containment tray in the event of bag leakage and is safety orange in color for easy identification as a biohazard containment system. Heat resistant up to 121 °C, chemical resistant. Dishwasher-safe; autoclavable at 121°C (250°F). | TTM |  | 55.1 | USD | 0 | General Cargo | 2 | 82 | 37 | 6 | 0.018204 | H-B Instrument / SP Scienceware (BelArt) F13192-0002 | 39232990 | https://www.belart.com/bel-art-f13192-0002-clavies-biohazard-bag-holder-14-x-14-x-21-5-holds-24-x-30-bags.html
106409 | Sharps containers with lids 2 lt | 10 units | N/A | Yellow sharps disposable container with lids made of polypropylene 2 lt. · Sharp container with temporary and permanent closure (permanent closure prevents waste leakage) · Practical opening system with wide opening for the disposal of larger objects · Needles removal and printed with biohazard symbols. | TTM |  | 11.45 | USD | 0 | General Cargo | 0.1 | 15.5 | 15.5 | 19.5 | 0.0046849 | Apmedical DISPO 2 | 39239000 | https://www.apmedical.it/eng/sharps-containers/dispo-line.html
5.2.2 Disinfectants
106821 | Chlorine disinfectant | 100 units | 23 months | For simultaneous disinfecting in the laboratory and medical field. Based on NaDCC (sodium dichloroisocyanurate). · Weight of tablet 6 g. Recommended active chlorine content for acinetobacter bactericidal, (Clostridium difficile (also spores), E. Coli, Enterococcus, Klebsiella, Pseudomonas and Staphylococcus (MRSA) (tested acc. to EN 1276). · Contains 1000 ppm - this corresponds to a dosage of one tablet in 1 litre water. | TTM |  | 43.2 | USD | 0 | Dangerous cargo UN 3077, Class 9, PG III | 2 | 21 | 29 | 32 | 0.019488 | Medentech/Kersia Klorspet 87 | 28011000 | http://www.medentech.com/index.php?id=759
106373 | Liquid soap for laboratory, 1 lt | 1 unit | 23 months | Liquid soap, mild, fragrance- and colorant free washing lotion with good skin tolerance for normal skin. cleanses with a mild formula. · Skin neutral pH-value: Skin-neutral pH value 5,5; Contains skin-mild; · Suitable also for sensitive skin; Free of fragrances and colorants; Free of parabens | TTM |  | 3.9 | USD | 0 | General Cargo | 1.1 | 21 | 29 | 32 | 0.0195 | Bode Chemie GmbH/Paul Hartmann 981329 | 34012090 | https://productcatalogue.bode-chemie.com/products/hands/baktolin_pure.php
106374 | Disinfectant for hands, 1 lt | 1 unit | 16 months | Disinfectant for hands 1lt bottle. Comprehensively effective against bacteria, yeasts and enveloped viruses as well as noro-, adeno- and rotavirus particularly skin-friendly and moisturizing. · Increases skin moisture, maintains skin elasticity and, with regular use, prevents rough and chapped skin1); excellent tolerance even with long-term use; · Maintains the skin's natural pH; excellent immediate effect; Free of dyes and fragrances. | TTM |  | 8.2 | USD | 0 | Dangerous cargo UN 1987, Class 3, PG III | 1.4 | 21 | 29 | 32 | 0.0195 | Bode Chemie GmbH/Paul Hartmann 980391 | 38089490 | https://productcatalogue.bode-chemie.com/products/hands/sterillium_classic_pure.php
106294 | Cleaner and disinfectant for instruments, 2 lt | 1 unit | 16 months | Aldehyde-free instrument disinfectant with excellent cleaning power with excellent cleaning performance. Compatible with preparations based on other active ingredients High material compatibility; pleasant smell; low use concentrations ensure high level of efficiency. | TTM |  | 46.55 | USD | 0 | Dangerous cargo UN 1903, Class 8, PG III | 3.6 | 21 | 29 | 32 | 0.0195 | Bode Chemie GmbH/Paul Hartmann 980320 | 38089490 | https://productcatalogue.bode-chemie.com/products/instruments/bomix_plus.php
106371 | Disinfectant for alcohol-resistant surfaces, 5 lt | 1 unit | 16 months | Bacillol AF, Alcohol-based rapid disinfectant for disinfecting alcohol-resistant surfaces, with extensive spectrum of activity. Suitable for spray-wipe disinfection method. pH 6 (20°C) and free from aldehyde. This item includes: a bottle of 5 lt. | TTM |  | 35.2 | USD | 0 | Dangerous cargo UN 1987, Class 3, PG III | 5.1 | 21 | 29 | 32 | 0.0195 | Bode Chemie GmbH/Paul Hartmann 980217. Bacillol AF, 5 l | 38089490 | https://productcatalogue.bode-chemie.com/products/surface/bacillol_af.php
106370 | Disinfectant for floors, 2 lt | 1 unit | 23 months | Disinfectant for floors, liquid, aldehyde-free and alcohol-free concentrate for disinfecting alcohol-resistant surfaces of medical products, especially device surfaces and objects in the vicinity of patients.. · Alcohol-free concentrate; Aldehyde free: Effectiveness: Bactericidal, yeasticidal; Tuberculocide effective; Limited virucidal PLUS effective; High material compatibility. | TTM |  | 17 | USD | 0 | Dangerous cargo UN 1903, Class 8, PG III | 5.2 | 19 | 15 | 25.5 | 0.0072675 | Meditrade GmbH 00972D. Medizid AF | 38089410 | https://www.meditrade.ch/medizid-af-desinfektionsm-2-ll---flasche-00972d-p.html
106747 | Spray dispenser bottles 1 lt | 2 units | N/A | Pressure less spray bottle made of LDPE with long-life hand pump with internal flexible tube. · Chemical resistant and long-life hand pump, media-contacting parts PP, PE and stainless steel · Nozzle 0.6 mm diameter, nozzle steplessly adjustable · Spray volume per stroke: 1.2 ml ± 0.1 with a nozzle diameter of 0,6 mm and hand-operated spray trigger · Pull-trigger with lockable function; stepless adjustable nozzle This item includes 2 bottles with 1 spray head per bottle | TTM |  | 10.3 | USD | 0 | General Cargo | 0.2 | 20 | 12 | 25 | 0.006 | Bürkle GmbH 0309-1100 | 39269097 | https://www.buerkle.de/en/spray-bottle#similar
106378 | Detergent liquid, 9.1 lt | 1 unit | 16 months | Liquid, highly alkaline cleaning concentrate for automated cleaning of laboratory glassware und reusable laboratory material. Rinsable without residues and guarantees therefore consistent and error-free results. With a broad cleaning spectrum and excellent results even at low concentrations. · High cleaning performance due to excellent dispersing capabilities, effective against stubborn deposits · Phosphate- and surfactant-free, no release of phosphates and surfactants into waste water. | TTM |  | 130.8 | USD | 0 | Dangerous cargo UN 1814, Class 8, PG II | 18 | 105 | 60 | 32 | 0.2016 | Borer Chemie AG DECONEX 22 HPF-x | 34029090 | https://www.borer.ch/pdf_files.php?type=prodinfo&file=pdf_files/Laboratory/english/Product%20datasheets/deconex_22_HPF-x_EN.pdf
5.3 WASTE MANAGEMENT
5.3.1 Waste management reagents
106331 | Phenol, 1 kg | 1 unit | 12 months | Phenol for laboratory use. Empirical formula C6H6O. CAS No 108-95-2. · Molar mass 94,11 g/mol; Density 1,07 g/cm³; Boiling point 182 °C; · Flash point 81 °C; Melting point 42 °C. Subject to availability from our supplier. Product lead-time may be longer than usual. Please contact your GDF Country Supply Officer for further information | TTM |  | 63.2 | USD | 0 | Dangerous cargo UN 1671, Class 6.1, PG II | 4 | 45 | 45 | 40 | 0.081 | Carl Roth GmbH 0040.2 | 29071100 | https://www.carlroth.com/com/en/determination-of-viscosity/phenol/p/0040.2
5.3.2 Waste management consumables
106253 | Safety box for liquids, 2 lt | 30 units | N/A | Safety box for liquid. Round and transparent disposable container made of Polyethylen (PE). Volume 2 Liter. Leakage proof with large opening with screw cap; With labelling. | TTM |  | 58.45 | USD | 0 | General Cargo | 4.3 | 61 | 37 | 52 | 0.11736 | Sarstedt AG 77.1515 | 39269097 | https://www.sarstedt.com/en/products/diagnostic/multi-safe/boxes/product/dmsarproducts/specificationPdf/Product/77.1515/
106718 | Bench-top biohazard waste container for solids, 10 lt | 12 units | N/A | Bench-top biohazard waste container for solids, 10 lt with inner biohazard plastic bag. . Printed with large, bright, biohazard graphics; · Suitable for closing and incineration after fill up the container; · When the box is full, pull the flap to a closed position and the entire unit is ready for incineration; · Packaged flat to save space and shipping costs, the carton and inner bag can be assembled in seconds. | TTM |  | 117.6 | USD | 0 | General Cargo | 2.1 | 51 | 43 | 11 | 0.024123 | Scienceware (Bel-Art) 132050002 | 48191000 | https://www.belart.com/bel-art-f13205-0002-cardboard-biohazard-incinerator-cartons-8-x-8-x-10-benchtop-model-pack-of-6.html
106256 | Waste container for liquids, 4 lt | 1 unit | N/A | Safety Waste System for Laboratory Liquid Chemical & Biological Waste, made of HDPE. For acid, base and alcohol waste. Capacity 4 liters. Translucent container and leak proof with screw cap with seal. This Item includes: One (1) removable funnel: Made of Polypropylene with diameter of funnel 140 mm. External ribbing or alternative system to prevent air lock. Closure/adapter includes vent with polypropylene/PTFE vent plug that releases pressure during pouring, and prevents emissions during use. | TTM |  | 71.6 | USD | 0 | General Cargo | 0.6 | 15 | 15 | 30 | 0.00675 | Nalge Nunc International Corp. 6379-0004 | 39269097 | https://www.thermofisher.com/order/catalog/product/6379-0004#/6379-0004
6. LABORATORY EQUIPMENT
6.1 UPS AND SOLAR PANELS
6.1.1 UPSs
106707 | UPS 700 VA with battery pack | 1 unit | N/A | On-Line UPS designed with double conversion technology. Capacity of 700 VA with pure sinewave voltage output, with protection for overcharge, with short circuit and with emergency power off (EPO); and with USB and RS 232; 4x IEC 320 C 13 output sockets; · 120 V Version 60/70/80 VAC - 144 VAC or 230 V Version 120/140/160 VAC - 276 VAC, 50/60 Hz frequency in/out with autosensing. · Audible alarms: low battery, working in backup mode, overcharge and faults · Product with 2 years warranty; maintenance free and sealed gel batteries This item includes 1 UPS and 1 battery pack | TTM |  | 761.15 | USD | 0 | General cargo UN 2800 Valve Regulated Lead Acid Battery: UN 2800 Class 8 Packing group not restricted as per special provision A48 / A67 | 15.3 | 55 | 28 | 34 | 0.05236 | Powercom Co. for Multimatic Edelstrom GmbH VGD-700 | 85044030 | https://www.upspowercom.com/UPS-VGD.jsp
106708 | Battery pack set for UPS 700 VA | 5 units | N/A | Battery pack for UPS 700 VA, to amplify the existing back up time of the UPS for up to 150 min duration with full load; with interface to connect the UPS via polarized terminals · Built-in battery charger; · Auto select 50/60Hz 220/230/240v +/- 2 % Sinus, with UPS software. · Product with 2 years warranty; maintenance free and sealed gel batteries This item included 5 x Battery pack | TTM |  | 1749.3 | USD | 0 | General cargo UN 2800 Valve Regulated Lead Acid Battery: UN 2800 Class 8 Packing group not restricted as per special provision A48 / A67 | 14 | 24 | 16 | 42 | 0.016128 | Powercom Co. for Multimatic Edelstrom GmbH BEM 1/4 | 85044030 | https://www.upspowercom.com/UPS-VGD.jsp
106709 | UPS 3000 VA with battery pack | 1 unit | N/A | On-Line UPS designed with pure sinewave voltage output. Capacity of 3000 VA with protection for overcharge and short circuit, with digital display for: In/out voltage and frequency, output power, output current, time left, load level, faults (shown on LCD display) - working mode shown by LED lights. · 120V version 60/70/80VAC-144 VAC or 230V version 120/140/160 VAC-276 VAC; · LAN ports protection with 1 network transient protector (2x RJ 45); · 8x IEC 320 C13 output sockets with interface communication ports: USB and RS 232; · Audible alarms: low battery, working in backup mode, overcharge and faults; · Product with 2 years warranty; maintenance free and sealed gel batteries · This item includes 1 UPS and 1 battery pack | TTM |  | 1933.55 | USD | 0 | General cargo UN 2800 Valve Regulated Lead Acid Battery: UN 2800 Class 8 Packing group not restricted as per special provision A48 / A67 | 34.6 | 52.6 | 34 | 48.5 | 0.086737 | Powercom Co. for Multimatic Edelstrom GmbH VGD-3000 | 85044030 | https://www.upspowercom.com/UPS-VGD.jsp
106710 | Battery pack set for UPS 3000 VA | 3 units | N/A | Battery pack set for UPS 3000 VA, to amplify the existing back up time of the UPS for up to 70 minutes with 3x battery packs; with interface to connect the UPS via polarized terminals · Built-in battery charger; ·Auto select 50/60Hz 220/230/240v +/- 2 % Sinus, with UPS software. · Product with 2 years warranty; maintenance free and sealed gel batteries · This item includes 3 battery pack | TTM |  | 2364.8 | USD | 0 | General cargo UN 2800 Valve Regulated Lead Acid Battery: UN 2800 Class 8 Packing group not restricted as per special provision A48 / A67 | 44.5 | 52.6 | 34 | 48.5 | 0.086737 | Powercom Co. for Multimatic Edelstrom GmbH BEM 2/2 | 85044030 | https://www.upspowercom.com/UPS-VGD.jsp
6.1.2 Solar panels
106493 | Solar panel generator system 750 W | 1 unit | N/A | Solar panel generator to power laboratory equipment, compatible with GeneXpert-4 module. Use per day of at least eight (8) hours of continuous operation. This item includes: · 4 photovoltaic panels: Monocrystalline silicon solar cells; Maximum voltage of 1000 VDC; Nominal power of 200 Wp; Number of cells: 72 (6x12); Panel efficiency of 13%; Operating temp range of -40°C to 85°C; Junction box IP67 with 3 bypass diodes and nominal power output 90% for at least 10 years; · Solid state photovoltaic charge controller: Automatic voltage selection 12 or 24 V; Operation temp range of -35°C to 60°C; Self-consumption < 10mA; Maximum load: 30A; With digital display; Battery overcharge protection; Circuit protection against overvoltage and undervoltage; · Battery bank: consisting of 8 batteries SB 6/330A; High Quality VRLA Solar Battery; Capacity 330 Ah; Voltage of 6 V and battery bank: 24 V; Excellent high rate charge/discharge; Max. depth of discharge to 70% of rated capacity; Life cycle at 30% capacity are > 3000 cycles that means more that 8 years; · Inverter: Pure sine wave output switchable 100 -120 VAC or 200 -240 VAC; Output frequency 50/60; Efficiency 94 %; Controlled by microprocessor; Maximum output p+E50ower of at least 1.5 kW (at 25°C); Input voltage (DC): 12 or 24 V; Output voltage (AC): 120 V or 220 V +/- 3%; · All needed cables and connectors and support stand. | TTM |  | 5375.65 | USD | 0 | General Cargo | 550 | 170 | 80 | 120 | 1.632 | TTM Technologie Transfer Marburg e.V. / Victron AT106493 special production | 85023980 | https://www.victronenergy.com/solar-pv-panels/bluesolar-panels
6.2 AUTOCLAVES
6.2.1 Horizontal Autoclaves
106750 | Horizontal benchtop Autoclave 90lt | 1 unit | N/A | Horizontal benchtop autoclave 90lt. Integrated, separate steam generator with housing, support frame and pressure vessel made of corrosion-resistant stainless steel; • Temperature and pressure range 140°C, 4 bar (except for Systec D-23); • LCD display and fully automatic microprocessor control; number of sterilization programs: Up to 25; • Code-secured access rights for changing parameters and further safety-relevant intervention; • Internal memory for storing up to 500 sterilization cycles; timer for starting programs; • Autofill: automatic demineralized water feed for steam generation; flexible PT-100 temperature sensor. • Additional temperature sensor in condense exhaust. • RS-232 and RS-485 interfaces for external data transmission (network-compatible). Please add to your request the UPS 3000 VA with its battery pack, the accessories (Air exhaust filters, Bucket waste sterilization, Wire-mesh basket printer) and the installation and training. The air compressor is an optional item to have a rapid re-cooling, for which support pressure is required. | TTM |  | 22181.15 | USD | 0 | General Cargo | 196 | 78 | 94 | 65 | 0.47658 | Systec, DX90 (2090, 9910, 9603, 1042, 2046, 8030) | 84192000 | https://www.systec-lab.de/wp-content/uploads/2021/03/12_2020%20Systec%20Laboratory%20Autoclaves%20V-%20and%20D-Series.pdf
106751 | Air exhaust filters for horizontal benchtop Autoclave 90lt | 2 units | N/A | For the sterilization of hazardous biological substances, Systec autoclaves can be fitted with an optional air exhaust filtration system. The autoclavable sterile filter, consisting of a filter cartridge with PTFE membrane of pore size 0.2 µm, is installed in a pressure-proof housing and can be quickly changed at any time. · The filter is also automatically sterilized inline during the sterilization process, monitored by the PT-100 temperature sensor. · The condensate is retained inside the pressure vessel during the heating and sterilization phases and thus also sterilized. Through air exhaust filtration and condensate inactivation, it is ensured that no microorganisms can escape before end of the sterilization phase. | TTM |  | 259.9 | USD | 0 | General Cargo | 15 | 70 | 50 | 80 | 0.28 | Systec, 10611 | 84202000 | https://www.systec-lab.de/wp-content/uploads/2021/03/12_2020%20Systec%20Laboratory%20Autoclaves%20V-%20and%20D-Series.pdf
106752 | Bucket waste sterilization for horizontal benchtop Autoclave 90l | 1 unit | N/A | Bucket for waste sterilization made of stainless steel for the horizontal benchtop autoclave 90lt. • Perforated with top handle; Stackable buckets. | TTM |  | 838.25 | USD | 0 | General Cargo | 1.5 | 10 | 30 | 40 | 0.012 | Systec GmbH 5075, 9005 | 84212000 | https://www.systec-lab.de/wp-content/uploads/2021/03/12_2020%20Systec%20Laboratory%20Autoclaves%20V-%20and%20D-Series.pdf
106753 | Wire-mesh basket for horizontal benchtop Autoclave 90lt | 1 unit | N/A | Perforated basket for sterilization made in stainless steel with top handle for the horizontal benchtop autoclave 90lt. | TTM |  | 549.8 | USD | 0 | General Cargo | 3 | 40 | 10 | 50 | 0.02 | Systec GmbH 5074 | 84222000 | https://www.systec-lab.de/wp-content/uploads/2021/03/12_2020%20Systec%20Laboratory%20Autoclaves%20V-%20and%20D-Series.pdf
106754 | Printer for horizontal benchtop Autoclave 90lt | 1 unit | N/A | Inbuilt thermal dot printer. Printing in paper roll for the horizontal benchtop autoclave 90lt. • The paper roll is 58 mm wide and has a diameter of 35 mm. At the start of a program the date, time and the kind of program with the set parameters will be documented. • Out of the sterilization phase the pressure, temperature, program phase and runtime will be printed in intervals of 2 minutes. This item includes: 20 paper rolls (can be exchanged from the front part of the printer) | TTM |  | 1259.5 | USD | 0 | General Cargo | 2 | 20 | 20 | 10 | 0.004 | Systec GmbH 8011, 8053 | 84232000 | https://www.systec-lab.de/wp-content/uploads/2021/03/12_2020%20Systec%20Laboratory%20Autoclaves%20V-%20and%20D-Series.pdf
106490 | Installation and Training for Autoclave | 1 unit | N/A | The installation of the autoclave is done in the laboratory with pre-installation work as plumbing and electricity. Also training on use and troubleshooting for the autoclave. | TTM |  | 4500 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | N/A | N/A
106808 | Horizontal benchtop Autoclave 55lt | 1 unit | N/A | Horizontal benchtop Autoclave 55lt with with front door and rapid cooling, support pressure and fan inside, for open and closed vessels for the avoidance temperature gradients in the chamber, with integrated vacuum device for solid substances. • Pressure max: 2.8 bar; Height or depth of vessel: 460 mm; Steam generator with automatic water fill; • Temperature max.: Adjustable temperature working range from 94° up to at least 137°C; 12 pre-defined programs and plus 10 user defined programs with LCD display; • Water reservoir of 16 Liter, consumption per cycle: 3,5 - 8 Liter; Safety systems for: Thermostat, Pressure switch, Valves; RS485 for data interface. Please add to your request the UPS 3000 VA with its battery pack, the accessories (Air exhaust filters, Bucket waste sterilization, Wire-mesh basket printer) and the installation and training. The air compressor is an optional item to have a rapid re-cooling, for which support pressure is required. | TTM |  | 22074.95 | USD | 0 | General Cargo | 196 | 78 | 94 | 65 | 0.47658 | SHP Laboklav, 55-MSLV | 84192000 | https://www.shp-steriltechnik.de/resources/PDF/Catalogue-Laboklav.pdf
106809 | Air exhaust filters for horizontal benchtop Autoclave 55lt | 1 unit | N/A | For the sterilization of hazardous biological substances, Systec autoclaves can be fitted with an optional air exhaust filtration system. The autoclavable sterile filter, consisting of a filter cartridge with PTFE membrane of pore size 0.2 µm, is installed in a pressure-proof housing and can be quickly changed at any time. · The filter is also automatically sterilized inline during the sterilization process, monitored by the PT-100 temperature sensor. · The condensate is retained inside the pressure vessel during the heating and sterilization phases and thus also sterilized. Through air exhaust filtration and condensate inactivation, it is ensured that no microorganisms can escape before end of the sterilization phase. | TTM |  | 723.9 | USD | 0 | General Cargo | 15 | 70 | 50 | 80 | 0.28 | Exhaust air filters for benchtop autoclave for SHP Laboklav, 55-MSLV. SHP 50.05.10015.012 | 84199015 | https://www.shp-steriltechnik.de/resources/PDF/Catalogue-Laboklav.pdf
106810 | Bucket waste sterilization for horizontal benchtop Autoclave 55l | 1 unit | N/A | Bucket for waste sterilization made of stainless steel for the horizontal benchtop autoclave 55lt. • Perforated with top handle; Stackable buckets. | TTM |  | 530.6 | USD | 0 | General Cargo | 1.5 | 10 | 30 | 40 | 0.012 | Bucket for sterilization for benchtop autoclave SPH Laboklav, 55-MSLV.SHP 30.1290.145.450 | 84199015 | https://www.shp-steriltechnik.de/resources/PDF/Catalogue-Laboklav.pdf
106811 | Wire-mesh basket for horizontal benchtop Autoclave 55lt | 1 unit | N/A | Perforated basket for sterilization made in stainless steel with top handle for the horizontal benchtop autoclave 55lt. | TTM |  | 513.1 | USD | 0 | General Cargo | 3 | 40 | 10 | 50 | 0.02 | Basket for sterilization for benchtop autoclave SPH Laboklav, 55-MSLV. | 84199015 | https://www.shp-steriltechnik.de/resources/PDF/Catalogue-Laboklav.pdf
106812 | Printer for horizontal benchtop Autoclave 55lt | 1 unit | N/A | Inbuilt thermal dot printer. Printing in paper roll for the horizontal benchtop autoclave 55lt. • The paper roll is 58 mm wide and has a diameter of 35 mm. At the start of a program the date, time and the kind of program with the set parameters will be documented. • Out of the sterilization phase the pressure, temperature, program phase and runtime will be printed in intervals of 2 minutes. This item includes: 10 printer paper rolls and ribbon (5x) | TTM |  | 1077.05 | USD | 0 | General Cargo | 2 | 20 | 20 | 10 | 0.004 | SHP Laboklav 55-MSLV Printer for Autoclave for horizonatl benchtop autoclave 55 litres instead of Systec GmH 8011, 8053 | 84199015 | https://www.shp-steriltechnik.de/resources/PDF/Catalogue-Laboklav.pdf
106490 | Installation and Training for Autoclave | 1 unit | N/A | The installation of the autoclave is done in the laboratory with pre-installation work as plumbing and electricity. Also training on use and troubleshooting for the autoclave. | TTM |  | 4500 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | N/A | N/A
6.2.2 Vertical Autoclaves
106813 | Vertical floor Autoclave 80lt | 1 unit | N/A | Vertical floor Autoclave 80lt top loading unit with rapid cooling, support pressure and fan inside, for open and closed vessels for the avoidance temperature gradients in the chamber, with integrated vacuum device for solid substances. - Solid sterilization and waste sterilization at 134°C. Water reservoir 30 Liter. - Max. permissible pressure: 2.8bar; Sterilizing temperature: 98°C to 138°C; - Motor driven and self-checking closing mechanism; Temp. measure precision not greater than ±2%; - Flexible temperature reference sensor PT100 as serial function; The lid can optionally be opened to the right, to the left or backwards. Adjustable timer range not smaller than 10 min to 60 min; - Fast re-cooling system for liquid sterilisation using a double jacket; - Adjustable temperature working range not smaller than from 98° up to at least 137°C; - 12 pre-defined programs and plus 10 user defined programs with LCD Display. Please add to your request the UPS 3000 VA with its battery pack, the accessories (Air exhaust filters, Bucket waste sterilization, Wire-mesh basket printer) and the installation and training. The air compressor is an optional item to have a rapid re-cooling, for which support pressure is required. | TTM |  | 28770.65 | USD | 0 | General Cargo | 150 | 100 | 80 | 100 | 0.8 | SHP Laboklav 80-MSLV | 84192000 | https://www.shp-steriltechnik.de/resources/PDF/Catalogue-Laboklav.pdf
106814 | Air exhaust filters for Vertical floor Autoclave 80lt | 1 unit | N/A | Filter cartridge air exhaust filter. The autoclavable sterile filter, consisting of a filter cartridge with PTFE membrane of pore size 0.2 µm, is installed in a pressure-proof housing. The filter is also automatically sterilized in line during the sterilization process, monitored by the PT-100 temperature sensor. · Filter removal capability to particles sizes of at less of 0.01 or less microns in gas, and at less than of 0.2 microns or less in liquid; Hydrophobic filter; Built in the autoclave. · Filter life: 12 months and 200 sterilization cycles. This item includes: One installed and one spare filter to be supplied with the autoclave; | TTM |  | 723.9 | USD | 0 | General Cargo | 15 | 70 | 50 | 80 | 0.28 | SHP Laboklav 80-MSLV Exhaust air filters for vertical autoclave | 84199015 | https://www.shp-steriltechnik.de/resources/PDF/Catalogue-Laboklav.pdf
106815 | Bucket waste sterilization for Vertical floor Autoclave 80l | 1 unit | N/A | Bucket for waste sterilization made of stainless steel for the Vertical floor Autoclave 80l • Perforated with top handle; Stackable buckets; Diameter 400 x 300 mm. | TTM |  | 845.2 | USD | 0 | General Cargo | 1.5 | 10 | 30 | 40 | 0.012 | SHP Laboklav 80-MSLV Bucket for sterilization for vertical autoclave 80 litres | 84199015 | https://www.shp-steriltechnik.de/resources/PDF/Catalogue-Laboklav.pdf
106816 | Wire-mesh basket for Vertical floor Autoclave 80lt | 1 unit | N/A | Perforated basket for sterilization made in stainless steel with top handle for the Vertical floor Autoclave 80l. Diameter of: 1/2 80, 400 x 300 mm | TTM |  | 553.2 | USD | 0 | General Cargo | 3 | 40 | 10 | 50 | 0.02 | SHP Laboklav 80-MSLV Baskets for sterilization for vertical autoclave 80 litres | 84199015 | https://www.shp-steriltechnik.de/resources/PDF/Catalogue-Laboklav.pdf
106817 | Printer for Vertical floor Autoclave 80lt | 1 unit | N/A | Inbuilt thermal dot printer. Printing in paper roll for the Vertical floor Autoclave 80l. • The paper roll is 58 mm wide and has a diameter of 35 mm. At the start of a program the date, time and the kind of program with the set parameters will be documented. • Out of the sterilization phase the pressure, temperature, program phase and runtime will be printed in intervals of 2 minutes. This item includes: 10 printer paper rolls and incl. ribbon (5x) | TTM |  | 1077.05 | USD | 0 | General Cargo | 2 | 20 | 20 | 10 | 0.004 | SHP Laboklav 80-MSLV Printer for Autoclave for vertical autoclave 80 litres | 84199015 | https://www.shp-steriltechnik.de/resources/PDF/Catalogue-Laboklav.pdf
106490 | Installation and Training for Autoclave | 1 unit | N/A | The installation of the autoclave is done in the laboratory with pre-installation work as plumbing and electricity. Also training on use and troubleshooting for the autoclave. | TTM |  | 4500 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | N/A | N/A
106824 | Vertical floor Autoclave 55lt | 1 unit | N/A | Vertical floor Autoclave 55lt top loading unit with rapid cooling, support pressure and fan inside, for open and closed vessels for the avoidance temperature gradients in the chamber, with integrated vacuum device for solid substances. - Solid sterilization and waste sterilization at 134°C. - Max. permissible pressure: 2.8bar; Sterilizing temperature: 98°C to 138°C; - Motor driven and self-checking closing mechanism; Temp. measure precision not greater than ±2%; - Flexible temperature reference sensor PT100 as serial function; The lid can optionally be opened to the right, to the left or backwards. Adjustable timer range not smaller than 10 min to 60 min; - Fast re-cooling system for liquid sterilisation using a double jacket; - Adjustable temperature working range not smaller than from 98° up to at least 137°C; - 12 pre-defined programs and plus 10 user defined programs with LCD Display. Please add to your request the UPS 3000 VA with its battery pack, the accessories (Air exhaust filters, Bucket waste sterilization, Wire-mesh basket printer) and the installation and training. The air compressor is an optional item to have a rapid re-cooling, for which support pressure is required. | TTM |  | 22074.95 | USD | 0 | General Cargo | 50 | 100 | 80 | 100 | 0.8 | SHP Laboklav 55-MSLV | 84192000 | https://www.shp-steriltechnik.de/resources/PDF/Catalogue-Laboklav.pdf
106825 | Air exhaust filters for Vertical floor Autoclave 55lt | 1 unit | N/A | Filter cartridge air exhaust filter. The autoclavable sterile filter, consisting of a filter cartridge with PTFE membrane of pore size 0.2 µm, is installed in a pressure-proof housing. The filter is also automatically sterilized in line during the sterilization process, monitored by the PT-100 temperature sensor. · Filter removal capability to particles sizes of at less of 0.01 or less microns in gas, and at less than of 0.2 microns or less in liquid; Hydrophobic filter; Built in the autoclave. · Filter life: 12 months and 200 sterilization cycles. This item includes: One installed and one spare filter to be supplied with the autoclave; | TTM |  | 723.9 | USD | 0 | General Cargo | 15 | 70 | 50 | 80 | 0.28 | SHP Laboklav 55-MSLV Exhaust air filters for vertical autoclave | 84199015 | https://www.shp-steriltechnik.de/resources/PDF/Catalogue-Laboklav.pdf
106826 | Bucket waste sterilization for Vertical floor Autoclave 55lt | 1 unit | N/A | Bucket for waste sterilization made of stainless steel for the Vertical floor Autoclave 55lt • Perforated with top handle; Stackable buckets; Diameter 400 x 300 mm. | TTM |  | 845.2 | USD | 0 | General Cargo | 1.5 | 10 | 30 | 40 | 0.012 | SHP Laboklav 55-MSLV Bucket for sterilization for vertical autoclave 80 litres | 84199015 | https://www.shp-steriltechnik.de/resources/PDF/Catalogue-Laboklav.pdf
106827 | Wire-mesh basket for Vertical floor Autoclave 55lt | 1 unit | N/A | Perforated basket for sterilization made in stainless steel with top handle for the Vertical floor Autoclave 55lt. Diameter of: 1/2 80, 400 x 300 mm | TTM |  | 505.9 | USD | 0 | General Cargo | 3 | 40 | 10 | 50 | 0.02 | SHP Laboklav 55-MSLV Baskets for sterilization for vertical autoclave 80 litres | 84199015 | https://www.shp-steriltechnik.de/resources/PDF/Catalogue-Laboklav.pdf
106828 | Printer for Vertical floor Autoclave 55lt | 1 unit | N/A | Inbuilt thermal dot printer. Printing in paper roll for the Vertical floor Autoclave 55lt.. • The paper roll is 58 mm wide and has a diameter of 35 mm. At the start of a program the date, time and the kind of program with the set parameters will be documented. • Out of the sterilization phase the pressure, temperature, program phase and runtime will be printed in intervals of 2 minutes. This item includes: 10 printer paper rolls and incl. ribbon (5x) | TTM |  | 1077.05 | USD | 0 | General Cargo | 2 | 20 | 20 | 10 | 0.004 | SHP Laboklav 55-MSLV Printer for Autoclave for vertical autoclave 80 litres | 84199015 | https://www.shp-steriltechnik.de/resources/PDF/Catalogue-Laboklav.pdf
106490 | Installation and Training for Autoclave | 1 unit | N/A | The installation of the autoclave is done in the laboratory with pre-installation work as plumbing and electricity. Also training on use and troubleshooting for the autoclave. | TTM |  | 4500 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | N/A | N/A
106580 | Air compressor for Autoclave | 1 unit | N/A | Oil-Free compact air compressor to have rapid re-cooling, for which support pressure is required. 100% oil free; 100% duty cycle; 5µm filter regulator with automatic drain. • Handle for transport and motor; receiver with internal coating; Air tank reservoir range 25 liters; max pressure of at least 5 bars; >10,000 hours cup life; receiver with internal coating • Air compressor with pressure regulator; manual condensation drainage and with thermal protection; • Motor power within 0.5 HP range; max flow at 5 bars at least 42 liters/minute; •Pressure up to 120 PSI protection; Quiet operation 61/63 dBA; easy to use; plug & play This item includes: all the hoses, adaptors and connectors to work with the autoclave. | TTM |  | 1905 | USD | 0 | General Cargo | 30 | 40 | 40 | 50 | 0.08 | JUN AIR 87R-25B | 84148022 | http://www.jun-air.com/support_documents/87R-25B%20Product%20Sheet.pdf
6.2.3 Autoclave water supply
106756 | Digital conductivity meter | 1 unit | N/A | Digital conductivity meter with connecting cable and plug power unit. Measuring cell and set of 2 hoses, each 1,5m, R ¾“. Compatible with item Reverse Osmosis Water purification system. • Microprocessor controlled automatic equipment with LED display for easy data visualization; • With visual and sound indicator for over limit value and red light for over limit value; Fault signalling: Display changes from green to red; • Conductivity measurement range: 0.1 – 300 µS/cm (automatically switches over at 9.99); • Temperature measurement range: 0.1 – 50 °C; Temperature compensation: automatic. | TTM |  | 480.85 | USD | 0 | General Cargo | 0.5 | 17 | 14 | 8 | 0.001904 | Stakpure GmbH DC400 | 90278011 | https://www.stakpure.de/ion-exchangers/conductivity-meter/?lang=en
6.3 BIOSAFETY CABINETS
6.3.1 Biosafety cabinet Class II - 3 filters
106784 | Biosafety cabinet 4 feet, 3 filters | 1 unit | N/A | Biosafety cabinet 4 feet, 3 filters. Cabinet features with panoramic front, side windows, electrical outlets, tabletops and light source for ultimate supervision, including support stand work height 75 + 5 mm. • V shaped HEPA-filter cartridges located below the table top – H-14 filter quality, 99.999% efficient against 0.3 µm particles, with large surface area for long life; • Under bench filters easily changeable, sealing without tape; • Spill-drains between each sectioned work surface ensures any spills are directed to the trough and not the HEPA filters; Low noise level < 55/56 dB(A) according to SS-EN ISO 4871 (1200/1800); • Low overall height. Will fit rooms with low ceilings easily. (Height is 1998 mm +50 mm on 750 mm stand). Leg room not compromised due to slimline filter housing; • The floor stand is in a 3 filter cabinet already inbuilt as the 3 filter is below the working surface. • UV -light mounted on the front top of the cabinet; Window in hardened glass with 3 holes for valves, Gas valve- G, Gas resetable solenoid valve;USB port for communication and with computer inclusive software package in Excel • Table top> Support stand 750-800 mm, Fluorescent dimmable light. • Electrical outlets - DE - 1 pair (1 left side & 1 right side)> Power supply> 230V 50/60Hz , Voltage 110V Biosafety cabinet with 3 filters is strongly recommended to be procured as better fit and protection in the TB diagnostics environment. Please add to your request the UPS 3000 VA with its battery pack, the accessories (vertical or horizontal ducting kit, fumigation kit for HPV, Double Exhaust, Exhaust Hepa H14, Standard Hepa and V-shaped Hepa filters) and installation and training. | TTM |  | 15643.55 | USD | 0 | General Cargo Large | 270 | 90 | 152 | 230 | 3.1464 | LaboGene ApS/Scanlaf Mars Pro 1200, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1302/mars-pro-1200-vers-2016.pdf
106785 | Vertical ducting kit for Biosafety cabinet 4 feet, 3 filters | 1 unit | N/A | Vertical ducting kit for Biosafety cabinet 4 feet with 3 filters. Automated lockable exhaust valve; • The ducting kit can be connected to an external ventilation system; • Exhaust filter is 610 mm x 457 mm x 69 mm. Double exhaust filter possible. This item includes: Ducting kit including antiblow back valve straight Ø200, signals for external ventilation system and 3 -sensor option with pressure switch. Whether a horizontal or vertical exhaust air outlet is required must be checked by the user before ordering. If the connection to the ventilation system is directly above the BSC, a vertical outlet is required, otherwise a horizontal outlet. Please order same quantity as the Biosafety cabinet 4 feet, 3 filters | TTM |  | 1591.4 | USD | 0 | General Cargo | 5 | 30 | 21 | 62 | 0.03906 | LaboGene ApS/Scanlaf Ducting kit - Mars Pro 1200, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1302/mars-pro-1200-vers-2016.pdf
106786 | Horizontal ducting kit for Biosafety cabinet 4 feet, 3 filters | 1 unit | N/A | Horizontal ducting kit for Biosafety cabinet 4 feet with 3 filters. Automated lockable exhaust valve; • The ducting kit can be connected to an external ventilation system; • Exhaust filter is 610 mm x 457 mm x 69 mm. Double exhaust filter possible. This item includes: Ducting kit including antiblow back valve straight Ø200, signals for external ventilation system and 3 -sensor option with pressure switch. Whether a horizontal or vertical exhaust air outlet is required must be checked by the user before ordering. If the connection to the ventilation system is directly above the BSC, a vertical outlet is required, otherwise a horizontal outlet. Please order same quantity as the Biosafety cabinet 4 feet, 3 filters | TTM |  | 1591.4 | USD | 0 | General Cargo | 5 | 30 | 21 | 62 | 0.03906 | LaboGene ApS/Scanlaf Ducting kit - Mars Pro 1200, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1302/mars-pro-1200-vers-2016.pdf
106787 | Fumigation kit for HPV for Biosafety cabinet 4 feet, 3 filters | 1 unit | N/A | Hydrogen Peroxide Vapor (HPV) fumigation kit to decontaminate the Biosafety cabinet 4 feet, 3 filters. This item includes:HVP fumigation plate kit with top plate, valves and sealing tape. If there is more than 1 BSC in the same laboratory, only 1 fumigation kit is required. Please order same quantity as the Biosafety cabinet 4 feet, 3 filters | TTM |  | 1109.45 | USD | 0 | General Cargo | 18 | 115 | 61 | 43 | 0.30164 | LaboGene ApS/Scanlaf Fumigation kit - Mars Pro 1200, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1302/mars-pro-1200-vers-2016.pdf
106788 | Double Exhaust filter kit for Biosafety cabinet 4feet, 3 filters | 1 kit | N/A | Double exhaust Hepa filter HP14 for Biosafety cabinet 4 feet, 3 filters. Frame made of powder coated steel. • Able to accept double Hepa filter HP14; • Double exhaust filter frame for Mars Pro 1200 dimension: 610 mm x 457 mm x 115mm.; • Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822). This item includes: A double Exhaust filter with 3 -sensor option with pressure switch and 1 spare filter. OPTIONAL ITEM: TO BE ORDERED ONLY IF REQUIRED BY THE EXISTING VENTILATION SYSTEM AND/OR LOCAL REGULATORY AUTHORITIES. THIS OPTION IS CURRENTLY ONLY COMMONLY USED IN THE UNITED KINGDOM (LOCAL REGULATIONS). The double exhaust filter is only required if connection to an external ventilation system is not possible. Please order same quantity as the Biosafety cabinet 4 feet, 3 filters | TTM |  | 1115.6 | USD | 0 | General Cargo | 15 | 160 | 68 | 25 | 0.272 | LaboGene ApS/Scanlaf Double Exhaust filter - Mars Pro 1200, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1302/mars-pro-1200-vers-2016.pdf
106789 | Exhaust Hepa filter H14 for Biosafety cabinet 4 feet, 3 filters | 2 units | N/A | Exhaust Hepa Air filter H14 for Biosafety cabinet 4 feet, 3 filters. Hepa filters exhaust filter is 610 mm x 457 mm x 69 mm; Class: H14 according with EN 1822:2009; Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822) This item includes: 2 exhaust Hepa Air filters. Please order same quantity as the Biosafety cabinet 4 feet, 3 filters | TTM |  | 1191.15 | USD | 0 | General Cargo | 8 | 91.5 | 45.7 | 69 | 0.28853 | LaboGene ApS/Scanlaf Exhaust filter filters - Mars Pro 1200, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1302/mars-pro-1200-vers-2016.pdf
106790 | Standard Hepa filter for Biosafety cabinet 4 feet, 3 filters | 1 unit | N/A | Standard HEPA Filter for Biosafety cabinet 4 feet, 3 filters. • Main recirculating filter; • Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822); • Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822); • Main filter dimension: 1220mm x 457 mm x 115 mm When ordering, please ensure that at least the same quantity of filters as the Exhaust Hepa filter H14 for Biosafety cabinet 4 feet, 3 filters is ordered so that it can be replaced at the same time. | TTM |  | 852.55 | USD | 0 | General Cargo | 15 | 152.5 | 45.7 | 11.5 | 0.080146 | LaboGene ApS/Scanlaf Standard. HEPA filter - Mars Pro 1200, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1302/mars-pro-1200-vers-2016.pdf
106791 | V-shaped Hepa filter for Biosafety cabinet 4 feet, 3 filters | 1 unit | N/A | V-shaped Hepa filter for Biosafety cabinet 4 feet, 3 filters, is operated with 4 filter cassettes (Hepa 14 V-shape filter). The V shape filter is the filter which is been built in the bottom of the BSC. This allows the change of the filters without specialist technicians. The advantage is that the filter consist of filter cassettes: 4 feet BSC: 4 cassettes. These cassettes can be changed individually. Meanwhile the BSC is operating. This allows throughout that technician and operator are been protected. A decontamination of the BSC for the change of the Filter is not required. • Type: H14 according to EN 1822:2009; Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822). The filters are not only providing an additional protection of the operator, but also protect the main and exhaust filter. This reduces the maintenance cost. Please order same quantity as the Biosafety cabinet 4 feet, 3 filters | TTM |  | 2025.5 | USD | 0 | General Cargo | 5 | 30 | 21 | 62 | 0.03906 | LaboGene ApS/Scanlaf V-shaped HEPA filter - Mars Pro 1200, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1302/mars-pro-1200-vers-2016.pdf
106489 | Installation and Training for Biosafety cabinet -Labo | 1 unit | N/A | The installation includes: On-site installation of Biosafety cabinet, including: installation of the filters and connection to the electricity. All standard accessories, consumables, parts required for the proper operation and all standard tools, cleaning and lubrication material shall be included in the offer. Final internal testing of all pre-set programs and print-out of test cycle performance data (internal printer). Certification of installation and operation by the company's service engineer The training includes: A detailed instruction of the laboratory personnel on use, function and maintenance of the equipment (user training) as well as a comprehensive maintenance plan (log-book for daily, troubleshooting and weekly, monthly and quarterly maintenance checklist) | TTM |  | 4500 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
106792 | Biosafety cabinet 5 feet, 3 filters | 1 unit | N/A | Biosafety cabinet 5 feet, 3 filters. Cabinet features with panoramic front, side windows, electrical outlets, tabletops and light source for ultimate supervision, including support stand work height 75 + 5 mm. • V shaped HEPA-filter cartridges located below the table top – H-14 filter quality, 99.999% efficient against 0.3 µm particles, with large surface area for long life; Under bench filters easily changeable, sealing without tape; • Spill-drains between each sectioned work surface ensures any spills are directed to the trough and not the HEPA filters; Low noise level < 55/56 dB(A) according to SS-EN ISO 4871 (1200/1800); • Low overall height. Will fit rooms with low ceilings easily. (Height is 1998 mm +50 mm on 750 mm stand). Leg room not compromised due to slimline filter housing; • The floor stand is in a 3 filter cabinet already inbuilt as the 3 filter is below the working surface. • UV -light mounted on the front top of the cabinet; Window in hardened glass with 3 holes for valves, Gas valve- G, Gas resettable solenoid valve; USB port for communication and with computer inclusive software package in Excel • Table top> Support stand 750-800 mm, Fluorescent dimmable light. • Electrical outlets - DE - 1 pair (1 left side & 1 right side)> Power supply> 230V 50/60Hz , Voltage 110V Biosafety cabinet with 3 filters is strongly recommended to be procured as better fit and protection in the TB diagnostics environment. Please add to your request the UPS 3000 VA with its battery pack, the accessories (vertical or horizontal ducting kit, fumigation kit for HPV, Double Exhaust, Exhaust Hepa H14, Standard Hepa and V-shaped Hepa filters) and installation and training. | TTM |  | 17877.05 | USD | 0 | General Cargo Large | 270 | 90 | 152 | 230 | 3.1464 | LaboGene ApS/Scanlaf Mars Pro 1500, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1303/mars-pro-1500-vers-2016.pdf
106793 | Vertical ducting kit for Biosafety cabinet 5 feet, 3 filters | 1 unit | N/A | Vertical ducting kit for Biosafety cabinet 5 feet with 3 filters. • Automated lockable exhaust valve; The ducting kit can be connected to an external ventilation system; • Exhaust filter is 915 x 457 x 115 mm. Double exhaust filter possible. This item includes: Ducting kit including antiblow back valve straight Ø200, signals for external ventilation system and 3 -sensor option with pressure switch. Whether a horizontal or vertical exhaust air outlet is required must be checked by the user before ordering. If the connection to the ventilation system is directly above the BSC, a vertical outlet is required, otherwise a horizontal outlet. Please order same quantity as the Biosafety cabinet 5 feet, 3 filters | TTM |  | 1591.4 | USD | 0 | General Cargo | 5 | 30 | 21 | 62 | 0.03906 | LaboGene ApS/Scanlaf Ducting kit - Mars Pro 1500, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1303/mars-pro-1500-vers-2016.pdf
106794 | Horizontal ducting kit for Biosafety cabinet 5 feet, 3 filters | 1 unit | N/A | Horizontal ducting kit for Biosafety cabinet 5 feet with 3 filters. • Automated lockable exhaust valve; The ducting kit can be connected to an external ventilation system; • Exhaust filter is 915 x 457 x 115 mm. Double exhaust filter possible. This item includes: Ducting kit including antiblow back valve straight Ø200, signals for external ventilation system and 3 -sensor option with pressure switch. Whether a horizontal or vertical exhaust air outlet is required must be checked by the user before ordering. If the connection to the ventilation system is directly above the BSC, a vertical outlet is required, otherwise a horizontal outlet. Please order same quantity as the Biosafety cabinet 5 feet, 3 filters | TTM |  | 1591.4 | USD | 0 | General Cargo | 5 | 30 | 21 | 62 | 0.03906 | LaboGene ApS/Scanlaf Ducting kit - Mars Pro 1500, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1303/mars-pro-1500-vers-2016.pdf
106795 | Fumigation kit for HPV for Biosafety cabinet 5 feet, 3 filters | 1 unit | N/A | Hydrogen Peroxide Vapor (HPV) fumigation kit to decontaminate the Biosafety cabinet 5 feet, 3 filters. This item includes: HVP fumigation plate kit with top plate, valves and sealing tape. If there is more than 1 BSC in the same laboratory, only 1 fumigation kit is required. Please order same quantity as the Biosafety cabinet 5 feet, 3 filters | TTM |  | 1109.45 | USD | 0 | General Cargo | 18 | 115 | 61 | 43 | 0.30164 | LaboGene ApS/Scanlaf Fumigation kit - Mars Pro 1500, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1303/mars-pro-1500-vers-2016.pdf
106796 | Double Exhaust filter kit for Biosafety cabinet 5feet, 3 filters | 1 kit | N/A | Exhaust Hepa filter HP14 for Biosafety cabinet 5 feet, 3 filters • Frame made of powder coated steel. Able to accept double Hepa filter HP14; • Double exhaust filter frame for Mars Pro 1500 dimension: 915 x 457 x 115 mm.; • Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822). This item includes: A double Exhaust filter with 3 -sensor option with pressure switch and 1 spare filter. OPTIONAL ITEM: TO BE ORDERED ONLY IF REQUIRED BY THE EXISTING VENTILATION SYSTEM AND/OR LOCAL REGULATORY AUTHORITIES. THIS OPTION IS CURRENTLY ONLY COMMONLY USED IN THE UNITED KINGDOM (LOCAL REGULATIONS). The double exhaust filter is only required if connection to an external ventilation system is not possible. Please order same quantity as the Biosafety cabinet 5 feet, 3 filters | TTM |  | 1115.6 | USD | 0 | General Cargo | 15 | 160 | 68 | 25 | 0.272 | LaboGene ApS/Scanlaf Double Exhaust filter - Mars Pro 1500, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1303/mars-pro-1500-vers-2016.pdf
106797 | Exhaust Hepa filter H14 for Biosafety cabinet 5 feet, 3 filters | 2 units | N/A | Exhaust Hepa Air filter H14 for Biosafety cabinet 5 feet, 3 filters5 • Hepa filters Exhaust filter is 915 x 457 x 69 mm; Class: H14 according with EN 1822:2009; • Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822) This item includes: 2 exhaust Hepa Air filters. Please order same quantity as the Biosafety cabinet 5 feet, 3 filters | TTM |  | 1191.15 | USD | 0 | General Cargo | 8 | 91.5 | 45.7 | 69 | 0.28853 | LaboGene ApS/Scanlaf Exhaust filter filters - Mars Pro 1500, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1303/mars-pro-1500-vers-2016.pdf
106798 | Standard Hepa filter for Biosafety cabinet 5 feet, 3 filters | 1 unit | N/A | Standard HEPA Filter for Biosafety cabinet 5 feet, 3 filters. • Main recirculating filter; • Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822); • Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822); • Main filter dimension: 1525x457x115 mm. When ordering, please ensure that at least the same quantity of filters as the Exhaust Hepa filter H14 for Biosafety cabinet 5 feet, 3 filters is ordered so that it can be replaced at the same time. | TTM |  | 1009.6 | USD | 0 | General Cargo | 15 | 152.5 | 45.7 | 11.5 | 0.080146 | LaboGene ApS/Scanlaf Standard. HEPA filter - Mars Pro 1500, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1303/mars-pro-1500-vers-2016.pdf
106799 | V-shaped Hepa filter for Biosafety cabinet 5 feet, 3 filters | 1 unit | N/A | V-shaped Hepa filter for Biosafety cabinet 5 feet, 3 filters, is operated with 4 filter cassettes (Hepa 14 V-shape filter). The V shape filter is the filter which is been built in the bottom of the BSC. This allows the change of the filters without specialist technicians. The advantage is that the filter consist of filter cassettes: 4 feet BSC: 5 cassettes. These cassettes can be changed individually. Meanwhile the BSC is operating. This allows throughout that technician and operator are been protected. A decontamination of the BSC for the change of the Filter is not required. • Type: H14 according to EN 1822:2009; Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822). The filters are not only providing an additional protection of the operator, but also protect the main and exhaust filter. This reduces the maintenance cost. Please order same quantity as the Biosafety cabinet 5 feet, 3 filters | TTM |  | 2531.85 | USD | 0 | General Cargo | 5 | 30 | 21 | 62 | 0.03906 | LaboGene ApS/Scanlaf V-shaped HEPA filter - Mars Pro 1500, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1303/mars-pro-1500-vers-2016.pdf
106489 | Installation and Training for Biosafety cabinet -Labo | 1 unit | N/A | The installation includes: On-site installation of Biosafety cabinet, including: installation of the filters and connection to the electricity. All standard accessories, consumables, parts required for the proper operation and all standard tools, cleaning and lubrication material shall be included in the offer. Final internal testing of all pre-set programs and print-out of test cycle performance data (internal printer). Certification of installation and operation by the company's service engineer The training includes: A detailed instruction of the laboratory personnel on use, function and maintenance of the equipment (user training) as well as a comprehensive maintenance plan (log-book for daily, troubleshooting and weekly, monthly and quarterly maintenance checklist) | TTM |  | 4500 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
106800 | Biosafety cabinet 6 feet, 3 filters | 1 unit | N/A | Biosafety cabinet 6 feet, 3 filters. Cabinet features with panoramic front, side windows, electrical outlets, tabletops and light source for ultimate supervision, including support stand work height 75 + 5 mm. • V shaped HEPA-filter cartridges located below the table top – H-14 filter quality, 99.999% efficient against 0.3 µm particles, with large surface area for long life; Under bench filters easily changeable, sealing without tape; • Spill-drains between each sectioned work surface ensures any spills are directed to the trough and not the HEPA filters; Low noise level < 55/56 dB(A) according to SS-EN ISO 4871 (1200/1800); • Low overall height. Will fit rooms with low ceilings easily. (Height is 1998 mm +50 mm on 750 mm stand). Leg room not compromised due to slimline filter housing; • The floor stand is in a 3 filter cabinet already inbuilt as the 3 filter is below the working surface. • UV -light mounted on the front top of the cabinet; Window in hardened glass with 3 holes for valves, Gas valve- G, Gas resettable solenoid valve; USB port for communication and with computer inclusive software package in Excel • Table top> Support stand 750-800 mm, Fluorescent dimmable light. • Electrical outlets - DE - 1 pair (1 left side & 1 right side)> Power supply> 230V 50/60Hz , Voltage 110V Biosafety cabinet with 3 filters is strongly recommended to be procured as better fit and protection in the TB diagnostics environment. Please add to your request the UPS 3000 VA with its battery pack, the accessories (vertical or horizontal ducting kit, fumigation kit for HPV, Double Exhaust, Exhaust Hepa H14, Standard Hepa and V-shaped Hepa filters) and installation and training. | TTM |  | 18906.05 | USD | 0 | General Cargo Large | 270 | 90 | 152 | 230 | 3.1464 | LaboGene ApS/Scanlaf Mars Pro 1800, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1304/mars-pro-1800-vers-2016.pdf
106801 | Vertical ducting kit for Biosafety cabinet 6 feet, 3 filters | 1 unit | N/A | Vertical ducting kit for Biosafety cabinet 6 feet with 3 filters. • Automated lockable exhaust valve; The ducting kit can be connected to an external ventilation system; • Exhaust filter is 915 x 457 x 115 mm. Double exhaust filter possible. This item includes: Ducting kit including antiblow back valve straight Ø200, signals for external ventilation system and 3 -sensor option with pressure switch. Whether a horizontal or vertical exhaust air outlet is required must be checked by the user before ordering. If the connection to the ventilation system is directly above the BSC, a vertical outlet is required, otherwise a horizontal outlet. Please order same quantity as the Biosafety cabinet 6 feet, 3 filters | TTM |  | 1591.4 | USD | 0 | General Cargo | 5 | 30 | 21 | 62 | 0.03906 | LaboGene ApS/Scanlaf Ducting kit - Mars Pro 1800, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1304/mars-pro-1800-vers-2016.pdf
106802 | Horizontal ducting kit for Biosafety cabinet 6 feet, 3 filters | 1 unit | N/A | Horizontal ducting kit for Biosafety cabinet 6 feet with 3 filters. • Automated lockable exhaust valve; The ducting kit can be connected to an external ventilation system; • Exhaust filter is 915 x 457 x 115 mm. Double exhaust filter possible. This item includes: Ducting kit including antiblow back valve straight Ø200, signals for external ventilation system and 3 -sensor option with pressure switch. Whether a horizontal or vertical exhaust air outlet is required must be checked by the user before ordering. If the connection to the ventilation system is directly above the BSC, a vertical outlet is required, otherwise a horizontal outlet. Please order same quantity as the Biosafety cabinet 6 feet, 3 filters | TTM |  | 1591.4 | USD | 0 | General Cargo | 5 | 30 | 21 | 62 | 0.03906 | LaboGene ApS/Scanlaf Ducting kit - Mars Pro 1800, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1304/mars-pro-1800-vers-2016.pdf
106803 | Fumigation kit for HPV for Biosafety cabinet 6 feet, 3 filters | 1 unit | N/A | Hydrogen Peroxide Vapor (HPV) fumigation kit to decontaminate the Biosafety cabinet 6 feet, 3 filters. This item includes: HVP fumigation plate kit with top plate, valves and sealing tape. If there is more than 1 BSC in the same laboratory, only 1 fumigation kit is required. Please order same quantity as the Biosafety cabinet 6 feet, 3 filters | TTM |  | 1109.45 | USD | 0 | General Cargo | 18 | 115 | 61 | 43 | 0.30164 | LaboGene ApS/Scanlaf Fumigation kit - Mars Pro 1800, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1304/mars-pro-1800-vers-2016.pdf
106804 | Double Exhaust filter kit for Biosafety cabinet 6feet, 3 filters | 1 kit | N/A | Double exhaust Hepa filter HP14 for Biosafety cabinet 6 feet, 3 filters • Frame made of powder coated steel. Able to accept double Hepa filter HP14; • Double exhaust filter frame for Mars Pro 1800 dimension: 915 x 457 x 115 mm.; • Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822). Thi item includes: A double Exhaust filter with 3 -sensor option with pressure switch and 1 spare filter. OPTIONAL ITEM: TO BE ORDERED ONLY IF REQUIRED BY THE EXISTING VENTILATION SYSTEM AND/OR LOCAL REGULATORY AUTHORITIES. THIS OPTION IS CURRENTLY ONLY COMMONLY USED IN THE UNITED KINGDOM (LOCAL REGULATIONS). The double exhaust filter is only required if connection to an external ventilation system is not possible. Please order same quantity as the Biosafety cabinet 6 feet, 3 filters | TTM |  | 1115.6 | USD | 0 | General Cargo | 15 | 160 | 68 | 25 | 0.272 | LaboGene ApS/Scanlaf Double Exhaust filter - Mars Pro 1800, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1304/mars-pro-1800-vers-2016.pdf
106805 | Exhaust Hepa filter H14 for Biosafety cabinet 6 feet, 3 filters | 2 units | N/A | Exhaust Hepa Air filter H14 for Biosafety cabinet 6 feet, 3 filters5 • Hepa filters Exhaust filter is 915 x 457 x 69 mm; Class: H14 according with EN 1822:2009; • Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822) This item includes: 2 exhaust Hepa Air filters. Please order same quantity as the Biosafety cabinet 6 feet, 3 filters | TTM |  | 1191.15 | USD | 0 | General Cargo | 8 | 91.5 | 45.7 | 69 | 0.28853 | LaboGene ApS/Scanlaf Exhaust filter filters - Mars Pro 1800, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1304/mars-pro-1800-vers-2016.pdf
106806 | Standard Hepa filter for Biosafety cabinet 6 feet, 3 filters | 1 unit | N/A | Standard HEPA Filter for Biosafety cabinet 6 feet, 3 filters. • Main recirculating filter; • Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822); • Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822); • Main filter dimension: 1525x457x115 mm. When ordering, please ensure that at least the same quantity of filters as the Exhaust Hepa filter H14 for Biosafety cabinet 6 feet, 3 filters is ordered so that it can be replaced at the same time. | TTM |  | 1009.6 | USD | 0 | General Cargo | 15 | 152.5 | 45.7 | 11.5 | 0.080146 | LaboGene ApS/Scanlaf Standard. HEPA filter - Mars Pro 1800, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1304/mars-pro-1800-vers-2016.pdf
106807 | V-shaped Hepa filter for Biosafety cabinet 6 feet, 3 filters | 1 unit | N/A | V-shaped Hepa filter for Biosafety cabinet 6 feet, 3 filters, is operated with 4 filter cassettes (Hepa 14 V-shape filter). The V shape filter is the filter which is been built in the bottom of the BSC. This allows the change of the filters without specialist technicians. The advantage is that the filter consist of filter cassettes: 4 feet BSC: 6 cassettes. These cassettes can be changed individually. Meanwhile the BSC is operating. This allows throughout that technician and operator are been protected. A decontamination of the BSC for the change of the Filter is not required. • Type: H14 according to EN 1822:2009; Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822). The filters are not only providing an additional protection of the operator, but also protect the main and exhaust filter. This reduces the maintenance cost. Please order same quantity as the Biosafety cabinet 6 feet, 3 filters | TTM |  | 3038.25 | USD | 0 | General Cargo | 5 | 30 | 21 | 62 | 0.03906 | LaboGene ApS/Scanlaf V-shaped HEPA filter - Mars Pro 1800, 3 filter | 90189084 | https://configurator.labogeneapps.com/media/1304/mars-pro-1800-vers-2016.pdf
106489 | Installation and Training for Biosafety cabinet -Labo | 1 unit | N/A | The installation includes: On-site installation of Biosafety cabinet, including: installation of the filters and connection to the electricity. All standard accessories, consumables, parts required for the proper operation and all standard tools, cleaning and lubrication material shall be included in the offer. Final internal testing of all pre-set programs and print-out of test cycle performance data (internal printer). Certification of installation and operation by the company's service engineer The training includes: A detailed instruction of the laboratory personnel on use, function and maintenance of the equipment (user training) as well as a comprehensive maintenance plan (log-book for daily, troubleshooting and weekly, monthly and quarterly maintenance checklist) | TTM |  | 4500 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
6.3.2 Biosafety cabinet Class II Type A2, 4 feet
106757 | Biosafety cabinet (BSC) Class II, Type A2, 4 feet -Labo | 1 unit | N/A | Biosafety cabinet Class II Type A2 with panoramic front, side windows, electrical outlets, tabletops and light source for ultimate supervision and glare-free lighting V-shaped angled inflow grills allowing a comfortable working position 110 mm HEPA filters with efficiency at 99,999 % against 0,3 µm particles, ensuring a clean sterile work chamber environment. • Air velocity, vertical flow: 0,35 m/s; Air velocity, deviation: 10 ±%; 2 Hepa filters: H-14; • Air velocity, vertical flow: 0,35 m/s; Air velocity, deviation: 10 ±%; Down Flow rate: 1100 m3 /h; • Noise level, ISO6081: <54 dB(A); Light intensity variable: 0-2000 Lux; Window material: hardened/laminated glass; Cabinet material: polyester coated /AISI 304 stainless steel; • Max. 2 pcs Electrical outlets in each side panel (left and right); 3 whole in each side window for gas valves in each side window UV-light. UV -light mounted on the front top of the cabinet • Window in hardened glass with 3 holes for valves, Gas valve- G, Gas resettable solenoid valve USB port for communication and with computer inclusive software package in Excel • Table top> Pre filter grid with EU-3, painted version (Model 2016), Fluorescent dimmable light. • Electrical outlets - DE - 1 pair (1 left side & 1 right side)> Power supply> 230V 50/60Hz , Voltage 110V. Please add to your request the UPS 3000 VA with its battery pack, the accessories (stand, vertical or horizontal ducting kit, fumigation kit for HPV, Double Exhaust, Exhaust Hepa H14, Standard filters) and installation and training. | TTM |  | 11272.9 | USD | 0 | General Cargo Large | 270 | 90 | 152 | 230 | 3.1464 | LaboGene ApS/Scanlaf Mars 1200, 2 filter | 90189084 | https://configurator.labogeneapps.com/media/1298/mars-1200-vers-2016.pdf
106758 | Stand for BSC Class II, Type A2, 4 feet -Labo | 1 unit | N/A | Adjustable stand for the Biosafety cabinet 4 feet with 2 filters. Floor stand adjustable for BSC; • Support stand 750 -800mm; Mounted on stabilizing adjustable feet; • Made of steel covered with epoxy coated paint, feet's made of stainless steel AISI 304. Needed if you do not have laboratory benches or the like where you can put the BSC on. | TTM |  | 582.25 | USD | 0 | General Cargo | 70 | 150 | 92 | 90 | 1.242 | LaboGene ApS/Scanlaf BSC Stand - Mars 1200, 2 filter. | 90189084 | https://configurator.labogeneapps.com/media/1298/mars-1200-vers-2016.pdf
106759 | Vertical ducting kit for BSC Class II, Type A2, 4 feet -Labo | 1 unit | N/A | Vertical ducting kit for Biosafety cabinet 4 feet with 2 filters. For a minimal room height of 250 cm • Automated lockable exhaust valve; The ducting kit can be connected to an external ventilation system; • Exhaust filter is 610 mm x 457 mm x 69 mm. Double exhaust filter possible. This item includes: Ducting kit including antiblow back valve straight Ø200, signals for external ventilation system and 3 -sensor option with pressure switch. Whether a horizontal or vertical exhaust air outlet is required must be checked by the user before ordering. If the connection to the ventilation system is directly above the BSC, a vertical outlet is required, otherwise a horizontal outlet. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 4 feet -Labo | TTM |  | 1540 | USD | 0 | General Cargo | 5 | 30 | 21 | 62 | 0.03906 | LaboGene ApS/Scanlaf ducting kit - Mars 1200, 2 filter. | 90189084 | https://configurator.labogeneapps.com/media/1298/mars-1200-vers-2016.pdf
106760 | Horizontal ducting kit for BSC Class II, Type A2, 4 feet -Labo | 1 unit | N/A | Horizontal ducting kit for Biosafety cabinet 4 feet with 2 filters. For a minimal room height of 250 cm • Automated lockable exhaust valve; The ducting kit can be connected to an external ventilation system; • Exhaust filter is 610 mm x 457 mm x 69 mm. Double exhaust filter possible. This item includes: Ducting kit including antiblow back valve straight Ø200, signals for external ventilation system and 3 -sensor option with pressure switch. Whether a horizontal or vertical exhaust air outlet is required must be checked by the user before ordering. If the connection to the ventilation system is directly above the BSC, a vertical outlet is required, otherwise a horizontal outlet. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 4 feet -Labo | TTM |  | 1540 | USD | 0 | General Cargo | 5 | 30 | 21 | 62 | 0.03906 | LaboGene ApS/Scanlaf ducting kit - Mars 1200, 2 filter. | 90189084 | https://configurator.labogeneapps.com/media/1298/mars-1200-vers-2016.pdf
106761 | Fumigation kit for HPV for BSC Class II, Type A2, 4 feet -Labo | 1 unit | N/A | Hydrogen Peroxide Vapor (HPV) fumigation kit to decontaminate the Biosafety cabinet 4 feet, 2 filters. This item includes: HVP fumigation plate kit with top plate, valves and sealing tape. If there is more than 1 BSC in the same laboratory, only 1 fumigation kit is required. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 4 feet -Labo | TTM |  | 1109.45 | USD | 0 | General Cargo | 18 | 115 | 61 | 43 | 0.30164 | LaboGene ApS/Scanlaf Fumigation kit - Mars 1200, 2 filter. | 90189084 | https://configurator.labogeneapps.com/media/1298/mars-1200-vers-2016.pdf
106762 | Double Exhaust filter kit for BSC Class II, Type A2, 4feet -Labo | 1 kit | N/A | Double exhaust Hepa filter HP14 for Biosafety cabinet 4 feet, 2 filters. Frame made of powder coated steel;. • Able to accept double Hepa filter HP14; Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822); Double exhaust filter frame for Mars 1200 dimension: 610 mm x 457 mm x 115 mm. This item includes: A double Exhaust filter with 3 -sensor option with pressure switch and 1 spare filter. OPTIONAL ITEM: TO BE ORDERED ONLY IF REQUIRED BY THE EXISTING VENTILATION SYSTEM AND/OR LOCAL REGULATORY AUTHORITIES. THIS OPTION IS CURRENTLY ONLY COMMONLY USED IN THE UNITED KINGDOM (LOCAL REGULATIONS). Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 4 feet -Labo | TTM |  | 1115.6 | USD | 0 | General Cargo | 15 | 160 | 68 | 25 | 0.272 | LaboGene ApS/Scanlaf Double Exhaust filter - Mars 1200, 2 filter. | 90189084 | https://configurator.labogeneapps.com/media/1298/mars-1200-vers-2016.pdf
106763 | Exhaust Hepa filter H14 for BSC Class II, Type A2, 4 feet -Labo | 2 units | N/A | Exhaust Hepa Air filter H14 for Biosafety cabinet 4 feet, 2 filters. Class: H14 according with EN 1822:2009. • Hepa filters exhaust filter is 610 mm x 457 mm x 69 mm; • Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822) This item includes: 2 exhaust Hepa Air filters. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 4 feet -Labo | TTM |  | 1191.15 | USD | 0 | General Cargo | 8 | 91.5 | 45.7 | 69 | 0.28853 | LaboGene ApS/Scanlaf Exhaust filter filters - Mars 1200, 2 filter. | 90189084 | https://configurator.labogeneapps.com/media/1298/mars-1200-vers-2016.pdf
106764 | Standard Hepa filter for BSC Class II, Type A2, 4 feet -Labo | 1 unit | N/A | Standard HEPA Filter for Biosafety cabinet 4 feet, 2 filters. Main recirculating filter; • Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822); • Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822); • Main filter dimension: 1220mm x 457 mm x 115 mm When ordering, please ensure that at least the same quantity of filters as the Exhaust Hepa filter H14 for BSC Class II, Type A2, 4 feet -Labo is ordered so that it can be replaced at the same time. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 4 feet -Labo | TTM |  | 852.55 | USD | 0 | General Cargo | 15 | 152.5 | 45.7 | 11.5 | 0.080146 | LaboGene ApS/Scanlaf Standard. HEPA filter - Mars 1200, 2 filter | 90189084 | https://configurator.labogeneapps.com/media/1298/mars-1200-vers-2016.pdf
106489 | Installation and Training for Biosafety cabinet -Labo | 1 unit | N/A | The installation includes: On-site installation of Biosafety cabinet, including: installation of the filters and connection to the electricity. All standard accessories, consumables, parts required for the proper operation and all standard tools, cleaning and lubrication material shall be included in the offer. Final internal testing of all pre-set programs and print-out of test cycle performance data (internal printer). Certification of installation and operation by the company's service engineer The training includes: A detailed instruction of the laboratory personnel on use, function and maintenance of the equipment (user training) as well as a comprehensive maintenance plan (log-book for daily, troubleshooting and weekly, monthly and quarterly maintenance checklist) | TTM |  | 4500 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | LaboGene ApS/Scanlaf installaiton and training | N/A
106860 | Biosafety cabinet (BSC) Class II, Type A2, 4 feet -Thermo | 1 unit | N/A | Biosafety cabinet Class II Type A2 with panoramic front & side windows for ultimate supervision and glare-free lighting allowing a comfortable working position 110 mm HEPA filters with efficiency at 99,995 % against 0,3 µm particles, ensuring a clean sterile work chamber environment. • Load Bearing Capacity: 50 kg; Max. ambient temp during operation: 40 °C (104 °F); Min. ambient temp during operation:10 °C (50 °F); Max. humidity/ operation: 90, non-condensing % r.H. • Above room temp with window closed: < 2 °K; Max. humidity/ storage: 95 % r.H; SmartPort: 182.9 mm (7.2 in.) • Fluorescent Light; Air Inflow Rate: 207 - 517 cfm; Filter Efficiency: 99.995% MPPS; Filter Type: H14 HEPA EN 1822, 99.995% MPPSl; The recommended UPS to use wit this item is: UPS 3000 VA with its battery pack. | AMEX |  | 7434 | USD | 0 | General Cargo Large | 496 | 92.5 | 141 | 173 | 2.26 | Thermo Scientific 1300 Series Class II, Type A2 | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106861 | Base for BSC Class II, Type A2, 4 feet -Thermo | 1 unit | N/A | Adjustable base for the Biosafety cabinet 4 feet with 2 filters. Floor stand adjustable for BSC; • Adjustable height range: 800 mm up to 850 mm; Mounted on stabilizing adjustable feet; • Made of steel covered with epoxy coated paint, feet's made of stainless steel AISI 304. | AMEX |  | 722.6 | USD | 0 | General Cargo | 40 |  |  |  | 0.5 | Thermo Scientific 1300 Series Class II, Type A2 - Stand | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106862 | Vertical ducting kit for BSC Class II, Type A2, 4 feet -Thermo | 1 unit | N/A | Vertical ducting kit for Biosafety cabinet 4 feet with 2 filters. Ducting kit and anti blowback valve of at least 200mm; Automatic damper lockable exhaust valve or similar automatic system. · Exhaust duct to be connected to the external ventilation system in the laboratory; · Able to accept frame for double HEPA filter of at least 90 x 45 cm; This item includes: accessories and parts for the installation: thimble, exhaust duct, sleeves, clamps, others. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 4 feet -Thermo | AMEX |  | 727.5 | USD | 0 | General Cargo | 25 |  |  |  | 0.35 | Thermo Scientific 1300 Series Class II, Type A2 -Vertical ducting kit | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106863 | Horizontal ducting kit for BSC Class II Type A2, 4 feet -Thermo | 1 unit | N/A | Horizontal ducting kit for Biosafety cabinet 4 feet with 2 filters. Ducting kit and anti blowback valve of at least 200mm; Automatic damper lockable exhaust valve or similar automatic system. · Exhaust duct to be connected to the external ventilation system in the laboratory; · Able to accept frame for double HEPA filter of at least 90 x 45 cm; This item includes: accessories and parts for the installation: thimble, exhaust duct, sleeves, clamps, others. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 4 feet -Thermo | AMEX |  | 727.5 | USD | 0 | General Cargo | 25 |  |  |  | 0.35 | Thermo Scientific 1300 Series Class II, Type A2 - Horizontal ducting kit | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106864 | Fumigation kit for HPV for BSC Class II Type A2, 4 feet -Thermo | 1 unit | N/A | Hydrogen Peroxide Vapor (HPV) fumigation kit to decontaminate the Biosafety cabinet 4 feet, 2 filters. This item includes: valves, top closing/plate, sealing kit and other accessories necessaries for the fumigation process. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 4 feet -Thermo | AMEX |  | 416.7 | USD | 0 | General Cargo | 10 |  |  |  | 0.06 | Thermo Scientific 1300 Series Class II, Type A2 - Fumigation kit | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106865 | Double Exhaust filter kit for BSC Class II Type A2, 4feet Thermo | 1 unit | N/A | Double exhaust Hepa filter HP14 for Biosafety cabinet 4 feet, 2 filters. Frame made of steel covered with coated paint or preferably of Stainless steel AISI 304. Able to accept double HEPA filter. · Frame dimension able to accept HEPA filter size of at least 90 x 45 cm; · Double HEPA filter, each: two Micro glass-fibers HEPA to block particles of at least 0.3 µm size and larger with an efficiency of at least 99.99%. Thi item includes: A double Exhaust filter OPTIONAL ITEM: TO BE ORDERED ONLY IF REQUIRED BY THE EXISTING VENTILATION SYSTEM AND/OR LOCAL REGULATORY AUTHORITIES. THIS OPTION IS CURRENTLY ONLY COMMONLY USED IN THE UNITED KINGDOM (LOCAL REGULATIONS). Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 4 feet -Thermo | AMEX |  | 860.6 | USD | 0 | General Cargo | 7.5 |  |  |  | 0.045 | Thermo Scientific 1300 Series Class II, Type A2 - Exhaust Hepa filter HP14 | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106866 | Exhaust Hepa filter H14 for BSC Class II Type A2, 4 feet Thermo | 2 units | N/A | Exhaust Hepa Air filter H14 for Biosafety cabinet 4 feet, 2 filters. HEPA filters size of at least 90 x 45 cm; Class: H14 according with EN 1822:2009; · Micro glass-fibers HEPA to block particles of at least 0.3 µm size and larger with an efficiency of at least 99.99% This item includes: 2 exhaust Hepa Air filters. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 4 feet -Thermo | AMEX |  | 524.7 | USD | 0 | General Cargo | 5 |  |  |  | 0.03 | Thermo Scientific 1300 Series Class II, Type A2 - Exhaust Hepa filter H14 | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106867 | Main Hepa filter for BSC Class II, Type A2, 4 feet -Thermo | 1 unit | N/A | Main HEPA Filter for Biosafety cabinet 4 feet, 2 filters. Main recirculating filter; · HEPA filter size of at least 90 x 45 cm; Type H14 according with EN 1822:2009; · Micro glass-fibers HEPA to block particles of at least 0.3 µm size and larger with an efficiency of at least 99.99% Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 4 feet -Thermo | AMEX |  | 669.1 | USD | 0 | General Cargo | 5 |  |  |  | 0.03 | Thermo Scientific 1300 Series Class II, Type A2 - Main Hepa filter | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106868 | Installation and Training for Biosafety cabinet -Thermo | 1 unit | N/A | The installation includes: On-site installation of Biosafety cabinet, including: installation of the filters and connection to the electricity. All standard accessories, consumables, parts required for the proper operation and all standard tools, cleaning and lubrication material shall be included in the offer. Final internal testing of all pre-set programs and print-out of test cycle performance data (internal printer). Certification of installation and operation by the company's service engineer The training includes: A detailed instruction of the laboratory personnel on use, function and maintenance of the equipment (user training) as well as a comprehensive maintenance plan (log-book for daily, troubleshooting and weekly, monthly and quarterly maintenance checklist) | AMEX |  | 1400 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | Thermo Scientific 1300 Series Installation and training | N/A | N/A
6.3.3 Biosafety cabinet Class II Type A2, 5 feet
106766 | Biosafety cabinet (BSC) Class II, Type A2, 5 feet -Labo | 1 unit | N/A | Biosafety cabinet Class II Type A2 with panoramic front, side windows, electrical outlets, tabletops and light source for ultimate supervision and glare-free lighting V-shaped angled inflow grills allowing a comfortable working position 110 mm HEPA filters with efficiency at 99,999 % against 0,3 µm particles, ensuring a clean sterile work chamber environment. • Air velocity, vertical flow: 0,35 m/s; Air velocity, deviation: 10 ±%; 2 Hepa filters: H-14; • Air velocity, vertical flow: 0,35 m/s; Air velocity, deviation: 10 ±%; Down Flow rate: 1100 m3 /h; • Noise level, ISO6081: <54 dB(A); Light intensity variable: 0-2000 Lux; Window material: hardened/laminated glass; Cabinet material: polyester coated /AISI 304 stainless steel; • Max. 2 pcs Electrical outlets in each side panel (left and right); 3 whole in each side window for gas valves in each side window UV-light. UV -light mounted on the front top of the cabinet • Window in hardened glass with 3 holes for valves, Gas valve- G, Gas resettable solenoid valve USB port for communication and with computer inclusive software package in Excel • Table top> Pre filter grid with EU-3, painted version (Model 2016), Fluorescent dimmable light. • Electrical outlets - DE - 1 pair (1 left side & 1 right side)> Power supply> 230V 50/60Hz , Voltage 110V Please add to your request the UPS 3000 VA with its battery pack, the accessories (stand, vertical or horizontal ducting kit, fumigation kit for HPV, Double Exhaust, Exhaust Hepa H14, Standard filters) and installation and training. | TTM |  | 12667.65 | USD | 0 | General Cargo Large | 270 | 90 | 152 | 230 | 3.1464 | LaboGene ApS/Scanlaf Mars 1500, 2 filter | 90189084 | https://configurator.labogeneapps.com/media/1299/mars-1500-vers-2016.pdf
106767 | Stand for BSC Class II, Type A2, 5 feet -Labo | 1 unit | N/A | Adjustable stand for the Biosafety cabinet 5 feet with 2 filters. Floor stand adjustable for BSC; • Support stand 750 -800mm; Mounted on stabilizing adjustable feet; • Made of steel covered with epoxy coated paint, feet's made of stainless steel AISI 304. Needed if you do not have laboratory benches or the like where you can put the BSC on. | TTM |  | 604.05 | USD | 0 | General Cargo | 70 | 150 | 92 | 90 | 1.242 | LaboGene ApS/Scanlaf BSC Stand - Mars 1500, 2 filter. | 90189084 | https://configurator.labogeneapps.com/media/1299/mars-1500-vers-2016.pdf
106768 | Vertical ducting kit for BSC Class II, Type A2, 5 feet -Labo | 1 unit | N/A | Vertical ducting kit for Biosafety cabinet 5 feet with 2 filters. For a minimal room height of 250 cm • Automated lockable exhaust valve; The ducting kit can be connected to an external ventilation system; • Exhaust filter is 610 mm x 457 mm x 69 mm. Double exhaust filter possible. This item includes: Ducting kit including antiblow back valve straight Ø200, signals for external ventilation system and 3 -sensor option with pressure switch. Whether a horizontal or vertical exhaust air outlet is required must be checked by the user before ordering. If the connection to the ventilation system is directly above the BSC, a vertical outlet is required, otherwise a horizontal outlet. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 5 feet -Labo | TTM |  | 1591.4 | USD | 0 | General Cargo | 5 | 30 | 21 | 62 | 0.03906 | LaboGene ApS/Scanlaf BSC ducting kit - Mars 1500, 2 filter. | 90189084 | https://configurator.labogeneapps.com/media/1299/mars-1500-vers-2016.pdf
106769 | Horizontal ducting kit for BSC Class II, Type A2, 5 feet -Labo | 1 unit | N/A | Horizontal ducting kit for Biosafety cabinet 5 feet with 2 filters. For a minimal room height of 250 cm • Automated lockable exhaust valve; The ducting kit can be connected to an external ventilation system; • Exhaust filter is 610 mm x 457 mm x 69 mm. Double exhaust filter possible. This item includes: Ducting kit including antiblow back valve straight Ø200, signals for external ventilation system and 3 -sensor option with pressure switch. Whether a horizontal or vertical exhaust air outlet is required must be checked by the user before ordering. If the connection to the ventilation system is directly above the BSC, a vertical outlet is required, otherwise a horizontal outlet. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 5 feet -Labo | TTM |  | 1591.4 | USD | 0 | General Cargo | 5 | 30 | 21 | 62 | 0.03906 | LaboGene ApS/Scanlaf BSC ducting kit - Mars 1500, 2 filter. | 90189084 | https://configurator.labogeneapps.com/media/1299/mars-1500-vers-2016.pdf
106770 | Fumigation kit for HPV for BSC Class II, Type A2, 5 feet -Labo | 1 unit | N/A | Hydrogen Peroxide Vapor (HPV) fumigation kit to decontaminate the Biosafety cabinet 4 feet, 2 filters. This item includes: HVP fumigation plate kit with top plate, valves and sealing tape. If there is more than 1 BSC in the same laboratory, only 1 fumigation kit is required. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 5 feet -Labo | TTM |  | 1109.45 | USD | 0 | General Cargo | 18 | 115 | 61 | 43 | 0.30164 | LaboGene ApS/Scanlaf Fumigation kit - Mars 1500, 2 filter. | 90189084 | https://configurator.labogeneapps.com/media/1299/mars-1500-vers-2016.pdf
106771 | Double Exhaust filter kit for BSC Class II, Type A2, 5feet -Labo | 1 kit | N/A | Double exhaust Hepa filter HP14 for Biosafety cabinet 5 feet, 2 filters. Frame made of powder coated steel;. • Able to accept double Hepa filter HP14; Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822); Double exhaust filter frame for Mars 1200 dimension: 915 mm x 457 mm x 115 mm. The double exhaust filter is only required if connection to an external ventilation system is not possible. This item includes: A double Exhaust filter with 3 -sensor option with pressure switch and 1 spare filter. OPTIONAL ITEM: TO BE ORDERED ONLY IF REQUIRED BY THE EXISTING VENTILATION SYSTEM AND/OR LOCAL REGULATORY AUTHORITIES. THIS OPTION IS CURRENTLY ONLY COMMONLY USED IN THE UNITED KINGDOM (LOCAL REGULATIONS). Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 5 feet -Labo | TTM |  | 1115.6 | USD | 0 | General Cargo | 15 | 160 | 68 | 25 | 0.272 | LaboGene ApS/Scanlaf Double Exhaust filter - Mars 1500, 2 filter. | 90189084 | https://configurator.labogeneapps.com/media/1299/mars-1500-vers-2016.pdf
106772 | Exhaust Hepa filter H14 for BSC Class II, Type A2, 5 feet -Labo | 2 units | N/A | Exhaust Hepa Air filter H14 for Biosafety cabinet 5 feet, 2 filters. Class: H14 according with EN 1822:2009. • Hepa filters exhaust filter is 915 mm x 457 mm x 69 mm; • Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822) This item includes: 2 exhaust Hepa Air filters. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 5 feet -Labo | TTM |  | 1191.15 | USD | 0 | General Cargo | 8 | 91.5 | 45.7 | 69 | 0.28853 | LaboGene ApS/Scanlaf Exhaust filter filters - Mars 1500, 2 filter. | 90189084 | https://configurator.labogeneapps.com/media/1299/mars-1500-vers-2016.pdf
106773 | Standard Hepa filter for BSC Class II, Type A2, 5 feet -Labo | 1 unit | N/A | Standard HEPA Filter for Biosafety cabinet 5 feet, 2 filters. Main recirculating filter; • Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822); • Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822); • Main filter dimension: 1225mm x 457 mm x 115 mm When ordering, please ensure that at least the same quantity of filters as item 106772 is ordered so that it can be replaced at the same time. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 5 feet -Labo | TTM |  | 1009.6 | USD | 0 | General Cargo | 15 | 152.5 | 45.7 | 11.5 | 0.080146 | LaboGene ApS/Scanlaf Standard. HEPA filter - Mars 1500, 2 filter | 90189084 | https://configurator.labogeneapps.com/media/1299/mars-1500-vers-2016.pdf
106489 | Installation and Training for Biosafety cabinet -Labo | 1 unit | N/A | The installation includes: On-site installation of Biosafety cabinet, including: installation of the filters and connection to the electricity. All standard accessories, consumables, parts required for the proper operation and all standard tools, cleaning and lubrication material shall be included in the offer. Final internal testing of all pre-set programs and print-out of test cycle performance data (internal printer). Certification of installation and operation by the company's service engineer The training includes: A detailed instruction of the laboratory personnel on use, function and maintenance of the equipment (user training) as well as a comprehensive maintenance plan (log-book for daily, troubleshooting and weekly, monthly and quarterly maintenance checklist) | TTM |  | 4500 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
106870 | Biosafety cabinet (BSC) Class II, Type A2, 5 feet -Thermo | 1 unit | N/A | Biosafety cabinet Class II Type A2 with panoramic front & side windows for ultimate supervision and glare-free lighting allowing a comfortable working position 110 mm HEPA filters with efficiency at 99,995 % against 0,3 µm particles, ensuring a clean sterile work chamber environment. • Load Bearing Capacity: 50 kg; Max. ambient temp during operation: 40 °C (104 °F); Min. ambient temp during operation:10 °C (50 °F); Max. humidity/ operation: 90, non-condensing % r.H. • Above room temp with window closed: < 2 °K; Max. humidity/ storage: 95 % r.H; SmartPort: 182.9 mm (7.2 in.) • Fluorescent Light; Air Inflow Rate: 207 - 517 cfm; Filter Efficiency: 99.995% MPPS; Filter Type: H14 HEPA EN 1822, 99.995% MPPSl; The recommended UPS to use wit this item is: UPS 3000 VA with its battery pack. | AMEX |  | 8635.3 | USD | 0 | General Cargo Large | 573 | 92.5 | 171 | 173 | 2.74 | Thermo Scientific 1300 Series Class II, Type A2 | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106871 | Base for BSC Class II, Type A2, 5 feet -Thermo | 1 unit | N/A | Adjustable base for the Biosafety cabinet 5 feet with 2 filters. Floor stand adjustable for BSC; • Adjustable height range: 800 mm up to 850 mm; Mounted on stabilizing adjustable feet; • Made of steel covered with epoxy coated paint, feet's made of stainless steel AISI 304. | AMEX |  | 786.7 | USD | 0 | General Cargo | 50 |  |  |  | 0.55 | Thermo Scientific 1300 Series Class II, Type A2 - Stand | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106872 | Vertical ducting kit for BSC Class II, Type A2, 5 feet -Thermo | 1 unit | N/A | Vertical ducting kit for Biosafety cabinet 5 feet with 2 filters. Ducting kit and anti blowback valve of at least 200mm; Automatic damper lockable exhaust valve or similar automatic system. · Exhaust duct to be connected to the external ventilation system in the laboratory; · Able to accept frame for double HEPA filter of at least 90 x 45 cm; This item includes: accessories and parts for the installation: thimble, exhaust duct, sleeves, clamps, others. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 5 feet -Thermo | AMEX |  | 727.5 | USD | 0 | General Cargo | 25 |  |  |  | 0.35 | Thermo Scientific 1300 Series Class II, Type A2 -Vertical ducting kit | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106873 | Horizontal ducting kit for BSC Class II, Type A2, 5 feet -Thermo | 1 unit | N/A | Horizontal ducting kit for Biosafety cabinet 5 feet with 2 filters. Ducting kit and anti blowback valve of at least 200mm; Automatic damper lockable exhaust valve or similar automatic system. · Exhaust duct to be connected to the external ventilation system in the laboratory; · Able to accept frame for double HEPA filter of at least 90 x 45 cm; This item includes: accessories and parts for the installation: thimble, exhaust duct, sleeves, clamps, others. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 5 feet -Thermo | AMEX |  | 727.5 | USD | 0 | General Cargo | 25 |  |  |  | 0.35 | Thermo Scientific 1300 Series Class II, Type A2 - Horizontal ducting kit | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106874 | Fumigation kit for HPV for BSC Class II, Type A2, 5 feet -Thermo | 1 unit | N/A | Hydrogen Peroxide Vapor (HPV) fumigation kit to decontaminate the Biosafety cabinet 5 feet, 2 filters. This item includes: valves, top closing/plate, sealing kit and other accessories necessaries for the fumigation process. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 5 feet -Thermo | AMEX |  | 416.7 | USD | 0 | General Cargo | 10 |  |  |  | 0.06 | Thermo Scientific 1300 Series Class II, Type A2 - Fumigation kit | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106875 | Double Exhaust filter kit for BSC Class II, Type A2 5feet Thermo | 1 unit | N/A | Double exhaust Hepa filter HP14 for Biosafety cabinet 5 feet, 2 filters. Frame made of steel covered with coated paint or preferably of Stainless steel AISI 304. Able to accept double HEPA filter. · Frame dimension able to accept HEPA filter size of at least 90 x 45 cm; · Double HEPA filter, each: two Micro glass-fibers HEPA to block particles of at least 0.3 µm size and larger with an efficiency of at least 99.99%. Thi item includes: A double Exhaust filter Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 5 feet -Thermo | AMEX |  | 860.6 | USD | 0 | General Cargo | 7.5 |  |  |  | 0.045 | Thermo Scientific 1300 Series Class II, Type A2 - Exhaust Hepa filter HP14 | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106876 | Exhaust Hepa filter H14 for BSC Class II Type A2, 5 feet Thermo | 2 units | N/A | Exhaust Hepa Air filter H14 for Biosafety cabinet 5 feet, 2 filters. HEPA filters size of at least 90 x 45 cm; Class: H14 according with EN 1822:2009; · Micro glass-fibers HEPA to block particles of at least 0.3 µm size and larger with an efficiency of at least 99.99% This item includes: 2 exhaust Hepa Air filters. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 5 feet -Thermo | AMEX |  | 524.7 | USD | 0 | General Cargo | 5 |  |  |  | 0.03 | Thermo Scientific 1300 Series Class II, Type A2 - Exhaust Hepa filter H14 | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106877 | Main Hepa filter for BSC Class II, Type A2, 5 feet -Thermo | 1 unit | N/A | Main HEPA Filter for Biosafety cabinet 5 feet, 2 filters. Main recirculating filter; · HEPA filter size of at least 90 x 45 cm; Type H14 according with EN 1822:2009; · Micro glass-fibers HEPA to block particles of at least 0.3 µm size and larger with an efficiency of at least 99.99% Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 5 feet -Thermo | AMEX |  | 669.1 | USD | 0 | General Cargo | 5 |  |  |  | 0.03 | Thermo Scientific 1300 Series Class II, Type A2 - Main Hepa filter | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106868 | Installation and Training for Biosafety cabinet -Thermo | 1 unit | N/A | The installation includes: On-site installation of Biosafety cabinet, including: installation of the filters and connection to the electricity. All standard accessories, consumables, parts required for the proper operation and all standard tools, cleaning and lubrication material shall be included in the offer. Final internal testing of all pre-set programs and print-out of test cycle performance data (internal printer). Certification of installation and operation by the company's service engineer The training includes: A detailed instruction of the laboratory personnel on use, function and maintenance of the equipment (user training) as well as a comprehensive maintenance plan (log-book for daily, troubleshooting and weekly, monthly and quarterly maintenance checklist) | AMEX |  | 1400 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
6.3.4 Biosafety cabinet Class II Type A2, 6 feet
106775 | Biosafety cabinet (BSC) Class II, Type A2, 6 feet -Labo | 1 unit | N/A | Biosafety cabinet Class II Type A2 with panoramic front, side windows, electrical outlets, tabletops and light source for ultimate supervision and glare-free lighting V-shaped angled inflow grills allowing a comfortable working position 110 mm HEPA filters with efficiency at 99,999 % against 0,3 µm particles, ensuring a clean sterile work chamber environment. • Air velocity, vertical flow: 0,35 m/s; Air velocity, deviation: 10 ±%; 2 Hepa filters: H-14; • Air velocity, vertical flow: 0,35 m/s; Air velocity, deviation: 10 ±%; Down Flow rate: 1100 m3 /h; • Noise level, ISO6081: <54 dB(A); Light intensity variable: 0-2000 Lux; Window material: hardened/laminated glass; Cabinet material: polyester coated /AISI 304 stainless steel; • Max. 2 pcs Electrical outlets in each side panel (left and right); 3 whole in each side window for gas valves in each side window UV-light. UV -light mounted on the front top of the cabinet • Window in hardened glass with 3 holes for valves, Gas valve- G, Gas resettable solenoid valve USB port for communication and with computer inclusive software package in Excel • Table top> Pre filter grid with EU-3, painted version (Model 2016), Fluorescent dimmable light. • Electrical outlets - DE - 1 pair (1 left side & 1 right side)> Power supply> 230V 50/60Hz , Voltage 110V Please add to your request the UPS 3000 VA with its battery pack, the accessories (stand, vertical or horizontal ducting kit, fumigation kit for HPV, Double Exhaust, Exhaust Hepa H14, Standard filters) and installation and training. | TTM |  | 13579.75 | USD | 0 | General Cargo Large | 270 | 90 | 152 | 230 | 3.1464 | LaboGene ApS/Scanlaf Mars 1800, 2 filter | 90189084 | https://configurator.labogeneapps.com/media/1300/mars-1800-vers-2016.pdf
106776 | Stand for BSC Class II, Type A2, 6 feet -Labo | 1 unit | N/A | Adjustable stand for the Biosafety cabinet 6 feet with 2 filters. Floor stand adjustable for BSC; • Support stand 750 -800mm; Mounted on stabilizing adjustable feet; • Made of steel covered with epoxy coated paint, feet's made of stainless steel AISI 304. Needed if you do not have laboratory benches or the like where you can put the BSC on. | TTM |  | 627.2 | USD | 0 | General Cargo | 70 | 150 | 92 | 90 | 1.242 | LaboGene ApS/Scanlaf BSC Stand - Mars 1800, 2 filter. | 90189084 | https://configurator.labogeneapps.com/media/1300/mars-1800-vers-2016.pdf
106777 | Vertical ducting kit for BSC Class II, Type A2, 6 feet -Labo | 1 unit | N/A | Vertical ducting kit for Biosafety cabinet 6 feet with 2 filters. For a minimal room height of 250 cm • Automated lockable exhaust valve; The ducting kit can be connected to an external ventilation system; • Exhaust filter is 610 mm x 457 mm x 69 mm. Double exhaust filter possible. This item includes: Ducting kit including antiblow back valve straight Ø200, signals for external ventilation system and 3 -sensor option with pressure switch. Whether a horizontal or vertical exhaust air outlet is required must be checked by the user before ordering. If the connection to the ventilation system is directly above the BSC, a vertical outlet is required, otherwise a horizontal outlet. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 6 feet -Labo | TTM |  | 1591.4 | USD | 0 | General Cargo | 5 | 30 | 21 | 62 | 0.03906 | LaboGene ApS/Scanlaf BSC ducting kit - Mars 1800, 2 filter. | 90189084 | https://configurator.labogeneapps.com/media/1300/mars-1800-vers-2016.pdf
106778 | Horizontal ducting kit for BSC Class II, Type A2, 6 feet -Labo | 1 unit | N/A | Horizontal ducting kit for Biosafety cabinet 6 feet with 2 filters. For a minimal room height of 250 cm • Automated lockable exhaust valve; The ducting kit can be connected to an external ventilation system; • Exhaust filter is 610 mm x 457 mm x 69 mm. Double exhaust filter possible. This item includes: Ducting kit including antiblow back valve straight Ø200, signals for external ventilation system and 3 -sensor option with pressure switch. Whether a horizontal or vertical exhaust air outlet is required must be checked by the user before ordering. If the connection to the ventilation system is directly above the BSC, a vertical outlet is required, otherwise a horizontal outlet. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 6 feet -Labo | TTM |  | 1591.4 | USD | 0 | General Cargo | 5 | 30 | 21 | 62 | 0.03906 | LaboGene ApS/Scanlaf BSC ducting kit - Mars 1800, 2 filter. | 90189084 | https://configurator.labogeneapps.com/media/1300/mars-1800-vers-2016.pdf
106779 | Fumigation kit for HPV for BSC Class II, Type A2, 6 feet -Labo | 1 unit | N/A | Hydrogen Peroxide Vapor (HPV) fumigation kit to decontaminate the Biosafety cabinet 6 feet, 2 filters. This item includes: HVP fumigation plate kit with top plate, valves and sealing tape. If there is more than 1 BSC in the same laboratory, only 1 fumigation kit is required. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 6 feet -Labo | TTM |  | 1109.45 | USD | 0 | General Cargo | 18 | 115 | 61 | 43 | 0.30164 | LaboGene ApS/Scanlaf Fumigation kit - Mars 1800, 2 filter. | 90189084 | https://configurator.labogeneapps.com/media/1300/mars-1800-vers-2016.pdf
106780 | Double Exhaust filter kit for BSC Class II, Type A2, 6feet -Labo | 1 kit | N/A | Double exhaust Hepa filter HP14 for Biosafety cabinet 6 feet, 2 filters. Frame made of powder coated steel. • Able to accept double Hepa filter HP14; Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822); Double exhaust filter frame for Mars 1200 dimension: 915 mm x 457 mm x 115 mm. This item includes: A double Exhaust filter with 3 -sensor option with pressure switch and 1 spare filter. OPTIONAL ITEM: TO BE ORDERED ONLY IF REQUIRED BY THE EXISTING VENTILATION SYSTEM AND/OR LOCAL REGULATORY AUTHORITIES. THIS OPTION IS CURRENTLY ONLY COMMONLY USED IN THE UNITED KINGDOM (LOCAL REGULATIONS). Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 6 feet -Labo | TTM |  | 1115.6 | USD | 0 | General Cargo | 15 | 160 | 68 | 25 | 0.272 | LaboGene ApS/Scanlaf Double Exhaust filter - Mars 1800, 2 filter. | 90189084 | https://configurator.labogeneapps.com/media/1300/mars-1800-vers-2016.pdf
106781 | Exhaust Hepa filter H14 for BSC Class II, Type A2, 6 feet -Labo | 2 units | N/A | Exhaust Hepa Air filter H14 for Biosafety cabinet 6 feet, 2 filters. Class: H14 according with EN 1822:2009. • Hepa filters exhaust filter is 915 mm x 457 mm x 69 mm; • Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822) This item includes: 2 exhaust Hepa Air filters. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 6 feet -Labo | TTM |  | 1191.15 | USD | 0 | General Cargo | 8 | 91.5 | 45.7 | 69 | 0.28853 | LaboGene ApS/Scanlaf Exhaust filter filters - Mars 1800, 2 filter. | 90189084 | https://configurator.labogeneapps.com/media/1300/mars-1800-vers-2016.pdf
106782 | Standard Hepa filter for BSC Class II, Type A2, 6 feet -Labo | 1 unit | N/A | Standard HEPA Filter for Biosafety cabinet 6 feet, 2 filters. Main recirculating filter; • Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822); • Standard. Hepa H14 filters (99,99995% MPPs, according to EN1822); • Main filter dimension: 1225mm x 457 mm x 115 mm When ordering, please ensure that at least the same quantity of filters as the Exhaust Hepa filter H14 for BSC Class II, Type A2, 6 feet -Labo is ordered so that it can be replaced at the same time. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 6 feet -Labo | TTM |  | 1052.45 | USD | 0 | General Cargo | 15 | 152.5 | 45.7 | 11.5 | 0.080146 | LaboGene ApS/Scanlaf Standard. HEPA filter - Mars 1800, 2 filter | 90189084 | https://configurator.labogeneapps.com/media/1300/mars-1800-vers-2016.pdf
106489 | Installation and Training for Biosafety cabinet -Labo | 1 unit | N/A | The installation includes: On-site installation of Biosafety cabinet, including: installation of the filters and connection to the electricity. All standard accessories, consumables, parts required for the proper operation and all standard tools, cleaning and lubrication material shall be included in the offer. Final internal testing of all pre-set programs and print-out of test cycle performance data (internal printer). Certification of installation and operation by the company's service engineer The training includes: A detailed instruction of the laboratory personnel on use, function and maintenance of the equipment (user training) as well as a comprehensive maintenance plan (log-book for daily, troubleshooting and weekly, monthly and quarterly maintenance checklist) | TTM |  | 4500 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
106878 | Biosafety cabinet (BSC) Class II, Type A2, 6 feet -Thermo | 1 unit | N/A | Biosafety cabinet Class II Type A2 with panoramic front & side windows for ultimate supervision and glare-free lighting allowing a comfortable working position 110 mm HEPA filters with efficiency at 99,995 % against 0,3 µm particles, ensuring a clean sterile work chamber environment. • Load Bearing Capacity: 50 kg; Max. ambient temp during operation: 40 °C (104 °F); Min. ambient temp during operation:10 °C (50 °F); Max. humidity/ operation: 90, non-condensing % r.H. • Above room temp with window closed: < 2 °K; Max. humidity/ storage: 95 % r.H; SmartPort: 182.9 mm (7.2 in.) • Fluorescent Light; Air Inflow Rate: 207 - 517 cfm; Filter Efficiency: 99.995% MPPS; Filter Type: H14 HEPA EN 1822, 99.995% MPPSl; The recommended UPS to use wit this item is: UPS 3000 VA with its battery pack. | AMEX |  | 8896.5 | USD | 0 | General Cargo Large | 694 | 92.5 | 201 | 173 | 3.22 | Thermo Scientific 1300 Series Class II, Type A2 | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106879 | Base for BSC Class II, Type A2, 6 feet -Thermo | 1 unit | N/A | Adjustable base for the Biosafety cabinet 6 feet with 2 filters. Floor stand adjustable for BSC; • Adjustable height range: 800 mm up to 850 mm; Mounted on stabilizing adjustable feet; • Made of steel covered with epoxy coated paint, feet's made of stainless steel AISI 304. | AMEX |  | 836.8 | USD | 0 | General Cargo | 60 |  |  |  | 0.6 | Thermo Scientific 1300 Series Class II, Type A2 - Stand | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106880 | Vertical ducting kit for BSC Class II, Type A2, 6 feet -Thermo | 1 unit | N/A | Vertical ducting kit for Biosafety cabinet 6 feet with 2 filters. Ducting kit and anti blowback valve of at least 200mm; Automatic damper lockable exhaust valve or similar automatic system. · Exhaust duct to be connected to the external ventilation system in the laboratory; · Able to accept frame for double HEPA filter of at least 90 x 45 cm; This item includes: accessories and parts for the installation: thimble, exhaust duct, sleeves, clamps, others. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 6 feet -Thermo | AMEX |  | 834.3 | USD | 0 | General Cargo | 25 |  |  |  | 0.35 | Thermo Scientific 1300 Series Class II, Type A2 -Vertical ducting kit | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106881 | Horizontal ducting kit for BSC Class II, Type A2, 6 feet -Thermo | 1 unit | N/A | Horizontal ducting kit for Biosafety cabinet 6 feet with 2 filters. Ducting kit and anti blowback valve of at least 200mm; Automatic damper lockable exhaust valve or similar automatic system. · Exhaust duct to be connected to the external ventilation system in the laboratory; · Able to accept frame for double HEPA filter of at least 90 x 45 cm; This item includes: accessories and parts for the installation: thimble, exhaust duct, sleeves, clamps, others. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 6 feet -Thermo | AMEX |  | 834.3 | USD | 0 | General Cargo | 25 |  |  |  | 0.35 | Thermo Scientific 1300 Series Class II, Type A2 - Horizontal ducting kit | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106882 | Fumigation kit for HPV for BSC Class II, Type A2, 6 feet -Thermo | 1 unit | N/A | Hydrogen Peroxide Vapor (HPV) fumigation kit to decontaminate the Biosafety cabinet 6 feet, 2 filters. This item includes: valves, top closing/plate, sealing kit and other accessories necessaries for the fumigation process. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 6 feet -Thermo | AMEX |  | 416.7 | USD | 0 | General Cargo | 10 |  |  |  | 0.06 | Thermo Scientific 1300 Series Class II, Type A2 - Fumigation kit | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106883 | Double Exhaust filter kit for BSC Class II Type A2, 6feet Thermo | 1 unit | N/A | Exhaust Hepa filter HP14 for Biosafety cabinet 6 feet, 2 filters. Frame made of steel covered with coated paint or preferably of Stainless steel AISI 304. Able to accept double HEPA filter. · Frame dimension able to accept HEPA filter size of at least 90 x 45 cm; · Double HEPA filter, each: two Micro glass-fibers HEPA to block particles of at least 0.3 µm size and larger with an efficiency of at least 99.99%. Thi item includes: A double Exhaust filter Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 6 feet -Thermo | AMEX |  | 1076.5 | USD | 0 | General Cargo | 7.5 |  |  |  | 0.045 | Thermo Scientific 1300 Series Class II, Type A2 - Exhaust Hepa filter HP14 | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106884 | Exhaust Hepa filter H14 for BSC Class II Type A2, 6 feet -Thermo | 2 units | N/A | Exhaust Hepa Air filter H14 for Biosafety cabinet 6 feet, 2 filters. HEPA filters size of at least 90 x 45 cm; Class: H14 according with EN 1822:2009; · Micro glass-fibers HEPA to block particles of at least 0.3 µm size and larger with an efficiency of at least 99.99% This item includes: 2 exhaust Hepa Air filters. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 6 feet -Thermo | AMEX |  | 678.7 | USD | 0 | General Cargo | 5 |  |  |  | 0.03 | Thermo Scientific 1300 Series Class II, Type A2 - Exhaust Hepa filter H14 | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106885 | Main Hepa filter for BSC Class II, Type A2, 6 feet -Thermo | 1 unit | N/A | Main HEPA Filter for Biosafety cabinet 6 feet, 2 filters. Main recirculating filter. · HEPA filter size of at least 90 x 45 cm; Type H14 according with EN 1822:2009; · Micro glass-fibers HEPA to block particles of at least 0.3 µm size and larger with an efficiency of at least 99.99%. Please order same quantity as the Biosafety cabinet (BSC) Class II, Type A2, 6 feet -Thermo | AMEX |  | 791.9 | USD | 0 | General Cargo | 5 |  |  |  | 0.03 | Thermo Scientific 1300 Series Class II, Type A2 - Main Hepa filter | 84213920 | https://www.thermofisher.com/order/catalog/product/1323TS#/1323TS
106868 | Installation and Training for Biosafety cabinet -Thermo | 1 unit | N/A | The installation includes: On-site installation of Biosafety cabinet, including: installation of the filters and connection to the electricity. All standard accessories, consumables, parts required for the proper operation and all standard tools, cleaning and lubrication material shall be included in the offer. Final internal testing of all pre-set programs and print-out of test cycle performance data (internal printer). Certification of installation and operation by the company's service engineer The training includes: A detailed instruction of the laboratory personnel on use, function and maintenance of the equipment (user training) as well as a comprehensive maintenance plan (log-book for daily, troubleshooting and weekly, monthly and quarterly maintenance checklist) | AMEX |  | 1400 | USD | 0 | Service | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
6.4 LABORATORY FURNITURE
6.4.1 Laboratory furniture
106082 | Laboratory chair | 1 unit | N/A | All chairs and stools are available with a star base which is made from high strength sectional steel tubing, which is particularly robust and durable. The flat design prevents trip hazards. · Complete hygienic design, easy to clean and disinfect · Seat height adjustment by gas-lift; Seat adjustable height in the range of at least 450 to 600 mm; · Seat tilt adjustment; Backrest adjustable in height and inclination · Seat adjustable height through by gas-lift; SoftTouch PU foam, resistant against aggressive liquids | TTM |  | 201.65 | USD | 0 | General Cargo | 15.6 | 68 | 65 | 31 | 0.13702 | VWR International GmbH 132-0406 | 94013900 | https://de.vwr.com/store/product/en/11730526/laborstuhle-all-in-one?languageChanged=en
106083 | Mobile drawer cabinet, 4 drawers | 1 unit | N/A | White mobile laboratory cabinet with 4 drawers with metal rails. Metal structure with epoxy coated finish. Top coated with high chemical resistant material. Top drawer lockable; Castor diameter of at least 70 mm. · Top made of hardwood Medium Density Fibreboard (MDF) of at least 2 cm; Formaldehyde free; · Top surface high resistant to acids, solvents and cleaning agents commonly used in laboratory · Top surface withstands temperatures up to at least 90°C. | TTM |  | 851.3 | USD | 0 | General Cargo | 30 | 60 | 60 | 50 | 0.18 | Köttermann GmbH Mobile Drawer LMS/Köttermann - special production | 94029000 | Not available

## 3 Diagnostics Planning tool
Diagnostics Logistics & planning summary of your order
Supplier | General Cargo |  | Cool Cargo |  | Dangerous Cargo |  | Cost of services | Cost of products EXW per supplier | Product discount | Total cost with product discount
| Weight (kg) | Volume (m3) | Weight (kg) | Volume (m3) | Weight (kg) | Volume (m3)
Abbott Rapid Dx International
AMEX
BD International
Cepheid HBDC
Euromex
Hain LifeScience GMBH
HUMAN
The Medical Export Group
Molbio Diagnostics
Nipro Corporation
Oxford Immunotec
QIAGEN GmbH
Svizera Europe B.V.
TAUNS Laboratories
TTM
VWR
Subtotals Weight (kg) & Volume (m3) | [8 empty] | Total EXW cost | 0
[9 empty] | Total weight (kg) | 0
[9 empty] | Total volume (m3) | 0

## Calculations - Diagnostics (hidden)
Client planning |  | 0
Green light provision |  | 0
In-country logistics |  | 0
GDF Quote provision |  | 0
GDF Order preparation and delivery |  | 3
Client funds reception |  | 0
3
Client |  | 0
GDF |  | 3
Global fund |  | 0
3
Supplier |  | Supply lead time in weeks (Supplier+4)
Alere |  | #REF! | 8
BD |  | 8 | 8
Carl Zeiss |  | 6 | 6
Cepheid |  | 2 | 2
DNA Genotek |  | #REF! | 11
HAIN |  | 6 | 6
HUMAN |  | 11 | 11
MEG |  | 14 | 14
SD |  | 4 | 4
SVIZERA |  | 6 | 6
TAUNS |  | 8 | 8
TTM |  | 6 | 6
Item code | Supplier | Quantities | Transport type | Total volume | Total weight | Total cost | Filtered title
1. DETECTION AND RESISTANCE OF MTB | 0 | 0 | 0 | 0 | 0 | 0 | 1. DETECTION AND RESISTANCE OF MTB
1.1 MOLECULAR TESTING: POC / NEAR-POC | 0 | 0 | 0 | 0 | 0 | 0 | 1.1 MOLECULAR TESTING: POC / NEAR-POC
1.1.1 Truenat MTB equipment | 0 | 0 | 0 | 0 | 0 | 0 | 1.1.1 Truenat MTB equipment
106689 | Molbio Diagnostics | 0 | General Cargo UN 3481 UN 3481, Section II, PI -967 as per ICAO/IATA Lithium-ion batteries contained in equipment | 0 | 0 | 0 | 1.1.2 Truenat MTB reagents
106690 | Molbio Diagnostics | 0 | General Cargo UN 3481 UN 3481, Section II, PI -967 as per ICAO/IATA Lithium-ion batteries contained in equipment | 0 | 0 | 0 | 1.1.3 Truenat consumables
106691 | Molbio Diagnostics | 0 | General Cargo UN 3481 UN 3481, Section II, PI -967 as per ICAO/IATA Lithium-ion batteries contained in equipment | 0 | 0 | 0 | 1.1.4 Truenat services
1.1.2 Truenat MTB reagents | 0 | 0 | 0 | 0 | 0 | 0 | 1.1.5 TB-LAMP equipment
106692 | Molbio Diagnostics | 0 | CRT cargo (+15-+25°C) | 0 | 0 | 0 | 1.1.6 TB-LAMP reagents
106959 | Molbio Diagnostics | 0 | CRT cargo (+15-+25°C) | 0 | 0 | 0 | 1.1.7 TB-LAMP consumables
106694 | Molbio Diagnostics | 0 | CRT cargo (+15-+25°C) | 0 | 0 | 0 | 1.1.8 TB-LAMP services
106695 | Molbio Diagnostics | 0 | CRT cargo (+15-+25°C) | 0 | 0 | 0 | 1.1.11 GeneXpert Equipment (10 color)
106960 | Molbio Diagnostics | 0 | CRT cargo (+15-+25°C) | 0 | 0 | 0 | 1.1.12 GeneXpert Satellite (10 color)
106697 | Molbio Diagnostics | 0 | CRT cargo (+15-+25°C) | 0 | 0 | 0 | 1.1.13 GeneXpert reagents (6 and 10 color)
106698 | Molbio Diagnostics | 0 | CRT cargo (+15-+25°C) | 0 | 0 | 0 | 1.1.14 GeneXpert consumables (6 and 10 color)
106699 | Molbio Diagnostics | 0 | CRT cargo (+15-+25°C) | 0 | 0 | 0 | 1.1.15 GeneXpert services (same for 6 and 10 color)
106700 | Molbio Diagnostics | 0 | CRT cargo (+15-+25°C) | 0 | 0 | 0 | 1.2 RAPID ANTIGEN DETECTION
1.1.3 Truenat consumables | 0 | 0 | 0 | 0 | 0 | 0 | 1.2.1 Rapid antigen detection tests (LAM)
106886 | Molbio Diagnostics | 0 | CRT cargo (+15-+25°C) | 0 | 0 | 0 | 1.3 SMEAR MICROSCOPY
106834 | Molbio Diagnostics | 0 | General Cargo | 0 | 0 | 0 | 1.3.1 Smear microscopy kits
106955 | Molbio Diagnostics | 0 | General Cargo | 0 | 0 | 0 | 1.3.2 Smear microscopy equipment
1.1.4 Truenat services | 0 | 0 | 0 | 0 | 0 | 0 | 1.3.3 Smear microscopy reagents
106701 | Molbio Diagnostics | 0 | Service | #VALUE! | #VALUE! | 0 | 1.3.4 Smear microscopy consumables
106956 | Molbio Diagnostics | 0 | Service | #VALUE! | #VALUE! | 0 | 1.4 MOLECULR TESTING: CENTRALIZED
106957 | Molbio Diagnostics | 0 | Service | #VALUE! | #VALUE! | 0 | 1.4.1 LPA equipment
106958 | Molbio Diagnostics | 0 | Service | #VALUE! | #VALUE! | 0 | 1.4.2 LPA reagents
1.1.5 TB-LAMP equipment | 0 | 0 | 0 | 0 | 0 | 0 | 1.4.3 LPA consumables
106632 | Human | 0 | General Cargo | 0 | 0 | 0 | 1.4.6 BD MAX equipment
106633 | Human | 0 | General Cargo | 0 | 0 | 0 | 1.4.7 BD MAX reagents
106636 | Human | 0 | General Cargo | 0 | 0 | 0 | 1.4.8 BD MAX Consumables
106663 | Human | 0 | General Cargo | 0 | 0 | 0 | 1.4.9 Fluorotype equipment
1.1.6 TB-LAMP reagents | 0 | 0 | 0 | 0 | 0 | 0 | 1.4.10 Fluorotype reagents
106634 | Human | 0 | General Cargo | 0 | 0 | 0 | 1.4.11 Fluorotype consumables
106635 | Human | 0 | General Cargo | 0 | 0 | 0 | 1.4.12 Fluorotype services
1.1.7 TB-LAMP consumables | 0 | 0 | 0 | 0 | 0 | 0 | 1.5 SPECIMEN CULTURE AND DST
106643 | Human | 0 | General Cargo | 0 | 0 | 0 | 1.5.1 Solid culture equipment
1.1.8 TB-LAMP services | 0 | 0 | 0 | 0 | 0 | 0 | 1.5.2 Solid culture reagents
106637 | Human | 0 | Service | #VALUE! | #VALUE! | 0 | 1.5.3 Solid culture consumables
106638 | Human | 0 | Service | #VALUE! | #VALUE! | 0 | 1.5.4 Liquid culture equipment
106639 | Human | 0 | Service | #VALUE! | #VALUE! | 0 | 1.5.5 Liquid culture reagents
106640 | Human | 0 | Service | #VALUE! | #VALUE! | 0 | 1.5.6 Liquid culture consumables
1.1.11 GeneXpert Equipment (10 color) | 0 | 0 | 0 | 0 | 0 | 0 | 1.5.7 Solid and liquid culture DST: FLD pure substances
GXII-1-D-10C | Cepheid HBDC | 0 | General Cargo | 0 | 0 | 0 | 1.5.8 Solid and liquid culture DST: SLD pure substances
GXII-2-D-10C | Cepheid HBDC | 0 | General Cargo | 0 | 0 | 0 | 1.6 MTB IDENTIFICATION
GXII-2-D-10C (Trade in IV-2) | Cepheid HBDC | 0 | General Cargo | 0 | 0 | 0 | 1.6.1 MTB Identification Rapid test
GXIV-2-D-10C | Cepheid HBDC | 0 | General Cargo | 0 | 0 | 0 | 2. DETECTION OF LATENT TB
GXIV-2-D-10C (Trade-in IV-2) | Cepheid HBDC | 0 | General Cargo | 0 | 0 | 0 | 2.1 IGRA
GXIV-4-D-10C | Cepheid HBDC | 0 | General Cargo | 0 | 0 | 0 | 2.1.1 IGRA Equipment
GXIV-4-D-10C (Trade-in IV-2) | Cepheid HBDC | 0 | General Cargo | 0 | 0 | 0 | 2.1.2 IGRA Testing
GXIV-4-D-10C (Trade-in IV-4) | Cepheid HBDC | 0 | General Cargo | 0 | 0 | 0 | 3. DETECTION OF COVID-19
GXXVI-8-D-10C | Cepheid HBDC | 0 | General Cargo | 0 | 0 | 0 | 3.1 COVID-19
GXXVI-8-D-10C (Trade-in IV-2) | Cepheid HBDC | 0 | General Cargo | 0 | 0 | 0 | 3.1.1 COVID-19 testing
GXXVI-8-D-10C (Trade-in IV-4) | Cepheid HBDC | 0 | General Cargo | 0 | 0 | 0 | 4. SAMPLE AND MEDIA PREPARATION
GXXVI-8-D-10C (Trade-in XVI-8) | Cepheid HBDC | 0 | General Cargo | 0 | 0 | 0 | 4.1 SPECIMEN TRANPORTATION AND PROCESSING
GXXVI-16-D-10C | Cepheid HBDC | 0 | General Cargo | 0 | 0 | 0 | 4.1.1 Specimen transport consumables
GXXVI-16-D-10C (rade-in IV-2) | Cepheid HBDC | 0 | General Cargo | 0 | 0 | 0 | 4.1.2 Specimen processing equipment
GXXVI-16-D-10C (Trade-in IV-4) | Cepheid HBDC | 0 | General Cargo | 0 | 0 | 0 | 4.1.3 Specimen processing reagents
GXXVI-16-D-10C (Trade-in XVI-8) | Cepheid HBDC | 0 | General Cargo | 0 | 0 | 0 | 4.1.4 Specimen processing consumables
GXXVI-16-D-10C (Trade-in XVI-16) | Cepheid HBDC | 0 | General Cargo | 0 | 0 | 0 | 4.2 MEDIA PREPARATION
GXII-1-L-10C | Cepheid HBDC | 0 | General cargo UN 3481 | 0 | 0 | 0 | 4.2.1 Media preparation equipment
GXII-2-L-10C | Cepheid HBDC | 0 | General cargo UN 3481 | 0 | 0 | 0 | 4.2.2 Media preparation plasticware
GXII-2-L-10C (Trade-in IV-2) | Cepheid HBDC | 0 | General cargo UN 3481 | 0 | 0 | 0 | 4.2.3 Media preparation glassware
GXIV-2-L-10C | Cepheid HBDC | 0 | General cargo UN 3481 | 0 | 0 | 0 | 4.2.4 Media preparation consumables
GXIV-2-L-10C (Trade-in IV-2) | Cepheid HBDC | 0 | General cargo UN 3481 | 0 | 0 | 0 | 5. INFECTION PREVENTION PRODUCTS
GXIV-4-L-10C | Cepheid HBDC | 0 | General cargo UN 3481 | 0 | 0 | 0 | 5.1 PERSONAL PROTECTIVE EQUIPMENT (PPE)
GXIV-4-L-10C (Trade-in IV-2) | Cepheid HBDC | 0 | General cargo UN 3481 | 0 | 0 | 0 | 5.1.1 PPE consumables
GXIV-4-L-10C (Trade-in IV-4) | Cepheid HBDC | 0 | General cargo UN 3481 | 0 | 0 | 0 | 5.2 BIOSAFETY PRODUCTS
GXXVI-8-L-10C | Cepheid HBDC | 0 | General cargo UN 3481 | 0 | 0 | 0 | 5.2.1 Biosafety consumables
GXXVI-8-L-10C (Trade-in IV-2) | Cepheid HBDC | 0 | General cargo UN 3481 | 0 | 0 | 0 | 5.2.2 Disinfectants
GXXVI-8-L-10C (Trade-in IV-4) | Cepheid HBDC | 0 | General cargo UN 3481 | 0 | 0 | 0 | 5.3 WASTE MANAGEMENT
GXXVI-8-L-10C (Trade-in XVI-8) | Cepheid HBDC | 0 | General cargo UN 3481 | 0 | 0 | 0 | 5.3.1 Waste management reagents
GXXVI-16-L-10C | Cepheid HBDC | 0 | General cargo UN 3481 | 0 | 0 | 0 | 5.3.2 Waste management consumables
GXXVI-16-L-10C (Trade-in IV-2) | Cepheid HBDC | 0 | General cargo UN 3481 | 0 | 0 | 0 | 6. LABORATORY EQUIPMENT
GXXVI-16-L-10C (Trade-in IV-4) | Cepheid HBDC | 0 | General cargo UN 3481 | 0 | 0 | 0 | 6.1 UPS AND SOLAR PANELS
GXXVI-16-L-10C (Trade-in XVI-8) | Cepheid HBDC | 0 | General cargo UN 3481 | 0 | 0 | 0 | 6.1.1 UPSs
GXXVI-16-L-10C (Trade-in XVI-16) | Cepheid HBDC | 0 | General cargo UN 3481 | 0 | 0 | 0 | 6.1.2 Solar panels
1.1.12 GeneXpert Satellite (10 color) | 0 | 0 | 0 | 0 | 0 | 0 | 6.2 AUTOCLAVES
900-0793 | Cepheid HBDC | 0 | General Cargo | 0 | 0 | 0 | 6.2.1 Horizontal Autoclaves
DAISYKIT | Cepheid HBDC | 0 | General cargo | 0 | 0 | 0 | 6.2.2 Vertical Autoclaves
1.1.13 GeneXpert reagents (6 and 10 color) | 0 | 0 | 0 | 0 | 0 | 0 | 6.2.3 Autoclave water supply
GXMTB/RIF-ULTRA-50 | Cepheid HBDC | 0 | Cool cargo (+2-+8°C) UN 2920 – CLASS 8(3)) in excepted quantity (EQ) | 0 | 0 | 0 | 6.3 BIOSAFETY CABINETS
GXMTB/XDR-10 | Cepheid HBDC | 0 | Cool cargo (+2-+8°C) UN 2920 – CLASS 8(3)) in excepted quantity (EQ) | 0 | 0 | 0 | 6.3.1 Biosafety cabinet Class II - 3 filters
PCK-MTB/XDR-50-L | Cepheid HBDC | 0 | Cool cargo (+2-+8°C) UN 2920 – CLASS 8(3)) in excepted quantity (EQ) | 0 | 0 | 0 | 6.3.2 Biosafety cabinet Class II Type A2, 4 feet
PCK-MTB/XDR-50-D | Cepheid HBDC | 0 | Cool cargo (+2-+8°C) UN 2920 – CLASS 8(3)) in excepted quantity (EQ) | 0 | 0 | 0 | 6.3.3 Biosafety cabinet Class II Type A2, 5 feet
1.1.14 GeneXpert consumables (6 and 10 color) | 0 | 0 | 0 | 0 | 0 | 0 | 6.3.4 Biosafety cabinet Class II Type A2, 6 feet
XPERTCHECKCE- 5 | Cepheid HBDC | 0 | Cool cargo (+2-+8°C) UN 2920 – CLASS 8(3)) in excepted quantity (EQ) | 0 | 0 | 0 | 6.4 LABORATORY FURNITURE
900-0508 | Cepheid HBDC | 0 | General Cargo | 0 | 0 | 0 | 6.4.1 Laboratory furniture
900-0491RS | Cepheid HBDC | 0 | General Cargo | 0 | 0 | 0 | #N/A
700-4510 | Cepheid HBDC | 0 | General Cargo | 0 | 0 | 0 | #N/A
900-0807 | Cepheid HBDC | 0 | General cargo | 0 | 0 | 0
CMPTR-DT-GXDX | Cepheid HBDC | 0 | General cargo | 0 | 0 | 0
CMPTR-LT-GXDX | Cepheid HBDC | 0 | General cargo | 0 | 0 | 0
1.1.15 GeneXpert services (same for 6 and 10 color) | 0 | 0 | 0 | 0 | 0 | 0
ACCESS_CARE (Tier 1) | Cepheid HBDC | 0 | N/A | #VALUE! | #VALUE! | 0
ACCESS_CARE (Tier 2) | Cepheid HBDC | 0 | N/A | #VALUE! | #VALUE! | 0
ACCESS_CARE (Tier 3) | Cepheid HBDC | 0 | N/A | #VALUE! | #VALUE! | 0
WXPLUS02Y1 (GXII-2) | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WXPLUS02Y3 (GXII-2) | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WXPLUS02Y1 (GXIV-2) | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WXPLUS02Y3 (GXIV-2) | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WXPLUS04Y1 | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WXPLUS04Y3 | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WXPLUS08Y1 | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WXPLUS08Y3 | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WXPLUS16Y1 | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WXPLUS16Y3 | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WX01RG12 | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WX02RG12 | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WX04RG12 | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WV08RG12 | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WV16RG12 | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WX01UP36 | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WX02UP36 | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WX04UP36 | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WV08UP36 | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WV16UP36 | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WX01RG36 | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WX02RG36 | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WX04RG36 | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WV08RG36 | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
WV16RG36 | Cepheid HBDC | 0 | Service | #VALUE! | #VALUE! | 0
1.2 RAPID ANTIGEN DETECTION | 0 | 0 | 0 | 0 | 0 | 0
1.2.1 Rapid antigen detection tests (LAM) | 0 | 0 | 0 | 0 | 0 | 0
106749 | Abbott Rapid Dx International | 0 | General cargo | 0 | 0 | 0
1.3 SMEAR MICROSCOPY | 0 | 0 | 0 | 0 | 0 | 0
1.3.1 Smear microscopy kits | 0 | 0 | 0 | 0 | 0 | 0
106938 | Svizera Europe B.V. | 0 | Dangerous cargo ZN carbol-fuchsin stain solution: UN2821 6.1 III ZN Acid-Alcohol decolorization solution: UN2924 3(8) II Chlorine disinfectant pellets: UN3077 9 III Immersion oil: UN 3082, 9, III | 0 | 0 | 0
106939 | Svizera Europe B.V. | 0 | Dangerous cargo ZN carbol-fuchsin stain solution: UN2821 6.1 III ZN Acid-Alcohol decolorization solution: UN2924 3(8) III | 0 | 0 | 0
106940 | Svizera Europe B.V. | 0 | Dangerous cargo Auramine stain solution: UN2821, 6.1, III Auramine Acid alcohol decolorisation solution: UN2924 3(8) II Chlorine disinfectant pellets: UN3077 9 III Immersion oil: UN 3082, 9, III | 0 | 0 | 0
106941 | Svizera Europe B.V. | 0 | Dangerous cargo UN 2821 & PG III | 0 | 0 | 0
106942 | Svizera Europe B.V. | 0 | General cargo | 0 | 0 | 0
1.3.2 Smear microscopy equipment | 0 | 0 | 0 | 0 | 0 | 0
106535 | The Medical Export Group | 0 | General Cargo | 0 | 0 | 0
106629 | The Medical Export Group | 0 | General Cargo | 0 | 0 | 0
106630 | The Medical Export Group | 0 | General Cargo | 0 | 0 | 0
106943 | Svizera Europe B.V. | 0 | General Cargo | 0 | 0 | 0
106944 | Svizera Europe B.V. | 0 | General Cargo | 0 | 0 | 0
106945 | Svizera Europe B.V. | 0 | General Cargo | 0 | 0 | 0
106950 | Euromex | 0 | General Cargo | 0 | 0 | 0
106951 | Euromex | 0 | General Cargo | 0 | 0 | 0
106952 | Euromex | 0 | Dangerous CAO Dangerous Cargo UN 3480, Cargo Aircraft Only (CAO) | 0 | 0 | 0
106979 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
106534 | 0 | 0 | 0 | 0 | 0 | 0
106980 | TTM | 0 | Dangerous cargo UN 2811, Class 6.1, PG III | 0 | 0 | 0
1.3.3 Smear microscopy reagents | TTM | 0 | Dangerous cargo UN 1170, Class 3, PG II | 0 | 0 | 0
106276 | TTM | 0 | Dangerous cargo UN 1671, Class 6.1, PG II | 0 | 0 | 0
106597 | Svizera Europe B.V. | 0 | General cargo UN 3082, 9, III (Excepted Quantity - Not restricted) | 0 | 0 | 0
106599 | 0 | 0 | 0 | 0 | 0 | 0
106627 | Svizera Europe B.V. | 0 | General cargo | 0 | 0 | 0
1.3.4 Smear microscopy consumables | Svizera Europe B.V. | 0 | General cargo | 0 | 0 | 0
106525 | Svizera Europe B.V. | 0 | General cargo | 0 | 0 | 0
106609 | Svizera Europe B.V. | 0 | General cargo | 0 | 0 | 0
106610 | Svizera Europe B.V. | 0 | General cargo | 0 | 0 | 0
106568 | Svizera Europe B.V. | 0 | General cargo | 0 | 0 | 0
106564 | Svizera Europe B.V. | 0 | General cargo | 0 | 0 | 0
106565 | Svizera Europe B.V. | 0 | General cargo | 0 | 0 | 0
106566 | Svizera Europe B.V. | 0 | General cargo | 0 | 0 | 0
106567 | Svizera Europe B.V. | 0 | General cargo | 0 | 0 | 0
106569 | Svizera Europe B.V. | 0 | General cargo | 0 | 0 | 0
106570 | Svizera Europe B.V. | 0 | General cargo | 0 | 0 | 0
106576 | Svizera Europe B.V. | 0 | General cargo | 0 | 0 | 0
106947 | Svizera Europe B.V. | 0 | General cargo | 0 | 0 | 0
106655 | Svizera Europe B.V. | 0 | General cargo | 0 | 0 | 0
106946 | TTM | 0 | General Cargo | 0 | 0 | 0
106571 | TTM | 0 | General Cargo | 0 | 0 | 0
106605 | 0 | 0 | 0 | 0 | 0 | 0
106606 | 0 | 0 | 0 | 0 | 0 | 0
1.4 MOLECULR TESTING: CENTRALIZED | Hain LifeScience GMBH | 0 | General Cargo | 0 | 0 | 0
1.4.1 LPA equipment | Hain LifeScience GMBH | 0 | General Cargo | 0 | 0 | 0
106050 | Hain LifeScience GMBH | 0 | General Cargo | 0 | 0 | 0
106064 | Hain LifeScience GMBH | 0 | General Cargo | 0 | 0 | 0
106438 | TTM | 0 | General Cargo | 0 | 0 | 0
106613 | TTM | 0 | General Cargo | 0 | 0 | 0
106732 | TTM | 0 | General Cargo | 0 | 0 | 0
106205 | TTM | 0 | General Cargo | 0 | 0 | 0
106048 | TTM | 0 | General Cargo | 0 | 0 | 0
106224 | TTM | 0 | General Cargo | 0 | 0 | 0
106225 | TTM | 0 | General Cargo | 0 | 0 | 0
106226 | TTM | 0 | General Cargo | 0 | 0 | 0
106227 | TTM | 0 | General Cargo | 0 | 0 | 0
106228 | TTM | 0 | General Cargo | 0 | 0 | 0
106229 | 0 | 0 | 0 | 0 | 0 | 0
106219 | Hain LifeScience GMBH | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
1.4.2 LPA reagents | Hain LifeScience GMBH | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106512 | Hain LifeScience GMBH | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106513 | Hain LifeScience GMBH | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106590 | Hain LifeScience GMBH | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106888 | Hain LifeScience GMBH | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106889 | Hain LifeScience GMBH | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106477 | Hain LifeScience GMBH | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106890 | TTM | 0 | General Cargo | 0 | 0 | 0
106439 | 0 | 0 | 0 | 0 | 0 | 0
106254 | TTM | 0 | General Cargo | 0 | 0 | 0
1.4.3 LPA consumables | TTM | 0 | General Cargo | 0 | 0 | 0
106053 | TTM | 0 | General Cargo | 0 | 0 | 0
106054 | TTM | 0 | General Cargo | 0 | 0 | 0
106055 | TTM | 0 | General Cargo | 0 | 0 | 0
106056 | TTM | 0 | General Cargo | 0 | 0 | 0
106057 | TTM | 0 | General Cargo | 0 | 0 | 0
106059 | TTM | 0 | General Cargo | 0 | 0 | 0
106060 | TTM | 0 | General Cargo | 0 | 0 | 0
106218 | TTM | 0 | General Cargo | 0 | 0 | 0
106554 | TTM | 0 | General Cargo | 0 | 0 | 0
106555 | TTM | 0 | General Cargo | 0 | 0 | 0
106388 | TTM | 0 | General Cargo | 0 | 0 | 0
106389 | TTM | 0 | General Cargo | 0 | 0 | 0
106390 | TTM | 0 | General Cargo | 0 | 0 | 0
106061 | TTM | 0 | General Cargo | 0 | 0 | 0
106396 | TTM | 0 | General Cargo | 0 | 0 | 0
106398 | TTM | 0 | General Cargo | 0 | 0 | 0
106412 | TTM | 0 | General Cargo | 0 | 0 | 0
106068 | TTM | 0 | General Cargo | 0 | 0 | 0
106446 | TTM | 0 | General Cargo | 0 | 0 | 0
106384 | TTM | 0 | General Cargo | 0 | 0 | 0
106340 | 0 | 0 | 0 | 0 | 0 | 0
106741 | BD International | 0 | General Cargo - large | 0 | 0 | 0
1.4.6 BD MAX equipment | BD International | 0 | Service | #VALUE! | #VALUE! | 0
106919 | 0 | 0 | 0 | 0 | 0 | 0
106920 | BD International | 0 | General Cargo | 0 | 0 | 0
1.4.7 BD MAX reagents | BD International | 0 | Dangerous cargo UN1824 Class 8; PG II | 0 | 0 | 0
106921 | 0 | 0 | 0 | 0 | 0 | 0
106922 | BD International | 0 | General Cargo | 0 | 0 | 0
106978 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
1.4.8 BD MAX Consumables | BD International | 0 | General Cargo | 0 | 0 | 0
106923 | BD International | 0 | General Cargo | 0 | 0 | 0
106925 | BD International | 0 | General Cargo | 0 | 0 | 0
106926 | BD International | 0 | General Cargo | 0 | 0 | 0
106927 | BD International | 0 | General Cargo | 0 | 0 | 0
106928 | BD International | 0 | General Cargo | 0 | 0 | 0
106929 | BD International | 0 | General Cargo | 0 | 0 | 0
106930 | BD International | 0 | General Cargo | 0 | 0 | 0
106931 | BD International | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106932 | BD International | 0 | General Cargo | 0 | 0 | 0
106685 | 0 | 0 | 0 | 0 | 0 | 0
106933 | Hain LifeScience GMBH | 0 | General Cargo | 0 | 0 | 0
1.4.9 Fluorotype equipment | Hain LifeScience GMBH | 0 | General Cargo | 0 | 0 | 0
106961 | Hain LifeScience GMBH | 0 | General Cargo | 0 | 0 | 0
106962 | Hain LifeScience GMBH | 0 | General Cargo | 0 | 0 | 0
106963 | Hain LifeScience GMBH | 0 | General Cargo | 0 | 0 | 0
106964 | 0 | 0 | 0 | 0 | 0 | 0
106965 | Hain LifeScience GMBH | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
1.4.10 Fluorotype reagents | Hain LifeScience GMBH | 0 | General cargo | 0 | 0 | 0
106966 | Hain LifeScience GMBH | 0 | Dangerous cargo UN 3316, class 9; PG II | 0 | 0 | 0
106967 | Hain LifeScience GMBH | 0 | Dangerous cargo UN 1845. Dry ice (-80 to -18°C) | 0 | 0 | 0
106968 | 0 | 0 | 0 | 0 | 0 | 0
106969 | Hain LifeScience GMBH | 0 | General Cargo | 0 | 0 | 0
1.4.11 Fluorotype consumables | Hain LifeScience GMBH | 0 | General Cargo | 0 | 0 | 0
106970 | 0 | 0 | 0 | 0 | 0 | 0
106971 | Hain LifeScience GMBH | 0 | Service | #VALUE! | #VALUE! | 0
1.4.12 Fluorotype services | Hain LifeScience GMBH | 0 | Service | #VALUE! | #VALUE! | 0
106972 | Hain LifeScience GMBH | 0 | Service | #VALUE! | #VALUE! | 0
106973 | 0 | 0 | 0 | 0 | 0 | 0
106974 | 0 | 0 | 0 | 0 | 0 | 0
1.5 SPECIMEN CULTURE AND DST | AMEX | 0 | General Cargo | 0 | 0 | 0
1.5.1 Solid culture equipment | AMEX | 0 | General Cargo | 0 | 0 | 0
106843 | TTM | 0 | General Cargo Large | 0 | 0 | 0
106844 | TTM | 0 | General Cargo | 0 | 0 | 0
106247 | 0 | 0 | 0 | 0 | 0 | 0
106248 | BD International | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
1.5.2 Solid culture reagents | BD International | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106038 | BD International | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106039 | BD International | 0 | General Cargo | 0 | 0 | 0
106559 | BD International | 0 | General Cargo | 0 | 0 | 0
106614 | BD International | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106615 | TTM | 0 | General Cargo | 0 | 0 | 0
106620 | TTM | 0 | General Cargo | 0 | 0 | 0
106259 | TTM | 0 | General Cargo | 0 | 0 | 0
106268 | TTM | 0 | Dangerous cargo UN 2811, Class 6.1, PG III | 0 | 0 | 0
106265 | TTM | 0 | General Cargo | 0 | 0 | 0
106267 | TTM | 0 | General Cargo | 0 | 0 | 0
106279 | 0 | 0 | 0 | 0 | 0 | 0
106303 | TTM | 0 | General Cargo | 0 | 0 | 0
1.5.3 Solid culture consumables | TTM | 0 | General Cargo | 0 | 0 | 0
106235 | TTM | 0 | General Cargo | 0 | 0 | 0
106234 | TTM | 0 | General Cargo | 0 | 0 | 0
106714 | TTM | 0 | General Cargo | 0 | 0 | 0
106098 | TTM | 0 | General Cargo | 0 | 0 | 0
106739 | TTM | 0 | General Cargo | 0 | 0 | 0
106339 | TTM | 0 | General Cargo | 0 | 0 | 0
106849 | TTM | 0 | General Cargo | 0 | 0 | 0
106432 | 0 | 0 | 0 | 0 | 0 | 0
106496 | BD International | 0 | General Cargo Large | 0 | 0 | 0
1.5.4 Liquid culture equipment | BD International | 0 | General Cargo | 0 | 0 | 0
106013 | BD International | 0 | Service | #VALUE! | #VALUE! | 0
106596 | BD International | 0 | General Cargo | 0 | 0 | 0
106408 | TTM | 0 | General Cargo | 0 | 0 | 0
106016 | TTM | 0 | General Cargo | 0 | 0 | 0
106451 | TTM | 0 | General Cargo | 0 | 0 | 0
106475 | TTM | 0 | General Cargo | 0 | 0 | 0
106085 | TTM | 0 | General Cargo | 0 | 0 | 0
106115 | 0 | 0 | 0 | 0 | 0 | 0
106180 | BD International | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
1.5.5 Liquid culture reagents | BD International | 0 | General Cargo | 0 | 0 | 0
106273 | 0 | 0 | 0 | 0 | 0 | 0
106558 | BD International | 0 | General Cargo | 0 | 0 | 0
1.5.6 Liquid culture consumables | BD International | 0 | General Cargo | 0 | 0 | 0
106272 | BD International | 0 | General Cargo | 0 | 0 | 0
106918 | BD International | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106026 | BD International | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106027 | BD International | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106028 | BD International | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106623 | BD International | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106582 | BD International | 0 | General Cargo | 0 | 0 | 0
106033 | BD International | 0 | Dangerous cargo UN 1824, Class 8, PG III | 0 | 0 | 0
106036 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
106037 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
106021 | BD International | 0 | General Cargo | 0 | 0 | 0
106020 | BD International | 0 | General Cargo | 0 | 0 | 0
106019 | BD International | 0 | General Cargo | 0 | 0 | 0
106018 | BD International | 0 | General Cargo | 0 | 0 | 0
106022 | BD International | 0 | General Cargo | 0 | 0 | 0
106616 | BD International | 0 | General Cargo | 0 | 0 | 0
106617 | BD International | 0 | General Cargo | 0 | 0 | 0
106618 | BD International | 0 | General Cargo | 0 | 0 | 0
106473 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
106474 | BD International | 0 | General Cargo | 0 | 0 | 0
106452 | BD International | 0 | General Cargo | 0 | 0 | 0
106468 | BD International | 0 | General Cargo | 0 | 0 | 0
106181 | TTM | 0 | General Cargo | 0 | 0 | 0
106209 | TTM | 0 | General Cargo | 0 | 0 | 0
106550 | TTM | 0 | General Cargo | 0 | 0 | 0
106241 | TTM | 0 | General Cargo | 0 | 0 | 0
106715 | TTM | 0 | General Cargo | 0 | 0 | 0
106822 | TTM | 0 | General Cargo | 0 | 0 | 0
106823 | TTM | 0 | General Cargo | 0 | 0 | 0
1.5.7 Solid and liquid culture DST: FLD pure substances | TTM | 0 | General Cargo | 0 | 0 | 0
106308 | 0 | 0 | 0 | 0 | 0 | 0
106309 | TTM | 0 | General Cargo | 0 | 0 | 0
106310 | TTM | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
1.5.8 Solid and liquid culture DST: SLD pure substances | TTM | 0 | General Cargo | 0 | 0 | 0
106584 | 0 | 0 | 0 | 0 | 0 | 0
106586 | BD International | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106587 | BD International | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106588 | BD International | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106975 | BD International | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106976 | BD International | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106977 | TTM | 0 | General Cargo | 0 | 0 | 0
106560 | TTM | 0 | General Cargo | 0 | 0 | 0
106653 | TTM | 0 | General cargo | 0 | 0 | 0
106654 | TTM | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106318 | TTM | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106313 | TTM | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106316 | TTM | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106315 | TTM | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106317 | TTM | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106319 | 0 | 0 | 0 | 0 | 0 | 0
1.6 MTB IDENTIFICATION | 0 | 0 | 0 | 0 | 0 | 0
1.6.1 MTB Identification Rapid test | TAUNS Laboratories | 0 | General Cargo | 0 | 0 | 0
106611 | TAUNS Laboratories | 0 | General Cargo | 0 | 0 | 0
106953 | TAUNS Laboratories | 0 | General Cargo | 0 | 0 | 0
106612 | 0 | 0 | 0 | 0 | 0 | 0
2. DETECTION OF LATENT TB | 0 | 0 | 0 | 0 | 0 | 0
2.1 IGRA | 0 | 0 | 0 | 0 | 0 | 0
2.1.1 IGRA Equipment | AMEX | 0 | General Cargo | 0 | 0 | 0
106841 | AMEX | 0 | General Cargo | 0 | 0 | 0
106842 | AMEX | 0 | General Cargo | 0 | 0 | 0
106845 | TTM | 0 | General Cargo | 0 | 0 | 0
106857 | 0 | 0 | 0 | 0 | 0 | 0
2.1.2 IGRA Testing | QIAGEN GmbH | 0 | General Cargo | 0 | 0 | 0
106936 | QIAGEN GmbH | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106937 | QIAGEN GmbH | 0 | Dangerous and Cool cargo (+2 -+8°C) UN 3264 Class: 8, Packing group: III - Corrosive liquid, acidic, inorganic, n.o.s. (sulfuric acid) | 0 | 0 | 0
106671 | QIAGEN GmbH | 0 | Dangerous and Cool cargo (+2 -+8°C) UN 3264 Class: 8, Packing group: III - Corrosive liquid, acidic, inorganic, n.o.s. (sulfuric acid) | 0 | 0 | 0
106672 | QIAGEN GmbH | 0 | CRT cargo (+15-+25°C) | 0 | 0 | 0
106673 | QIAGEN GmbH | 0 | CRT cargo (+15-+25°C) | 0 | 0 | 0
106674 | QIAGEN GmbH | 0 | CRT cargo (+15-+25°C) | 0 | 0 | 0
106675 | QIAGEN GmbH | 0 | CRT cargo (+15-+25°C) | 0 | 0 | 0
106676 | Oxford Immunotec | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106665 | Oxford Immunotec | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106667 | Oxford Immunotec | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106668 | Oxford Immunotec | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106669 | Oxford Immunotec | 0 | Cool cargo (+2-+8°C) | 0 | 0 | 0
106670 | 0 | 0 | 0 | 0 | 0 | 0
3. DETECTION OF COVID-19 | 0 | 0 | 0 | 0 | 0 | 0
3.1 COVID-19 | 0 | 0 | 0 | 0 | 0 | 0
3.1.1 COVID-19 testing | BD International | 0 | Keep cool 2-8°C | 0 | 0 | 0
106684 | BD International | 0 | General Cargo | 0 | 0 | 0
106685 | Cepheid HBDC | 0 | Cool cargo (+2-+8°C) UN 2920 – CLASS 8(3)) in excepted quantity (EQ) | 0 | 0 | 0
XP3SARS-COV2-10 | 0 | 0 | 0 | 0 | 0 | 0
4. SAMPLE AND MEDIA PREPARATION | 0 | 0 | 0 | 0 | 0 | 0
4.1 SPECIMEN TRANPORTATION AND PROCESSING | 0 | 0 | 0 | 0 | 0 | 0
4.1.1 Specimen transport consumables | TTM | 0 | General Cargo | 0 | 0 | 0
106736 | TTM | 0 | General Cargo | 0 | 0 | 0
106244 | 0 | 0 | 0 | 0 | 0 | 0
4.1.2 Specimen processing equipment | TTM | 0 | General Cargo | 0 | 0 | 0
106220 | TTM | 0 | General Cargo | 0 | 0 | 0
106222 | 0 | 0 | 0 | 0 | 0 | 0
4.1.3 Specimen processing reagents | TTM | 0 | General Cargo | 0 | 0 | 0
106283 | TTM | 0 | General Cargo | 0 | 0 | 0
106284 | TTM | 0 | General Cargo | 0 | 0 | 0
106287 | 0 | 0 | 0 | 0 | 0 | 0
4.1.4 Specimen processing consumables | TTM | 0 | General Cargo | 0 | 0 | 0
106237 | TTM | 0 | General Cargo | 0 | 0 | 0
106238 | TTM | 0 | General Cargo | 0 | 0 | 0
106239 | TTM | 0 | General Cargo | 0 | 0 | 0
106230 | TTM | 0 | General Cargo | 0 | 0 | 0
106231 | TTM | 0 | General Cargo | 0 | 0 | 0
106509 | TTM | 0 | General Cargo | 0 | 0 | 0
106355 | TTM | 0 | Dangerous cargo UN 1263-Paint, Class 3, PG II | 0 | 0 | 0
106738 | TTM | 0 | General Cargo | 0 | 0 | 0
106356 | 0 | 0 | 0 | 0 | 0 | 0
4.2 MEDIA PREPARATION | 0 | 0 | 0 | 0 | 0 | 0
4.2.1 Media preparation equipment | TTM | 0 | General Cargo | 0 | 0 | 0
106257 | TTM | 0 | General Cargo | 0 | 0 | 0
106099 | TTM | 0 | General Cargo | 0 | 0 | 0
106100 | TTM | 0 | General Cargo | 0 | 0 | 0
106136 | TTM | 0 | General Cargo | 0 | 0 | 0
106127 | TTM | 0 | General Cargo | 0 | 0 | 0
106137 | TTM | 0 | General Cargo | 0 | 0 | 0
106594 | TTM | 0 | General Cargo | 0 | 0 | 0
106595 | TTM | 0 | General Cargo | 0 | 0 | 0
106130 | TTM | 0 | General Cargo | 0 | 0 | 0
106562 | TTM | 0 | General Cargo | 0 | 0 | 0
106132 | TTM | 0 | General Cargo | 0 | 0 | 0
106720 | 0 | 0 | 0 | 0 | 0 | 0
4.2.2 Media preparation plasticware | TTM | 0 | General Cargo | 0 | 0 | 0
106157 | TTM | 0 | General Cargo | 0 | 0 | 0
106158 | TTM | 0 | General Cargo | 0 | 0 | 0
106201 | TTM | 0 | General Cargo | 0 | 0 | 0
106202 | TTM | 0 | General Cargo | 0 | 0 | 0
106204 | TTM | 0 | General Cargo | 0 | 0 | 0
106171 | TTM | 0 | General Cargo | 0 | 0 | 0
106172 | TTM | 0 | General Cargo | 0 | 0 | 0
106173 | TTM | 0 | General Cargo | 0 | 0 | 0
106174 | TTM | 0 | General Cargo | 0 | 0 | 0
106175 | TTM | 0 | General Cargo | 0 | 0 | 0
106402 | 0 | 0 | 0 | 0 | 0 | 0
4.2.3 Media preparation glassware | TTM | 0 | General Cargo | 0 | 0 | 0
106134 | TTM | 0 | General Cargo | 0 | 0 | 0
106135 | TTM | 0 | General Cargo | 0 | 0 | 0
106203 | TTM | 0 | General Cargo | 0 | 0 | 0
106152 | TTM | 0 | General Cargo | 0 | 0 | 0
106146 | TTM | 0 | General Cargo | 0 | 0 | 0
106147 | TTM | 0 | General Cargo | 0 | 0 | 0
106148 | TTM | 0 | General Cargo | 0 | 0 | 0
106149 | TTM | 0 | General Cargo | 0 | 0 | 0
106159 | TTM | 0 | General Cargo | 0 | 0 | 0
106160 | TTM | 0 | General Cargo | 0 | 0 | 0
106161 | TTM | 0 | General Cargo | 0 | 0 | 0
106162 | TTM | 0 | General Cargo | 0 | 0 | 0
106163 | TTM | 0 | General Cargo | 0 | 0 | 0
106164 | TTM | 0 | General Cargo | 0 | 0 | 0
106165 | TTM | 0 | General Cargo | 0 | 0 | 0
106166 | TTM | 0 | General Cargo | 0 | 0 | 0
106167 | TTM | 0 | General Cargo | 0 | 0 | 0
106168 | TTM | 0 | General Cargo | 0 | 0 | 0
106169 | TTM | 0 | General Cargo | 0 | 0 | 0
106170 | TTM | 0 | General Cargo | 0 | 0 | 0
106721 | TTM | 0 | General Cargo | 0 | 0 | 0
106177 | TTM | 0 | General Cargo | 0 | 0 | 0
106178 | TTM | 0 | General Cargo | 0 | 0 | 0
106179 | TTM | 0 | General Cargo | 0 | 0 | 0
106183 | TTM | 0 | General Cargo | 0 | 0 | 0
106184 | TTM | 0 | General Cargo | 0 | 0 | 0
106185 | TTM | 0 | General Cargo | 0 | 0 | 0
106186 | TTM | 0 | General Cargo | 0 | 0 | 0
106187 | 0 | 0 | 0 | 0 | 0 | 0
4.2.4 Media preparation consumables | TTM | 0 | General Cargo | 0 | 0 | 0
106112 | TTM | 0 | General Cargo | 0 | 0 | 0
106363 | TTM | 0 | General Cargo | 0 | 0 | 0
106150 | TTM | 0 | General Cargo | 0 | 0 | 0
106746 | TTM | 0 | General Cargo | 0 | 0 | 0
106415 | TTM | 0 | General Cargo | 0 | 0 | 0
106353 | TTM | 0 | General Cargo | 0 | 0 | 0
106354 | TTM | 0 | General Cargo | 0 | 0 | 0
106494 | TTM | 0 | General Cargo | 0 | 0 | 0
106063 | TTM | 0 | General Cargo | 0 | 0 | 0
106084 | TTM | 0 | General Cargo | 0 | 0 | 0
106086 | 0 | 0 | 0 | 0 | 0 | 0
5. INFECTION PREVENTION PRODUCTS | 0 | 0 | 0 | 0 | 0 | 0
5.1 PERSONAL PROTECTIVE EQUIPMENT (PPE) | 0 | 0 | 0 | 0 | 0 | 0
5.1.1 PPE consumables | TTM | 0 | General Cargo | 0 | 0 | 0
106915 | TTM | 0 | General Cargo | 0 | 0 | 0
106818 | TTM | 0 | General Cargo | 0 | 0 | 0
106819 | TTM | 0 | General Cargo | 0 | 0 | 0
106820 | TTM | 0 | General Cargo | 0 | 0 | 0
106723 | TTM | 0 | General Cargo | 0 | 0 | 0
106724 | TTM | 0 | General Cargo | 0 | 0 | 0
106725 | TTM | 0 | General Cargo | 0 | 0 | 0
106435 | TTM | 0 | General Cargo | 0 | 0 | 0
106436 | TTM | 0 | General Cargo | 0 | 0 | 0
106437 | TTM | 0 | General Cargo | 0 | 0 | 0
106726 | TTM | 0 | General Cargo | 0 | 0 | 0
106727 | 0 | 0 | 0 | 0 | 0 | 0
5.2 BIOSAFETY PRODUCTS | 0 | 0 | 0 | 0 | 0 | 0
5.2.1 Biosafety consumables | TTM | 0 | General Cargo | 0 | 0 | 0
106079 | TTM | 0 | General Cargo | 0 | 0 | 0
106711 | TTM | 0 | General Cargo | 0 | 0 | 0
106745 | TTM | 0 | General Cargo | 0 | 0 | 0
106722 | TTM | 0 | General Cargo | 0 | 0 | 0
106094 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
106349 | TTM | 0 | General Cargo | 0 | 0 | 0
106728 | TTM | 0 | General Cargo | 0 | 0 | 0
106445 | TTM | 0 | General Cargo | 0 | 0 | 0
106409 | TTM | 0 | General Cargo | 0 | 0 | 0
5.2.2 Disinfectants | 0 | 0 | 0 | 0 | 0 | 0
106821 | TTM | 0 | Dangerous cargo UN 3077, Class 9, PG III | 0 | 0 | 0
106373 | TTM | 0 | General Cargo | 0 | 0 | 0
106374 | TTM | 0 | Dangerous cargo UN 1987, Class 3, PG III | 0 | 0 | 0
106294 | TTM | 0 | Dangerous cargo UN 1903, Class 8, PG III | 0 | 0 | 0
106371 | TTM | 0 | Dangerous cargo UN 1987, Class 3, PG III | 0 | 0 | 0
106370 | TTM | 0 | Dangerous cargo UN 1903, Class 8, PG III | 0 | 0 | 0
106747 | TTM | 0 | General Cargo | 0 | 0 | 0
106378 | TTM | 0 | Dangerous cargo UN 1814, Class 8, PG II | 0 | 0 | 0
5.3 WASTE MANAGEMENT | 0 | 0 | 0 | 0 | 0 | 0
5.3.1 Waste management reagents | 0 | 0 | 0 | 0 | 0 | 0
106331 | TTM | 0 | Dangerous cargo UN 1671, Class 6.1, PG II | 0 | 0 | 0
5.3.2 Waste management consumables | 0 | 0 | 0 | 0 | 0 | 0
106253 | TTM | 0 | General Cargo | 0 | 0 | 0
106718 | TTM | 0 | General Cargo | 0 | 0 | 0
106256 | TTM | 0 | General Cargo | 0 | 0 | 0
6. LABORATORY EQUIPMENT | 0 | 0 | 0 | 0 | 0 | 0
6.1 UPS AND SOLAR PANELS | 0 | 0 | 0 | 0 | 0 | 0
6.1.1 UPSs | 0 | 0 | 0 | 0 | 0 | 0
106707 | TTM | 0 | General cargo UN 2800 Valve Regulated Lead Acid Battery: UN 2800 Class 8 Packing group not restricted as per special provision A48 / A67 | 0 | 0 | 0
106708 | TTM | 0 | General cargo UN 2800 Valve Regulated Lead Acid Battery: UN 2800 Class 8 Packing group not restricted as per special provision A48 / A67 | 0 | 0 | 0
106709 | TTM | 0 | General cargo UN 2800 Valve Regulated Lead Acid Battery: UN 2800 Class 8 Packing group not restricted as per special provision A48 / A67 | 0 | 0 | 0
106710 | TTM | 0 | General cargo UN 2800 Valve Regulated Lead Acid Battery: UN 2800 Class 8 Packing group not restricted as per special provision A48 / A67 | 0 | 0 | 0
6.1.2 Solar panels | 0 | 0 | 0 | 0 | 0 | 0
106493 | TTM | 0 | General Cargo | 0 | 0 | 0
6.2 AUTOCLAVES | 0 | 0 | 0 | 0 | 0 | 0
6.2.1 Horizontal Autoclaves | 0 | 0 | 0 | 0 | 0 | 0
106750 | TTM | 0 | General Cargo | 0 | 0 | 0
106751 | TTM | 0 | General Cargo | 0 | 0 | 0
106752 | TTM | 0 | General Cargo | 0 | 0 | 0
106753 | TTM | 0 | General Cargo | 0 | 0 | 0
106754 | TTM | 0 | General Cargo | 0 | 0 | 0
106490 | TTM | 0 | Service | #VALUE! | #VALUE! | 0
106808 | TTM | 0 | General Cargo | 0 | 0 | 0
106809 | TTM | 0 | General Cargo | 0 | 0 | 0
106810 | TTM | 0 | General Cargo | 0 | 0 | 0
106811 | TTM | 0 | General Cargo | 0 | 0 | 0
106812 | TTM | 0 | General Cargo | 0 | 0 | 0
106490 | TTM | 0 | Service | #VALUE! | #VALUE! | 0
6.2.2 Vertical Autoclaves | 0 | 0 | 0 | 0 | 0 | 0
106813 | TTM | 0 | General Cargo | 0 | 0 | 0
106814 | TTM | 0 | General Cargo | 0 | 0 | 0
106815 | TTM | 0 | General Cargo | 0 | 0 | 0
106816 | TTM | 0 | General Cargo | 0 | 0 | 0
106817 | TTM | 0 | General Cargo | 0 | 0 | 0
106490 | TTM | 0 | Service | #VALUE! | #VALUE! | 0
106824 | TTM | 0 | General Cargo | 0 | 0 | 0
106825 | TTM | 0 | General Cargo | 0 | 0 | 0
106826 | TTM | 0 | General Cargo | 0 | 0 | 0
106827 | TTM | 0 | General Cargo | 0 | 0 | 0
106828 | TTM | 0 | General Cargo | 0 | 0 | 0
106490 | TTM | 0 | Service | #VALUE! | #VALUE! | 0
106580 | TTM | 0 | General Cargo | 0 | 0 | 0
6.2.3 Autoclave water supply | 0 | 0 | 0 | 0 | 0 | 0
106756 | TTM | 0 | General Cargo | 0 | 0 | 0
6.3 BIOSAFETY CABINETS | 0 | 0 | 0 | 0 | 0 | 0
6.3.1 Biosafety cabinet Class II - 3 filters | 0 | 0 | 0 | 0 | 0 | 0
106784 | TTM | 0 | General Cargo Large | 0 | 0 | 0
106785 | TTM | 0 | General Cargo | 0 | 0 | 0
106786 | TTM | 0 | General Cargo | 0 | 0 | 0
106787 | TTM | 0 | General Cargo | 0 | 0 | 0
106788 | TTM | 0 | General Cargo | 0 | 0 | 0
106789 | TTM | 0 | General Cargo | 0 | 0 | 0
106790 | TTM | 0 | General Cargo | 0 | 0 | 0
106791 | TTM | 0 | General Cargo | 0 | 0 | 0
106489 | TTM | 0 | Service | #VALUE! | #VALUE! | 0
106792 | TTM | 0 | General Cargo Large | 0 | 0 | 0
106793 | TTM | 0 | General Cargo | 0 | 0 | 0
106794 | TTM | 0 | General Cargo | 0 | 0 | 0
106795 | TTM | 0 | General Cargo | 0 | 0 | 0
106796 | TTM | 0 | General Cargo | 0 | 0 | 0
106797 | TTM | 0 | General Cargo | 0 | 0 | 0
106798 | TTM | 0 | General Cargo | 0 | 0 | 0
106799 | TTM | 0 | General Cargo | 0 | 0 | 0
106489 | TTM | 0 | Service | #VALUE! | #VALUE! | 0
106800 | TTM | 0 | General Cargo Large | 0 | 0 | 0
106801 | TTM | 0 | General Cargo | 0 | 0 | 0
106802 | TTM | 0 | General Cargo | 0 | 0 | 0
106803 | TTM | 0 | General Cargo | 0 | 0 | 0
106804 | TTM | 0 | General Cargo | 0 | 0 | 0
106805 | TTM | 0 | General Cargo | 0 | 0 | 0
106806 | TTM | 0 | General Cargo | 0 | 0 | 0
106807 | TTM | 0 | General Cargo | 0 | 0 | 0
106489 | TTM | 0 | Service | #VALUE! | #VALUE! | 0
6.3.2 Biosafety cabinet Class II Type A2, 4 feet | 0 | 0 | 0 | 0 | 0 | 0
106757 | TTM | 0 | General Cargo Large | 0 | 0 | 0
106758 | TTM | 0 | General Cargo | 0 | 0 | 0
106759 | TTM | 0 | General Cargo | 0 | 0 | 0
106760 | TTM | 0 | General Cargo | 0 | 0 | 0
106761 | TTM | 0 | General Cargo | 0 | 0 | 0
106762 | TTM | 0 | General Cargo | 0 | 0 | 0
106763 | TTM | 0 | General Cargo | 0 | 0 | 0
106764 | TTM | 0 | General Cargo | 0 | 0 | 0
106489 | TTM | 0 | Service | #VALUE! | #VALUE! | 0
106860 | AMEX | 0 | General Cargo Large | 0 | 0 | 0
106861 | AMEX | 0 | General Cargo | 0 | 0 | 0
106862 | AMEX | 0 | General Cargo | 0 | 0 | 0
106863 | AMEX | 0 | General Cargo | 0 | 0 | 0
106864 | AMEX | 0 | General Cargo | 0 | 0 | 0
106865 | AMEX | 0 | General Cargo | 0 | 0 | 0
106866 | AMEX | 0 | General Cargo | 0 | 0 | 0
106867 | AMEX | 0 | General Cargo | 0 | 0 | 0
106868 | AMEX | 0 | Service | #VALUE! | #VALUE! | 0
6.3.3 Biosafety cabinet Class II Type A2, 5 feet | 0 | 0 | 0 | 0 | 0 | 0
106766 | TTM | 0 | General Cargo Large | 0 | 0 | 0
106767 | TTM | 0 | General Cargo | 0 | 0 | 0
106768 | TTM | 0 | General Cargo | 0 | 0 | 0
106769 | TTM | 0 | General Cargo | 0 | 0 | 0
106770 | TTM | 0 | General Cargo | 0 | 0 | 0
106771 | TTM | 0 | General Cargo | 0 | 0 | 0
106772 | TTM | 0 | General Cargo | 0 | 0 | 0
106773 | TTM | 0 | General Cargo | 0 | 0 | 0
106489 | TTM | 0 | Service | #VALUE! | #VALUE! | 0
106870 | AMEX | 0 | General Cargo Large | 0 | 0 | 0
106871 | AMEX | 0 | General Cargo | 0 | 0 | 0
106872 | AMEX | 0 | General Cargo | 0 | 0 | 0
106873 | AMEX | 0 | General Cargo | 0 | 0 | 0
106874 | AMEX | 0 | General Cargo | 0 | 0 | 0
106875 | AMEX | 0 | General Cargo | 0 | 0 | 0
106876 | AMEX | 0 | General Cargo | 0 | 0 | 0
106877 | AMEX | 0 | General Cargo | 0 | 0 | 0
106868 | AMEX | 0 | Service | #VALUE! | #VALUE! | 0
6.3.4 Biosafety cabinet Class II Type A2, 6 feet | 0 | 0 | 0 | 0 | 0 | 0
106775 | TTM | 0 | General Cargo Large | 0 | 0 | 0
106776 | TTM | 0 | General Cargo | 0 | 0 | 0
106777 | TTM | 0 | General Cargo | 0 | 0 | 0
106778 | TTM | 0 | General Cargo | 0 | 0 | 0
106779 | TTM | 0 | General Cargo | 0 | 0 | 0
106780 | TTM | 0 | General Cargo | 0 | 0 | 0
106781 | TTM | 0 | General Cargo | 0 | 0 | 0
106782 | TTM | 0 | General Cargo | 0 | 0 | 0
106489 | TTM | 0 | Service | #VALUE! | #VALUE! | 0
106878 | AMEX | 0 | General Cargo Large | 0 | 0 | 0
106879 | AMEX | 0 | General Cargo | 0 | 0 | 0
106880 | AMEX | 0 | General Cargo | 0 | 0 | 0
106881 | AMEX | 0 | General Cargo | 0 | 0 | 0
106882 | AMEX | 0 | General Cargo | 0 | 0 | 0
106883 | AMEX | 0 | General Cargo | 0 | 0 | 0
106884 | AMEX | 0 | General Cargo | 0 | 0 | 0
106885 | AMEX | 0 | General Cargo | 0 | 0 | 0
106868 | AMEX | 0 | Service | #VALUE! | #VALUE! | 0
6.4 LABORATORY FURNITURE | 0 | 0 | 0 | 0 | 0 | 0
6.4.1 Laboratory furniture | 0 | 0 | 0 | 0 | 0 | 0
106082 | TTM | 0 | General Cargo | 0 | 0 | 0
106083 | TTM | 0 | General Cargo | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0

## 4 Medical devices categories
| GDF TB Medical devices Product categories | Please click any category to view the full product specifications
1. TB SCREENING / TRIAGE
1.1 CHEST RADIOGRAPHY (CXR)
1.1.1 Ultraportable Digital X-ray systems
1.2 CAD TECHNOLOGIES
1.2.1 CAD software for automated interpretation of CXRs
1.3 ACCESSORIES
1.3.1 Solar panel for Delft light and CAD4TB box

## 5 Medical devices ordering list
|  | GDF TB Medical devices Ordering List May, 2024 |  |  |  | Quantity to order | Product price |  |  | Product transport type and packaging | [5 empty] | Product information |  | Products links
GDF code | Product Name | Pack size | Minimum remaining shelf life when requesting shelf life approval | Product technical specifications | Supplier |  | Indicative FCA Price | Currency | Total Cost | Transport type | Weight Kg | Depth cm | Width cm | Height cm | Volume m3 | Manufacturer product code | HS Code | Product picture, Product information link
1. TB SCREENING / TRIAGE
1.1 CHEST RADIOGRAPHY (CXR)
1.1.1 Ultraportable Digital X-ray systems
200015 | FDR Xair System Full kit (with accessories) (166679271) | 1 Kit | N/A | Lightweight and compact ultraportable X-ray system easy to transport, mount and maintain, running on light built-in lithium polymer batteries allowing an autonomy of around 100 shots. This item is composed of: 1. One (1) FDR Xair Generator XD2000 Portable E (16667927) 2. One (1) FDR Xair Generator Stand EX-647 (50063) 3. One (1) FDR Xair Detector D-EVO II DR-ID 1211 SE A E (16477023) 4. One (1) FDR Xair Detector Stand EX-647 (50063001) 5. One (1) FDR Xair Console advance DR-ID 300CL w/ Application software DX-CL (16697609) 6. One (1) FDR Xair System Carrying case (417014680) 7. One (1) FDR Xair Power bank (RP-PB054Pro) 8. One (1) Protective Lead Apron with Thyroid shield (UF111722) 9. One (1) Portable radiation warning sign (S-8154-PL-10) 10. Ten (10) Shock stickers (ShockWatch2_10G) 11. One (1) year warranty is included, starting on the day of successful installation and user acceptance 12. One (1) USB Type C converter to allow connection with the CAD4TB Box. For any purchase of this item, a FDR Xair Full Kit Installation and Training, a FDR Xair Full Kit Warranty Extension 1 year or 3 years, must be procured in the same quantity. | Fujifilm Corporation |  | 47000 | USD | 0 | Dangerous cargo (CAO): 1. All components (except the power bank): GC UN3481 PI966 Section II. 2. Power bank: Dangerous cargo UN 3480 - Cargo aircraft only (CAO). | 52.8 | 96 | 91 | 60 | 0.52 | 166679273 | 90221400 | www.fujifilm.com/products/medical/digital_radiography/fdr_xair/
200016 | FDR Xair Full Kit Installation and Training (16667927113) | 1 unit | N/A | Installation includes: - In-country FDR Xair System installation and calibration based on user requirements including the provision of a User acceptance report Training includes: - On-site 1-day theoretical and practical training of at least 6 operators on safety, utilization, transportation and maintenance, delivered in English or French - Provision of hard and digital copies of training materials for each trainee This item should be procured in the same quantities as the FDR Xair System Full kit | Fujifilm Corporation |  | 2000 | USD | 0 | N/A | N/A | N/A | N/A | N/A | N/A | 16667927113 | N/A | N/A
200017 | FDR Xair Full Kit Warranty Extension 1 year (16667927011) | 1 unit | N/A | This item extends the initial warranty period of one FDR Xair System for 1 year, covers all the kit components and accessories and includes: - Visits for preventive and corrective maintenance including provision and shipment to site of spare parts to be replaced, shipment to site, disposal of faulty parts, cost of replacement work and personnel transport arrangements - Software patches, upgrades and updates when relevant - Service provided by Fujifilm own technicians or their local agents - One session of theoretical and practical training of at least 2 operators / users per system on safety, utilization, transportation and maintenance, delivered in English or French. No country is excluded. This item should be procured in the same quantities as the FDR Xair System Full kit. This item can be procured together with the FDR Xair Full Kit Warranty Extension 3 years for a total of 5 years warranty period of the FDR Xair System. | Fujifilm Corporation |  | 5000 | USD | 0 | N/A | N/A | N/A | N/A | N/A | N/A | 16667927011 | N/A | N/A
200018 | FDR Xair Full Kit Warranty extension 3 years (16667927013) | 1 unit | N/A | This item extends the initial warranty period of one FDR Xair System for 3 years, covers all the kit components and accessories and includes: - Visits for preventive and corrective maintenance including provision and shipment to site of spare parts to be replaced, shipment to site, disposal of faulty parts, cost of replacement work and personnel transport arrangements - Software patches, upgrades and updates when relevant - Service provided by Fujifilm own technicians or their local agents - One session per year of theoretical and practical training of at least 2 operators / users per system on safety, utilization, transportation and maintenance, delivered in English or French. No country is excluded. This item should be procured in the same quantities as the FDR Xair System Full kit. This item can be procured together with the FDR Xair Full Kit Warranty Extension 1 year for a total of 5 years.warranty period of the FDR Xair System. | Fujifilm Corporation |  | 15000 | USD | 0 | N/A | N/A | N/A | N/A | N/A | N/A | 16667927013 | N/A | N/A
200006 | Delft Light Full Kit (with accessories) | 1 Kit | N/A | The system provides constant high image quality at the lowest possible dose and runs on rechargeable Li-ion batteries (for Generator and Detector) capable of autonomous operation of around 200 exposures / 8 hours. Total weight is 32kg (29kg for core system and 3kg for accessories). This item is composed of: - One Delft Light Generator TR90/20 and One Delft Light Generator Mobile Stand - One Delft Light Detector Canon CXDI 703 C Wired/Wireless and One Delft Light Detector VersariX Portable Panelhanger - One Delft Light Workstation Canon NE and One Delft Light Software Canon NE - One Delft Light Backpack and One Mobisun Pro Solar Panel This item includes following accessories: - One Protective Lead Apron ; 5 Portable radiation warning signs and 10 Shock stickers 1-year warranty is included (including batteries) starting on the day of successful installation and user acceptance or 15 months from the date of invoice whichever date comes first. The terms and conditions of the initial warranty are the same as the ones of 'Delft Light Full kit Warranty Extension 1 year". For countries where EU/US sanctions apply: a confirmation should be asked to the supplier to ensure delivery is possible For any purchase of this item, a Delft Light Full Kit Installation and Training must be procured in the same quantity. For any purchase of this item together with the CAD4TB Software Perpetual License 15-month Support & Maintenance, the CAD4TB box is provided free of charge. Three (3) remote refresher training sessions on safety, utilization, transportation, maintenance, and upgrades of the system for unlimited number of operators. Two (2) remote preventive maintenance sessions for unlimited number of operators based on 'Delft Light Acceptance Maintenance checklist' including the provision of a maintenance report. | Delft Imaging Systems B.V. |  | 66750 | USD | 0 | Dangerous cargo: 1. All components (except the Solar panel): GC UN3481 PI966 section I. 2. Only the solar panel included in the Delft Light full kit: Dangerous cargo UN3481 PI967 section I. | 45 | 120 | 40 | 50 | 0.25 | DL-FK-2 | 90221400 | https://www.delft.care/x-ray-systems/
200007 | Delft Light Full Kit Installation and Training | 1 unit | N/A | Installation includes: - Remote Delft Light System installation, safety, calibration, operation checks and "Acceptance Tests" based on user requirements including the provision of an 'Acceptance report' signed by the client Training includes: - Remote session of theoretical and practical training of up to 10 operators on Delft Light System safety, utilisation, transportation and maintenance, delivered in English or French - Provision of hard and digital copies of training materials and access to Delft eLearning online platform for each trainee Installation and training is done remotely in 2 days sessions, provided by trained and licensed staff either from Delft Imaging HQ (Netherlands), Delft Imaging official branches (Ghana and South Africa) or from duly authorized service provider - In case the service is requested by the client to be performed on-site, travel and accommodation expenses are included for the Delft Imaging personnel. For the procurement of ≥ 10 Delft Light Full kit, the installation and training on-site are recommended. This item should be procured in the same quantities as the Delft Light Full Kit | Delft Imaging Systems B.V. |  | 2500 | USD | 0 | N/A | N/A | N/A | N/A | N/A | N/A | DL-FK-START-1 | N/A | N/A
200008 | Delft Light Full Kit Warranty Extension 1 year | 1 unit | N/A | This item includes: - One remote session for preventive maintenance including the provision of a maintenance report - Software patches, upgrades, and updates when relevant - Delft Imaging helpdesk 24x7 availability for service calls ensuring issue resolution within 5 working days maximum from the date the issue was reported - Replacement / repair of any faulty component including associated transport costs. Batteries are excluded. - Service provided by trained and licensed staff either from Delft Imaging HQ (Netherlands), Delft Imaging official branches (Ghana and South Africa) or from duly authorized service provider - One 1-day remote session of theoretical and practical training of up to 2 operators on Delft Light System safety, utilisation, transportation and maintenance, delivered in English or French - Provision of hard and digital copies of training materials and access to Delft eLearning online platform for each trainee - Three (3) remote refresher training sessions on safety, utilization, transportation, maintenance, and upgrades of the system for unlimited number of operators. Remote refresher training is currently available in English, French, and Russian. Other languages might be available upon request. - Two (2) remote preventive maintenance sessions for unlimited number of operators based on 'Delft Light Acceptance Maintenance checklist' including the provision of a maintenance report. Remote preventive maintenance is currently available in English, French and Russian. Other languages might be available upon request This item should be procured in the same quantities as the Delft Light Full kit. It is not possible to procure multiple 1-year warranties at once for the same Delft Light Full Kit (with accessories). This item can be procured together with the Delft Light Full Ki Warranty Extension 3 years for a total of 5 years warranty period of the Delft Light System. | Delft Imaging Systems B.V. |  | 4460 | USD | 0 | N/A | N/A | N/A | N/A | N/A | N/A | DL-FK-WTY-EXT-1 | N/A | N/A
200009 | Delft Light Full Kit Warranty Extension 3 years | 1 unit | N/A | This item includes: - One remote session per year for preventive maintenance including the provision of a maintenance report - Software patches, upgrades and updates when relevant - 24/7 availability of Delft Imaging helpdesk for service calls ensuring issue resolution within 5 working days maximum from the date the issue was reported - Replacement / repair of any faulty component including associated transport costs. Batteries are excluded - Service provided by trained and licensed staff either from Delft Imaging HQ (Netherlands), Delft Imaging official branches (Ghana and South Africa) or from duly authorized service provider - One 1-day remote session per year of theoretical and practical training of up to 2 operators on Delft Light System safety, utilisation, transportation and maintenance, delivered in English or French - Provision of hard and digital copies of training materials and access to Delft eLearning online platform for each trainee - Three (3) remote refresher training sessions on safety, utilization, transportation, maintenance, and upgrades of the system for unlimited number of operators. Remote refresher training is currently available in English, French, and Russian. Other languages might be available upon request. - Two (2) remote preventive maintenance sessions for unlimited number of operators based on 'Delft Light Acceptance Maintenance checklist' including the provision of a maintenance report. Remote preventive maintenance is currently available in English, French and Russian. Other languages might be available upon request This item should be procured in the same quantities as the Delft Light Full kit. This item can be procured together with the Delft Light Full Ki Warranty Extension 1 year for a total of 5 years warranty period of the Delft Light System. | Delft Imaging Systems B.V. |  | 27834 | USD | 0 | N/A | N/A | N/A | N/A | N/A | N/A | DL-FK-WTY-EXT-3 | N/A | N/A
1.2 CAD TECHNOLOGIES
1.2.1 CAD software for automated interpretation of CXRs
200001 | InferRead DR Chest Software (Perpetual License) | 1 unit | N/A | Perpetual license open to use for unlimited users accessing the specified hardware with no usage limitation. Product maintenance period (Support Life Cycle) 10 years, then offline version can be used but online version will cease operation. Product able to process both Anterior-Posterior (AP) and Posterior-Anterior (PA) chest x-ray (CXR) images. Built-in anonymisation function. Able to communicate with PACS and RIS systems. Results display: heat map, continuous values for TB-probability and binary outcome (pre-set threshold score adjustable by user anytime). Compatibility with X-ray system: not vendor specific. Able to integrate all digital X-ray models generating CXR images according to DICOM format and transmits images via DICOM protocol. Data storage / backup: all data can be uploaded to a specific cloud / local server. Auto-clearing function Central server provided by Infervision with no additional cost (500GB or approximately 35,000 CXR images). For every 5 licenses purchased, Infervision provides a free cloud back-up server of 1TB (approximately 70,000 CXR images) up to 5 years. Requirements of hardware for offline version: CPU = Intel i7 6,850K processor and above; GPU = NVIDIA GeForce 1,080 and above or V100 and above; RAM = DDR4 is recommended with a capacity of at least 16G; solid state system hard disks with capacity of at least 120G; high-speed mechanical hard disk with a capacity of at least 1T and rotation speed of 7,200 rpm as the storage disk. Requirements of server for online version: CPU at least Intel i3, RAM at least 2GB, Hard disk 50GB; Browser Chrome Requirements for online version requiring a Central server within a specific territory: CPU Intel i7 at least 6,850K processor, GPU NVIDIA GeForce at least 1,080 or at least V100; RAM DDR4 recommended with a capacity of at least 16G, solid state hard disks with capacity of at least 120G, high-speed mechanical hard disk with a capacity of at least 1T and rotation speed of 7,200 rpm recommended as storage disk AND+E18 workstation with CPU Intel i3; RAM at least 2GB; Hard disk: 50GB; Browser Chrome 1-year of Support and Maintenance is included Minimum quantity to order: 5 software For any purchase of this item, an InferRead DR Chest Laptop Server Schenker Key 15 and an InferRead DR Chest Installation and Training, InferRead DR Chest Support and Maintenance Extension 1 year or 3 years must be procured in the same quantity. This perpetual Software will be delivered already installed in the hardware: InferRead DR Chest Laptop Server (Schenker Key 15). | Infervision Medical Technology Co. Ltd |  | 2700 | USD | 0 | N/A | N/A | N/A | N/A | N/A | N/A | INDRCHESTPL | N/A | https://global.infervision.com/product/5/
200002 | InferRead DR Chest Laptop Server (Schenker Key 15) | 1 unit | N/A | Hardware device (laptop) allowing the analysis of chest X-ray images from InferRead DR Chest Software to be processed offline and able to provide server / cloud data synchronization. - 1 x 15.6", Ultra HD OLED, 100% sRGB/DCI- 1x NVIDIA GeForce RTX 2070 Max-Q Refresh, 8 GB GD 1x Intel Core i7-10750H - 2 x 8 GB (16 GB) DDR4-3200 Samsung - 1 x 1 TB M.2 Samsung 970 EVO Plus, PCIe 3.0 x4, N 1x Intel Wi-Fi 6 AX200, Bluetooth 5.1 - 1 x Lighted Keyboard US (ISO) Data storage / backup: Embedded 3 Cell Polymer Battery Pack, 73WH, lasting up to 600 minutes (10 hours). Central server provided by Infervision with no additional cost (500GB or approximately 35,000 Chest X-ray images) 2-year warranty included. No country excluded as a final destination Minimum quantity to order: 5 units For any purchase of this item, an InferRead DR Chest Software (Perpetual License), an InferRead DR Chest Installation and Training, must be procured in the same quantity. For any purchase of this item, a Inferread DR Chest Installation and Training, a InferRead DR Chest Support and Maintenance Extension 1 year or 3 years, must be procured in the same quantity. | Infervision Medical Technology Co. Ltd |  | 2082 | USD | 0 | Dangerous cargo UN 3481 Class 9 | 2.5 | 49 | 41 | 10 | 0.02 | INHWSCHK15 | 847130 | https://www.schenker-tech.de/en/schenker-key-15-en
200003 | InferRead DR Chest Installation and Training | 1 unit | N/A | Remote Installation and Training of software (InferRead DR Chest Software) and hardware (InferRead DR Chest Laptop Server - Schenker Key 15) Installation includes: - Installation guidelines - Software operating environment and server configuration - Threshold calibration as per protocol and toolkit recommended per WHO Training includes: - 1-hour online training on software usage, administration and utilization to maximum 20 participants, delivered by 2 trainers in English or French - Provision of hard and digital copies of training material for each trainee This item should be procured in the same quantity as the InferRead DR Chest Software Perpetual License, the InferRead DR Chest Laptop Server and the InferRead DR Chest Support and Maintenance Extension 1 year or 3 years. | Infervision Medical Technology Co. Ltd |  | 248 | USD | 0 | N/A | N/A | N/A | N/A | N/A | N/A | INDRCHESTIT | N/A | N/A
200004 | InferRead DR Chest Support and Maintenance Extension 1 year | 1 unit | N/A | Extension of support, maintenance and warranty covering software (InferRead DR Chest Software Perpetual License) and hardware (InferRead DR Chest Laptop Server Schenker Key 15) for 1 year. - Free software upgrade, patches and update conducted remotely - Remote support for issue resolution of reported problem within one working day from the date the issue was reported - Corrective maintenance as required, provision and shipment to site of hardware to be replaced, disposal of faulty hardware, cost of replacement work, personnel transport arrangements - One session of 1-hour online training on software usage, administration and utilization to maximum 20 participants delivered by 2 trainers in English or French - Provision of hard and digital copies of training material for each trainee This item should be procured in the same quantity as the InferRead DR Chest Software Perpetual License, the InferRead DR Chest Laptop Server, InferRead DR Chest Installation and Training and the InferRead DR Chest Support and Maintenance Extension 1 year or 3 years. If this item is procured together with the InferRead DR Chest Software (Perpetual License) and the InferRead DR Chest Laptop Server, this item covers both software and hardware. If this item is procured separately from the InferRead DR Chest Software (Perpetual License) and the InferRead DR Chest Laptop Server, this item only covers the software: InferRead DR Chest Software Perpetual License. | Infervision Medical Technology Co. Ltd |  | 232 | USD | 0 | N/A | N/A | N/A | N/A | N/A | N/A | INDRCHEST1Y | N/A | N/A
200005 | InferRead DR Chest Support and Maintenance Extension 3 years | 1 unit | N/A | Extension of support, maintenance and warranty covering software (InferRead DR Chest Software Perpetual License) and hardware (InferRead DR Chest Laptop Server Schenker Key 15) for 3 years. - Free software upgrade, patches and update conducted remotely - Remote support for issue resolution of reported problem within one working day from the date the issue was reported - Corrective maintenance as required, provision and shipment to site of hardware to be replaced, disposal of faulty hardware, cost of replacement work, personnel transport arrangements - One session per year of 1-hour online training on software usage, administration and utilization to maximum 20 participants delivered by two 2 trainers in English or French - Provision of hard and one digital copies of training material for each trainee This item should be procured in the same quantity as the InferRead DR Chest Software Perpetual License, the InferRead DR Chest Laptop Server, InferRead DR Chest Installation and Training and the InferRead DR Chest Support and Maintenance Extension 1 year or 3 years. If this item is procured together with the InferRead DR Chest Software (Perpetual License) and the InferRead DR Chest Laptop Server, this item covers both software and hardware. If this item is procured separately from the InferRead DR Chest Software (Perpetual License) and the InferRead DR Chest Laptop Server, this item only covers the software: InferRead DR Chest Software Perpetual License. | Infervision Medical Technology Co. Ltd |  | 522 | USD | 0 | N/A | N/A | N/A | N/A | N/A | N/A | INDRCHEST3Y | N/A | N/A
200010 | CAD4TB Software Perpetual License 12-month Support & Maintenance | 1 unit | N/A | Software tool for detecting TB lung abnormalities in frontal chest radiograph images. One perpetual license (installed on a ‘software’ on the CAD4TBbox) providing offline and online access for one Ultraportable X-ray System to an unlimited number of Posterior-Anterior (PA) chest X-ray (CXR) readings. This item includes: - 5 viewing licenses; Options for data and images hosting and storage on local server, local cloud and/or supplier's server up to client's preference. Hosting and storage up to 300 images per license and per day offered by Delft on its servers (using Amazon Web Services) located in Ireland, India, South Africa and Brazil. - Option to automatically de-identify and (pseudo) anonymize patient's DICOM data; Performance: sensitivity of 90% at a specificity of 70% with AUC > 80% demonstrated by independent studies - Output: heatmap and 0-100 score translated into a binary (pos/neg) output based on an adjustable threshold; Compatibility: any digital X-ray system providing a DICOM image. Interoperability with PACS and RIS system through integrated API functionality - 1 year of Support and Maintenance is included, starting on the day of successful installation and user acceptance. - Three (3) remote refresher training sessions on software administration and utilization for unlimited number of operators/radiographers. Remote refresher training is currently available in English, French, and Russian. Other languages might be available upon request; Two (2) remote preventive maintenance sessions for preventive maintenance sessions for unlimited number of operators based on 'Delft Light Acceptance Maintenance checklist' including the provision of a maintenance report. Remote preventive maintenance is currently available in English, French and Russian. Other languages might be available upon request. - For countries where EU/US sanctions apply: a confirmation should be asked to the supplier to ensure delivery is possible. - If this item is purchased together with Delft Light Full Kit, this item should be replaced by CAD4TB Software Perpetual License 15-month Support & Maintenance to have 3 additional months of Support & Maintenance. - For any purchase of this item, a CAD4TB box, a CAD4TB Installation and Training, a CAD4TB Support and Maintenance Extension 1 year or 3 years must be procured in equal quantities. Upon expiration of the Support and Maintenance contract, the client will be able to continue to use the Software Perpetual License. But note that supplier cannot warrant the continued optimal/effective use of the CAD4TB software if it is no supported by a Support and Maintenance contract. Therefore, the clients must procure 1 or 3 years CAD4TB Support and Maintenance Extension that assures access to technical support and latest CAD4TB updates and upgrades that will further enhance the presumptive TB case finding impact of the intervention. | Delft Imaging Systems B.V. |  | 12750 | USD | 0 | N/A | N/A | N/A | N/A | N/A | N/A | CAD4TB-LICENSE-1 | N/A | https://www.delft.care/cad4tb/
200019 | CAD4TB Software Perpetual License 15-month Support & Maintenance | 1 unit | N/A | Software tool for detecting TB lung abnormalities in frontal chest radiograph images. One perpetual license (installed on a ‘software’ on the CAD4TBbox) providing offline and online access for one Ultraportable X-ray System to an unlimited number of Posterior-Anterior (PA) chest X-ray (CXR) readings. This item includes: - 5 viewing licenses; Options for data and images hosting and storage on local server, local cloud and/or supplier's server up to client's preference. Hosting and storage up to 300 images per license and per day offered by Delft on its servers (using Amazon Web Services) located in Ireland, India, South Africa and Brazil. - Option to automatically de-identify and (pseudo) anonymize patient's DICOM data; Performance: sensitivity of 90% at a specificity of 70% with AUC > 80% demonstrated by independent studies. - Output: heatmap and 0-100 score translated into a binary (pos/neg) output based on an adjustable threshold; Compatibility: any digital X-ray system providing a DICOM image. Interoperability with PACS and RIS system through integrated API functionality - 1 year of Support and Maintenance is included, starting on the day of successful installation and user acceptance. - Three (3) remote refresher training sessio+E22ns on software administration and utilization for unlimited number of operators/radiographers. Remote refresher training is currently available in English, French, and Russian. Other languages might be available upon request. - Two (2) remote preventive maintenance sessions for preventive maintenance sessions for unlimited number of operators based on 'Delft Light Acceptance Maintenance checklist' including the provision of a maintenance report. Remote preventive maintenance is currently available in English, French and Russian. Other languages might be available upon request. - For countries where EU/US sanctions apply: a confirmation should be asked to the supplier to ensure delivery is possible. - Please use this item in replacement of the CAD4TB Software Perpetual License 12-month Support & Maintenance only if your order includes the Delft Light Full Kit. - For any purchase of this item, a CAD4TB box will be provided free of charge. - For any purchase of this item, a CAD4TB Installation and Training, a CAD4TB Support and Maintenance Extension 1 year or 3 years must be procured in equal quantities. Upon expiration of the Support and Maintenance contract, the client will be able to continue to use the Software Perpetual License. But note that supplier cannot warrant the continued optimal/effective use of the CAD4TB software if it is no supported by a Support and Maintenance contract. Therefore, the clients must procure 1 or 3 years CAD4TB Support and Maintenance Extension that assures access to technical support and latest CAD4TB updates and upgrades that will further enhance the presumptive TB case finding impact of the intervention. | Delft Imaging Systems B.V. |  | 12750 | USD | 0 | N/A | N/A | N/A | N/A | N/A | N/A |  | N/A | https://www.delft.care/cad4tb/
200011 | CAD4TB box | 1 unit | N/A | Hardware device allowing the analysis of chest X-ray images from CAD4TB Software License to be processed offline and able to provide server / cloud data synchronization. - IntelNUC, i-5 Quad core - 8 GB RAM, 256 GB SSD (storage capacity approximately 20,000 images) - Integrated API functionality enables backup to be located on client's own server - No power autonomy: should be connected to a power source - Supplied with a rugged USB disk (1TB) for an offline backup facility and two (2) LAN cables to allow the connection with the CAD4TBbox - The CAD4TBbox will be delivered within the CAD4TB suitcase that protects the hardware components from the challenging weather conditions. - 1-year warranty is included, starting on the day of successful installation and user acceptance. For countries where EU/US sanctions apply: a confirmation should be asked to the supplier to ensure delivery is possible. Free of charge in the same quantity if your order includes a Delft Light Full Kit and a CAD4TB Software Perpetual License 15-month Support & Maintenance. For any purchase of this item, a CAD4TB Software Perpetual License, a CAD4TB Installation and Training, a CAD4TB Support and Maintenance Extension 1 year or 3 years must be procured in equal quantities. For connection to a FDR Xair System from Fujifilm, the client may add as an optional item a Mobisun Solar Panel (from Delft) for a full autonomy. CAD4TB box can be procured for any digital X-ray system that provides a “DICOM” image as an output. If the client wants to procure the CAD4TB box for an X-ray system already available in the country, Stop TB GDF can confirm compatibility on a case-by-case basis if the client shares anonymized DICOM images taken from their X-ray system. The Product Support Life Cycle of CAD4TB Box is five (5) years starting from the day of successful user acceptance. Beyond that period, the client has to procure a new CAD4TB box as the supplier cannot support the hardware for over 5 years (due to the unavailability of spare parts and change in the technology). | Delft Imaging Systems B.V. |  | 2750 | USD | 0 | General cargo UN3481 (PI 967 section II). Lithium battery contained in equipment. | 1 | 15 | 15 | 10 | 0.002 | CAD4TBbox-1 | 84718000 | https://www.delft.care/cad4tb/
200012 | CAD4TB Installation and Training | 1 unit | N/A | Remote Installation and Training of CAD4TB software (CAD4TB Software License) and hardware (CAD4TB box). Installation includes: - Software operating environment and server configuration - Remote installation (software and hardware) and software calibration study in the setting of intended use based on WHO protocol & toolkit Training includes: - One 1-day remote session of theoretical and practical training of up to 10 operators on software administration and utilization - Provision of hard and digital copies of training materials and access to Delft eLearning online platform for each trainee - The training is currently available in English, French, Russian, Spanish, Portuguese, or Arabic. Other languages might be available upon request. For any purchase of this item, a CAD4TB Box, a CAD4TB Software Perpetual License, a CAD4TB Support and Maintenance Extension 1 year or 3 years must be procured in equal quantities. | Delft Imaging Systems B.V. |  | 1150 | USD | 0 | N/A | N/A | N/A | N/A | N/A | N/A | CAD4TB-START-1 | N/A | N/A
200013 | CAD4TB Support and Maintenance Extension 1 year | 1 unit | N/A | Extension of Support, maintenance and warranty covering CAD4TB software (CAD4TB Software Perpetual License) and hardware (CAD4TB box) for 1 year. - Free software patches, upgrades and update - Remote support for issue resolution of reported problem within one working day from the date the issue was reported. - Corrective maintenance as required, provision and shipment to site of hardware to be replaced, disposal of faulty hardware, cost of replacement work, personnel transport arrangements - One 1-day remote session of theoretical and practical training of up to 10 operators on software administration and utilization. - Three (3) remote refresher training sessions on software administration and utilization for unlimited number of operators/radiographers. Remote refresher training is currently available in English, French, and Russian. Other languages might be available upon request. - Two (2) remote preventive maintenance sessions for preventive maintenance sessions for unlimited number of operators based on 'Delft Light Acceptance Maintenance checklist' including the provision of a maintenance report. Remote preventive maintenance is currently available in English, French and Russian. Other languages might be available upon request. - Provision of hard and digital copies of training materials and access to Delft eLearning online platform for each trainee For any purchase of this item, a CAD4TB Box, a CAD4TB Software Perpetual License must be procured in equal quantities. | Delft Imaging Systems B.V. |  | 5100 | USD | 0 | N/A | N/A | N/A | N/A | N/A | N/A | CAD4TB-WTY-EXT-1 | N/A | N/A
200014 | CAD4TB Support and Maintenance Extension 3 years | 1 unit | N/A | Extension of support, maintenance and warranty covering CAD4TB software (CAD4TB Software Perpetual License) and hardware (CAD4TB box) for 3 years. - Free software patches, upgrades and update - Remote support for issue resolution of reported problem within one working day from the date the issue was reported - Corrective maintenance as required, provision and shipment to site of hardware to be replaced, disposal of faulty hardware, cost of replacement work, personnel transport arrangements - One 1-day remote session per year of theoretical and practical training of up to 10 operators on software administration and utilisation. - Three (3) remote refresher training sessions per Year on software administration and utilization for unlimited number of operators/radiographers. Remote refresher training is currently available in English, French, and Russian. Other languages might be available upon request. - Two (2) remote preventive maintenance sessions per Year for preventive maintenance sessions for unlimited number of operators based on 'Delft Light Acceptance Maintenance checklist' including the provision of a maintenance report. Remote preventive maintenance is currently available in English, French and Russian. Other languages might be available upon request. - Provision of hard and digital copies of training materials and access to Delft eLearning online platform for each trainee For any purchase of this item, a CAD4TB Box, a CAD4TB Software Perpetual License, a CAD4TB Support and Maintenance Extension 1 year must be procured in equal quantities. | Delft Imaging Systems B.V. |  | 11475 | USD | 0 | N/A | N/A | N/A | N/A | N/A | N/A | CAD4TB-WTY-EXT-3 | N/A | N/A
1.3 ACCESSORIES
1.3.1 Solar panel for Delft light and CAD4TB box
200020 | Mobisun Pro Solar Panel | 1 unit | N/A | Lightweight portable solar panel with integrated battery and compact power converter that delivers direct high voltage (230V) current and power (250W). Able to power the 'Delft Light Full kit together with the Delft CAD4TB box. Fully charged in approximately 16 hours of sunshine or via the mains in approximately 2,5 hours. Solar panel with battery capacity 70000 mAh, 230 V output, max. 250W continuous, max. 500W peak. 2 x 5V USB output, max.3,1A. LiFePO4 battery 256 Wh, capacity 70 Ah. Water resistant IPX4. The Mobisun Pro Solar Panel can also work as a UPS when the items are charged through the integrated power bank, it has a voltage stabilizer that can protect the batteries from voltage spikes. This item is included by default to the Delft Light Full kit. Can be procured additionally to the Delft CAD4TB box to ensure a full autonomy when connected to the FDR Xair System or any ultraportable X-ray system. | Delft Imaging Systems B.V. |  | 1750 | USD | 0 | Dangerous cargo | 8 | 60 | 32 | 14 | 0.025 | Delft Light Mobisun Pro Solar Panel DL-SLR-2 | 85044030 | https://www.mobisun.com/mobisun-pro-portable-solar-panel-battery-generator-230v/

## 6 Medical devices Planning tool
Medical devices Logistics & planning summary of your order
Supplier | General Cargo |  | Dangerous cargo |  | Cost of services | Cost of products EXW per supplier | Product discount | Total cost with product discount
| Weight (kg) | Volume (m3) | Weight (kg) | Volume (m3)
Delft Imaging Systems B.V.
Infervision Medical Technology Co. Ltd
Fujifilm Corporation
Subtotals weight (kg) & volume (m3) | [6 empty] | Total EXW cost | 0
[7 empty] | Total weight (kg) | 0
[7 empty] | Total volume (m3) | 0

## Calculations - Medical devices (hidden)
Item code | Supplier | Quantities | Transport type | Total Weight | Total volume | Total cost | Filtered title
1. TB SCREENING / TRIAGE | 0 | 0 | 0 | 0 | 0 | 0 | 1. TB SCREENING / TRIAGE
1.1 CHEST RADIOGRAPHY (CXR) | 0 | 0 | 0 | 0 | 0 | 0 | 1.1 CHEST RADIOGRAPHY (CXR)
1.1.1 Ultraportable Digital X-ray systems | 0 | 0 | 0 | 0 | 0 | 0 | 1.1.1 Ultraportable Digital X-ray systems
200015 | Fujifilm Corporation | 0 | Dangerous cargo (CAO): 1. All components (except the power bank): GC UN3481 PI966 Section II. 2. Power bank: Dangerous cargo UN 3480 - Cargo aircraft only (CAO). | 0 | 0 | 0 | 1.2 CAD TECHNOLOGIES
200016 | Fujifilm Corporation | 0 | N/A | #VALUE! | #VALUE! | 0 | 1.2.1 CAD software for automated interpretation of CXRs
200017 | Fujifilm Corporation | 0 | N/A | #VALUE! | #VALUE! | 0 | 1.3 ACCESSORIES
200018 | Fujifilm Corporation | 0 | N/A | #VALUE! | #VALUE! | 0 | 1.3.1 Solar panel for Delft light and CAD4TB box
200006 | Delft Imaging Systems B.V. | 0 | Dangerous cargo: 1. All components (except the Solar panel): GC UN3481 PI966 section I. 2. Only the solar panel included in the Delft Light full kit: Dangerous cargo UN3481 PI967 section I. | 0 | 0 | 0
200007 | Delft Imaging Systems B.V. | 0 | N/A | #VALUE! | #VALUE! | 0
200008 | Delft Imaging Systems B.V. | 0 | N/A | #VALUE! | #VALUE! | 0
200009 | Delft Imaging Systems B.V. | 0 | N/A | #VALUE! | #VALUE! | 0
1.2 CAD TECHNOLOGIES | 0 | 0 | 0 | 0 | 0 | 0
1.2.1 CAD software for automated interpretation of CXRs | 0 | 0 | 0 | 0 | 0 | 0
200001 | Infervision Medical Technology Co. Ltd | 0 | N/A | #VALUE! | #VALUE! | 0
200002 | Infervision Medical Technology Co. Ltd | 0 | Dangerous cargo UN 3481 Class 9 | 0 | 0 | 0
200003 | Infervision Medical Technology Co. Ltd | 0 | N/A | #VALUE! | #VALUE! | 0
200004 | Infervision Medical Technology Co. Ltd | 0 | N/A | #VALUE! | #VALUE! | 0
200005 | Infervision Medical Technology Co. Ltd | 0 | N/A | #VALUE! | #VALUE! | 0
200010 | Delft Imaging Systems B.V. | 0 | N/A | #VALUE! | #VALUE! | 0
200019 | Delft Imaging Systems B.V. | 0 | N/A | #VALUE! | #VALUE! | 0
200011 | Delft Imaging Systems B.V. | 0 | General cargo UN3481 (PI 967 section II). Lithium battery contained in equipment. | 0 | 0 | 0
200012 | Delft Imaging Systems B.V. | 0 | N/A | #VALUE! | #VALUE! | 0
200013 | Delft Imaging Systems B.V. | 0 | N/A | #VALUE! | #VALUE! | 0
200014 | Delft Imaging Systems B.V. | 0 | N/A | #VALUE! | #VALUE! | 0
1.3 ACCESSORIES | 0 | 0 | 0 | 0 | 0 | 0
1.3.1 Solar panel for Delft light and CAD4TB box | 0 | 0 | 0 | 0 | 0 | 0
200020 | Delft Imaging Systems B.V. | 0 | Dangerous cargo | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 | 0 | 0 | 0