Content

- Scope

- Definitions and abbreviations

- Personnel qualifications

3.1 Medical fitness

3.2 Education and training

- Procedure

4.1 Principle

4.2 Samples

4.3 Equipment and materials

4.4 Reagents and solutions

4.5 Detailed instructions for use and disinfection

4.6 Reading and recording

4.7 Quality control and maintenance

4.8 Waste management

- Related documents

Annex 1. Incubator logbook

Annex 2. Incubator maintenance logbook

|  | Compiled by |  | Examined by | Approved by | Replaced | New version |
| --- | --- | --- | --- | --- | --- | --- |
| Name |  |  |  |  | Code: | Code: |
| Date |  |  |  |  |  |  |
| Signature |  |  |  |  |  |  |
| Laboratory area: |  | No of copies: |  | Reason for change: |  |  |

- Scope

This SOP describes the use and maintenance of an incubator through proper installation, regular monitoring and maintenance.

- Definitions and abbreviations

°C:	degrees Celsius

UPS:	uninterrupted power supply.

- Personnel qualifications

- Medical fitness

In accordance with national laws and practices, arrangements should be made for  appropriate health surveillance of TB laboratory workers:

- before enrolment in the TB laboratory;

- at regular intervals thereafter, annually or bi-annually;

- after any biohazard incident;

- at the onset of TB symptoms.

Ideally, individual medical records shall be kept for up to 10 years following the end of occupational exposure.

Laboratory workers should be educated about the symptoms of TB and provided with ready access to free medical care if symptoms arise.

Confidential HIV counselling and testing should be offered to laboratory workers. Options for reassignment of HIV-positive or immuno-suppressed individuals away from the high-risk areas of the TB laboratory should be considered.

All cases of disease or death identified in accordance with national laws and/or practice as resulting from occupational exposure to biological agents shall be notified to the competent authority.

- Education and training

Education and training must be given on the following topics:

- potential risks to health (symptoms of TB disease and transmission);

- precautions to be taken to minimize aerosol formation and prevent exposure;

- hygiene requirements;

- wearing and use of protective equipment and clothing;

- handling of potentially infectious materials;

- prevention of incidents and steps to be taken by workers in the case of incidents (biohazard incidents, chemical, electrical and fire hazards);

- good laboratory practice;

- organization of work flow;

- waste handling;

- use of equipment (operation, identification of malfunctions, maintenance).

The training shall be:

- given before a staff member takes up his/her post;

- strictly supervised;

- adapted to take account of new or changed conditions; and

- repeated periodically, preferably every year.

- Procedure

4.1 Principle

The incubator is intended to ensure optimum growth conditions for the culture of Mycobacterium tuberculosis. Incubators are available in various sizes, from small (benchtop) models to large incubator rooms equipped with fans to achieve a homogeneous temperature.

Since M. tuberculosis growth is inhibited above 37 °C, the incubator temperature should be set at 36 ± 1 °C.

4.2 Samples

Mycobacterial specimens and cultures.

4.3 Equipment and materials

The incubator should be installed  and operated according to the instructions in the manufacturer’s manual (specific to each incubator). It should be located on a level surface in a dry, well-ventilated area, away from heat sources.

4.4 Reagents and solutions

70% ethanol.

4.5 Detailed instructions for use and disinfection

4.5.1  General

Keep door(s) closed to prevent heat loss to the environment.

Ensure that rack positions are clearly marked.

Switch off the incubator when it is not in use.

Do not overload.

4.5.2  Procedure for disinfection

Every 6 months or in case of spillage of infectious material within the incubator:

- Disconnect the incubator from the electrical supply.

- Identify an adequate volume of available safe space for temporary storage of biohazardous material. Ensure that materials kept in the incubator are clearly marked.

- Clean metallic surfaces (racks, floor, walls and doors) with 70% ethanol.

- Clean any materials located in the incubator that may have been contaminated by spillage.

- Reconnect the incubator to the electricity supply and switch it on.

- Replace incubator contents once the temperature has reached 36 °C.

- Notify staff of the replacement of the contents.

4.6  Reading and reporting

Use a record from to keep a daily record of the incubator temperature (Annex 1).

Record regular maintenance of the incubator on the maintenance logbook (Annex 2).

4.7 Quality control and maintenance

Perform monthly temperature spot checks by placing(maximum/minimum thermometers at various positions in the incubator to verify that the internal incubator temperature is constant at 36  1°C.

Repairs should be performed by a qualified service technician.

Calibration

The temperature should be calibrated:

- before use;

- after temperature changes have been detected and corrected;

- following a power failure;

- after cleaning of spillages.

Calibration process

Ensure that the door is closed and that the incubator is switched on.

Set the required temperature using the temperature control and leave the incubator to run for 1 hour.

Place a thermometer in the centre of the incubator with the probe away from the heating element.

Take the temperature reading after 30 minutes; if the temperature is not 36 ± 1 °C, adjust the control.

Repeat the process every 30 minutes until the required temperature is reached.

Continue to take readings until two consecutive readings (30 minutes apart) are  36 ± 1 °C.

Record readings in the incubator logbook. The incubator is ready for use only at the calibrated temperature.

4.8 Waste management

All inoculated tubes and vials, whether negative or contaminated, should be autoclaved as potentially infectious material.

- Related documents

Manufacturer’s manual, specific to each incubator.

Basics of quality assurance for intermediate and peripheral laboratories, 2nd ed. Cairo, WHO Regional Office for the Eastern Mediterranean, 2002.

Laboratory services in tuberculosis control. Part III: Culture. Geneva, World Health Organization, 1998 (WHO/TB/98.258).

Maintenance and repair of laboratory, diagnostic imaging and hospital equipment. Geneva, World Health Organization, 1994.

Maintenance manual for laboratory equipment, 2nd ed. Geneva, World Health Organization, 2008 (available at www.who.int/entity/diagnostics_laboratory/documents/guidance/guidance2/en/)

Annex 1. Incubator logbook

Temperature required:              36 ± 1 °C (acceptable variation)

| Trimester: |  |  |  | #### Year: |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ## Month | Temp. °C | Operator (initials) | ## Month | Temp. °C | Operator (initials) | ## Month | Temp. °C | Operator (initials) |
| 1 |  |  | 1 |  |  | 1 |  |  |
| 2 |  |  | 2 |  |  | 2 |  |  |
| 3 |  |  | 3 |  |  | 3 |  |  |
| 4 |  |  | 4 |  |  | 4 |  |  |
| 5 |  |  | 5 |  |  | 5 |  |  |
| 6 |  |  | 6 |  |  | 6 |  |  |
| 7 |  |  | 7 |  |  | 7 |  |  |
| 8 |  |  | 8 |  |  | 8 |  |  |
| 9 |  |  | 9 |  |  | 9 |  |  |
| 10 |  |  | 10 |  |  | 10 |  |  |
| 11 |  |  | 11 |  |  | 11 |  |  |
| 12 |  |  | 12 |  |  | 12 |  |  |
| 13 |  |  | 13 |  |  | 13 |  |  |
| 14 |  |  | 14 |  |  | 14 |  |  |
| 15 |  |  | 15 |  |  | 15 |  |  |
| 16 |  |  | 16 |  |  | 16 |  |  |
| 17 |  |  | 17 |  |  | 17 |  |  |
| 18 |  |  | 18 |  |  | 18 |  |  |
| 19 |  |  | 19 |  |  | 19 |  |  |
| 20 |  |  | 20 |  |  | 20 |  |  |
| 21 |  |  | 21 |  |  | 21 |  |  |
| 22 |  |  | 22 |  |  | 22 |  |  |
| 23 |  |  | 23 |  |  | 23 |  |  |
| 24 |  |  | 24 |  |  | 24 |  |  |
| 25 |  |  | 25 |  |  | 25 |  |  |
| 26 |  |  | 26 |  |  | 26 |  |  |
| 27 |  |  | 27 |  |  | 27 |  |  |
| 28 |  |  | 28 |  |  | 28 |  |  |
| 29 |  |  | 29 |  |  | 29 |  |  |
| 30 |  |  | 30 |  |  | 30 |  |  |
| 31 |  |  | 31 |  |  | 31 |  |  |

Annex 2. Incubator maintenance logbook

| ## ITEM IDENTIFICATION |  |  |
| --- | --- | --- |
| Equipment: INCUBATOR | Brand name: |  |
| Purchase date: | Model/type: |  |
| Location within laboratory: | Serial no. |  |
| Warranty expiry date: |  |  |
| Manufacturer: |  | Tel: |
| Address:. |  |  |
| Contact person: |  |  |
| Technical service representative: |  | Tel: |

| FAILURE EVENTS |  |  |  |
| --- | --- | --- | --- |
| Date | Event | Corrective action taken | Operator |

| Disinfection date |  |
| --- | --- |
| Disinfection reason (regular, spillage) |  |
| Operator's name |  |
| Calibration date |  |
| Calibration reason (temperature correction, power failure, disinfection) |  |
| Operator's name |  |

| Disinfection date |  |
| --- | --- |
| Disinfection reason (regular, spillage) |  |
| Operator's name |  |
| Calibration date |  |
| Calibration reason (temperature correction, power failure, disinfection) |  |
| Operator's name |  |

| Disinfection date |  |
| --- | --- |
| Disinfection reason (regular, spillage) |  |
| Operator's name |  |
| Calibration date |  |
| Calibration reason (temperature correction, power failure, disinfection) |  |
| Operator's name |  |