Content

- Scope

- Definitions and abbreviations

- Personnel qualifications

3.1 Medical fitness

3.2 Education and training

- Procedure

4.1 Principle

4.2 Samples

4.3 Equipment and materials

4.4 Reagents and solutions

4.5 Detailed instructions

4.6 Quality control

- Related documents

|  | Compiled by |  | Examined by | Approved by | Replaced | New version |
| --- | --- | --- | --- | --- | --- | --- |
| Name |  |  |  |  | Code: | Code: |
| Date |  |  |  |  |  |  |
| Signature |  |  |  |  |  |  |
| Laboratory area: |  | No of copies: |  | Reason for change: |  |  |

- Scope

This SOP describes the preparation of all reagents used during specimen processing for culture and identification of Mycobacterium tuberculosis.

- Definitions and abbreviations

CPC:	cetylpyridinium chloride

MW:	molecular weight

NA:	not applicable

NALC:	N-acetyl L-cysteine

- Personnel qualifications

- Medical fitness

NA

- Education and training

Basic education and training must be given on the following topics:

- use, calibration, identification of malfunctions and maintenance of autoclave, balance, pH meter and all equipment used in a media preparation laboratory;

- prevention of incidents and steps to be taken by workers in the case of incidents (biohazard incidents, chemical, electrical and fire hazards);

- good laboratory practice;

- knowledge of aseptic technique;

- organization of work flow;

- importance of media and reagent quality for laboratory results and patient management.

The training shall be:

- given before a staff member takes up his/her post;

- strictly supervised;

- adapted to take account of new or changed conditions; and

- repeated periodically, preferably every year.

- Procedure

- 4.1 Principle

Reagents should be prepared in small aliquots to avoid reuse of stock solutions.

- 4.2 Samples

NA

- 4.3 Equipment and materials

Balance (sensitivity 0.01 g), spoons or spatulas, weighing paper

Autoclave

pH meter

Dispenser device

Appropriate sterile glassware and sufficient sterile screw-cap tubes for dispensing the reagents.

Dark bottles

Refrigerator

General equipment of a media preparation laboratory; (see SOP on Preparation of plain egg based media)

- 4.4 Reagents and solutions

#### 4.4.1 Reagents for the sodium hydroxide (modified Petroff) method

- Sodium hydroxide (NaOH) solution, 4%

- Sodium hydroxide pellets (analytical grade)	4 g

- Distilled water	100 ml

- Dissolve NaOH in the distilled water. Aliquot in 2-ml quantities. Sterilize by autoclaving at 121 ºC for 20 minutes.

- Phosphate buffer, 0.067 mol/litre, pH 6.8

- Stock solution A: disodium phosphate, 0.067 mol/litre

Dissolve 9.47 g of anhydrous Na2HPO4 in 1 litre of distilled water.

- Stock solution B: monopotassium phosphate, 0.067 mol/litre

Dissolve 9.07 g of KH2PO4 in 1 litre of distilled water

- Mix 50 ml of solution A and 50 ml of solution B. Use a pH meter to confirm that the correct pH for the buffer is reached. Adjust as necessary, using 10% phosphoric acid or 10% sodium hydroxide.

- Aliquot in volumes required for adding to a single centrifugation tube (e.g. 50-ml amounts if 50-ml centrifuge tubes are used), discarding the extra volume. Sterilize by autoclaving at 121 ºC for 20 minutes

- Left-over volumes of buffer can then be pooled and sterilized again for further use.

#### 4.4.2 Reagents for the NALC–NaOH method

- Sodium hydroxide (NaOH), 4%

- Sodium hydroxide pellets (analytical grade)	4 g

- Distilled water	        100 ml

Dissolve NaOH in the distilled water. Sterilize by autoclaving at 121 ºC for 20 minutes.

- Trisodium citrate, 2.94%

- Trisodium citrate·2H2O, analytical grade	2.94  g

HOC(COONa)(CH2COONa)2·2H2O

(C6H5Na3O7.2H2O, MW 294.10)

or

- Trisodium citrate anhydrous	2.6 g

(C6H5Na3O7, MW 258.07)

- Distilled water 	100 ml

Dissolve trisodium citrate in the distilled water. Sterilize by autoclaving at 121 ºC for 20 minutes.

- NALC–NaOH solution freshly prepared

Mix equal volumes of solutions 1 and 2. Add 0.5 g N-acetyl L-cysteine (NALC) just before use. Aliquot in 4-ml amounts.

- Phosphate buffer, 0.067 mol/litre, pH 6.8

See section 4.4.1 above.

#### 4.4.3 Reagents for the trisodium phosphate method

10% trisodium phosphate

- Trisodium phosphate (Na3PO4·12H2O, MW 380.13)  	23 g

- Distilled water	100 ml

Aliquot in 4-ml amounts.

#### 4.4.4 Reagents for the 5% oxalic method

Oxalic acid, 5%

- Oxalic acid dehydrate	5 g

(HOOCCOOH·2H2O, MW 126.01)

- Distilled water	100 ml

Aliquot in 4-ml amounts. Sterilize by autoclaving at 121 ºC for 20 minutes.

#### 4.4.5 Reagents for the CPC/NaCl method

1% CPC–2% NaCl solution

- Cetylpyridinium chloride 	10 g

- Sodium chloride	 20 g

- Sterile distilled water. 	1000 ml

Distribute in 4-ml aliquots in sterile bottles. CPC solution is self-sterilizing and remains stable if protected from light, extreme heat and evaporation. The stock solution should be stored in dark bottles at room temperature.

#### 4.4.6 Reagents for decontamination using the modified Kudoh  method

Sodium hydroxide (NaOH), 4%

- Sodium hydroxide pellets (analytical grade)	 4 g

- Distilled water	100 ml

Dissolve NaOH in the distilled water. Aliquot in 4-ml amounts. Sterilize by autoclaving at 121 ºC for 20 minutes.

#### 4.4.7 Other reagents

- Saline solution .

- Sodium chloride (NaCl)	 0.9 g

- Distilled water	100 ml

Sterilize by autoclaving at 121 °C for 20 minutes.

- Kirchner medium

- Disodium o-phosphate·12H2O	19.0 g

- Monopotassium phosphate, anhydrous	2.5 g

- Magnesium sulfate.7H2O	0.6 g

- Trisodium citrate	2.5 g

- Asparagine	5.0 g

- Glycerol	20.0 ml

- Phenol red, 0.4 %	3.0 ml

- Distilled water 	up to 1000 ml

Steam to dissolve. The pH should be 7.4–7.6. Autoclave at 115 °C for 10 minutes. When cool, add 10% of horse serum. Incubate at 37 °C for 24 hours to test sterility.

- Blood agar plates (chocolate plate)

- Nutrient substrate (heart extract and peptones	20.0 g

- Sodium chloride 	5.0 g

- Agar 	15.0 g

Note: A nutrient agar base, with the composition given above, is commercially available. Once the bottle is opened, the base may be used up to its expiry date if stored dry, tightly closed and protected from light at 15–25 °C.

- Distilled water 	1000 ml

Autoclave at 121 °C for 15 minutes. Cool to 45–50 °C and add 5–8 % sheep blood. Check that pH is 6.8 ± 0.2 at 25 °C. After adding the blood, heat the culture medium for 10 minutes at 80 °C with gentle swirling until it turns brownish (chocolate colour). Poured blood plates can be stored for a maximum of 3 months at 4 °C.

- Skim milk 10%

Prepare skim milk powder at 10% in sterile distilled water.

Autoclave at 115 °C for 10 minutes.

- McFarland turbidity standard No.1

The McFarland No.1 standard is the most widely used turbidity standard for the visual comparison of bacterial suspensions. The density of the resulting barium sulfate precipitate is used as a proxy to approximate the colony count of bacterial suspensions. McFarland No.1 is approximately equivalent to 3 x 108 CFU/ml.

Sulfuric acid (H2SO4), 1%

H2SO4		 		1 ml

Distilled water			99 ml

Always add acid to water.

Barium chloride (BaCl2), 1%

BaCl2				1 g

Distilled water			100 ml

McFarland preparation

Barium chloride solution, 1%	0.1 ml

Sulfuric acid, 1%		9.9 ml

Label the suspension and indicate the last day of use (1 month after preparation).

- 4.6  Quality control

Use high-grade reagents, sterile distilled water and a balance that is regularly checked for sensitivity.

New reagent batches used for decontamination procedures can be checked by processing a control tube containing all reagents except specimen, centrifuging it and inoculating a culture medium tube/vial with some of the liquid left in the tube. Growth of organisms would lead to the reagent batch being discarded.

- Related documents

Collins C, Grange J, Yates M. Organization and practice in tuberculosis bacteriology. London, Butterworths, 1985.

Kent PT, Kubica GP. Public health mycobacteriology: a guide for the level III laboratory. Atlanta, GA, U.S. Department of Health and Human Services, Centers for Disease Control, 1985.

Kudoh S, Kudoh T. A simple technique for culturing tubercle bacilli. Bulletin of the World Health Organization, 1974, 51:71–82

Laboratory services in tuberculosis control. Part III: Culture. Geneva, World Health Organization, 1998 (WHO/TB/98/258).