## Slide 1: MONITORING QUALITY INDICATORS
MODULE 5 (M5)

## Slide 2: MODULE CONTENTS
Quality (performance) indicator monitoring
Xpert MTB/RIF (Ultra)* quality indicators
Collecting and analysing quality indicator data
Connectivity quality indicators
* Refers to Xpert MTB/RIF and Xpert MTB/RIF Ultra

## Slide 3: LEARNING OBJECTIVES
At the end of this module, you will be able to:
  Understand the importance of quality indicator monitoring
  List quality indicators that are needed to monitor the quality of TB laboratory tests
  List the quality indicators that should be collected for Xpert MTB/RIF (Ultra)*
  Describe the process of analysing and reporting quality indicators

## Slide 4: QUALITY INDICATOR MONITORING
Routine monitoring of quality indicators allows the laboratory to detect issues early that may impact on the quality of results
The laboratory should collect and analyse quality indicators on at least a monthly basis, using a standardised format
Targets should be set for all indicators monitored, and any unexplained change in quality indicators, such as increase in error rates, a change in MTB positivity rate or rifampicin resistance rate, or a significant change in volume of tests conducted, should be documented and investigated
A standard set of quality indicators should be used for all sites conducting a particular test to allow for comparison
Laboratories must work with clinicians and programme managers to develop and monitor quality indicators that reflect the whole diagnostic process

## Slide 5: QUALITY INDICATOR MONITORING
Quality indicators should be reviewed by the laboratory manager and must always be linked to root cause analysis and corrective actions if any unexpected results or trends are observed
Documentation of corrective actions and subsequent improvement and normalization of laboratory indicators following the corrective actions are critical
A system should be in place for centralized reporting of monthly quality indicators to the NRL or NTP
See Module PM6- How to Monitor and Evaluate your Programme

## Slide 6: QUALITY INDICATOR MONITORING
General indicators apply to all technologies and should be collected, analysed on a monthly basis, and disaggregated according to tests*
Adapt according to NTP guidelines in your country
Indicator | Target
Number of tests performed, by type of test | -
Service interruptions | No interruptions
Stock outs | No stock outs leading to service interruption
Equipment down time | No equipment downtime leading to service interruption
Turnaround time (TAT) | 90% of results meet test-specific TAT
Test statistics (quality indicator) report | 100% reports completed by defined due date
EQA results | >90% EQA panels are passed
QC results | >90% QC results meet expected criteria
Specimen rejection | <1% specimens rejected
Customer satisfaction | >80% surveyed customers are satisfied
* http://www.stoptb.org/wg/gli/assets/documents/GLI_practical_guide.pdf

## Slide 7: GENEXPERT QUALITY INDICATOR MONITORING
The recommended quality indicators for Xpert MTB/RIF (Ultra) are listed below
  These indicators should be collected in addition to the general quality indicators listed above
  Targets provided in the tables are intended as a guide, and laboratories should determine their own targets
  Targets vary based on local situation, patient population tested, and other relevant factors
  Deviations from the usual rates should be investigated

## Slide 8: GENEXPERT QUALITY INDICATOR MONITORING
Indicator | Description | Target | Reference
Number and proportion of specimens with MTB detected, rifampicin resistance detected | Number of specimens with MTB detected rifampicin resistance detected / Total number of specimens tested | Dependent on population tested and country drug- resistance prevalence and patterns | Xpert MTB/RIF implementation manual. World Health Organization. http://www.who.int/ tb/laboratory/xpert_ launchupdate/en/ GLI Practical Guide to TB Laboratory Strengthening. http://www.stoptb.org/wg/gli/assets/documents/GLI_practical_guide.pdf
Number and proportion of specimens with MTB detected rifampicin indeterminate | Number of specimens with MTB detected rifampicin indeterminate / Total number of specimens tested | Dependent on population tested and country drug- resistance prevalence and patterns | 
Number and proportion of specimens with MTB not detected | Number of specimens with MTB not detected / Total number of specimens tested | Dependent on population tested and country drug resistance prevalence and patterns | 

## Slide 9
Indicator | Description | Target | Reference
Number and proportion of specimens with errors | Number of specimens with errors / Total number of specimens tested | <3% | Xpert MTB/RIF implementation manual. World Health Organization. http://www.who.int/ tb/laboratory/xpert_ launchupdate/en/ GLI Practical Guide to TB Laboratory Strengthening. http://www.stoptb.org/wg/gli/assets/documents/GLI_practical_guide.pdf
Number and proportion of specimens with invalid results | Number of specimens with invalid results / Total number of specimens tested | <1% | 
Number and proportion of specimens with no results | Number of specimens with no results / Total number of specimens tested | <1% | 
Laboratory turnaround time | Time between receipt of specimen for Xpert MTB/RIF at the laboratory and result reporting | 2–24hrs | 
GENEXPERT QUALITY INDICATOR MONITORING

## Slide 10
Indicator | Description | Reference
Number and proportion of trace calls, disaggregated by patient group | Number of specimens with trace call results / Total number of specimens tested | GLI planning for country transition to Xpert MTB/RIF Ultra cartridges. http://www.stoptb.org/wg/gli/assets/documents/GLI_ultra.pdf
Number and proportion of patients whose first sample produces a trace result and who have a repeat test conducted, disaggregated by patient group | Number of first specimens with trace call results with a repeat test / Total number of repeat specimens tested | 
Number and proportion of patients who have a repeat test conducted whose second sample gives a result for MTB detection and rifampicin resistance, disaggregated by patient group | Number of first specimens with trace call results with a repeat test that gives a MTB detection and rifampicin resistance / Total number of repeat specimens tested | 
GENEXPERT ULTRA QUALITY INDICATOR MONITORING

## Slide 11: DIAGNOSTIC CONNECTIVITY INDICATOR MONITORING
Connected diagnostic test devices such as the GeneXpert, can transmit data to a variety of users and can provide a highly cost-effective way to ensure proper functioning of the diagnostic device network
Diagnostics connectivity solutions typically comprise:
A connectable diagnostic device that produces electronic
A software platform that receives and interprets data
A means to transmit data from the device to the software platform and to a server

## Slide 12: DIAGNOSTIC CONNECTIVITY INDICATOR MONITORING
With remote monitoring, designated persons can use any internet-enabled computer to access the software platform, providing them with an overview of the facilities, devices and commodities in their network
For an updated list of connectivity solutions: http://tinyurl.com/gliconnectivity
For more information see the GLI Quick Guide to TB Diagnostics Connectivity Solutions. Global Laboratory Initiative. 2016. http://stoptb.org/wg/gli/assets/documents/gli_connectivity_guide.pdf

## Slide 13: MANUAL DATA ANALYSIS
If no connected diagnostic solution is available, Xpert MTB/RIF (Ultra) result data can be collected and analysed manually using an Excel-based tool
Testing site data should be sent to the NRL or NTP for aggregation and analysis for reporting to the programme
General indicators (e.g. TAT) can’t be monitored by a connected diagnostic solution. Manual data collection and analysis of these indicators are required

## Slide 14: COLLECTING AND ANALYSING PERFORMANCE INDICATOR DATA
Collect quality indicator data monthly
Review the quality indicator data monthly for trends
Track the quality indicator data over time (quarterly and annually) for trends that can’t be identified on a monthly basis
Compare the quality indicator data to pre-determined targets
Perform root cause analysis, take corrective actions and document the outcomes if problems are identified (e.g. unexpected increased or decreases)

## Slide 15: EXAMPLE: INCREASED ERROR RATE
Target error rate < 3%
Recorded error rate 10%
Action:
Review error record log and compare to the GeneXpert results
  Determine which module errors occur on
  Determine which batch of reagents
  Determine staff performing testing
Review for any patterns and determine root cause
Determine corrective action and how to determine success of action
Implement action
Monitor action
Document and record all steps

## Slide 16: EXAMPLE: INCREASED INVALID RATE
Target invalid rate < 1%
Recorded error rate 4%
Action:
Review invalid record log and compare to the GeneXpert results
  Determine which batch of reagents, is this a new batch?
  Determine staff performing testing
Review for any patterns and determine root cause
Determine corrective action and how to determine success of action
Implement action
Monitor action
Document and record all steps

## Slide 17: EXAMPLE: REVIEW OF RECORDING AND REPORTING
Determine how many tests were reported – review a randomly selected 10% of results recorded
Check log with GeneXpert results. If errors are found:
  Review all results
  Ensure requesting client is notified
  Determine why results were incorrectly recorded
  Determine corrective action and how to determine success of action
  Implement action
  Monitor action
  Document and record all steps

## Slide 18: ANALYSING QUALITY INDICATORS OVER TIME
On a graph plot various quality indicators and look for trends or outliers e.g.
  Change in testing volume
  Increase in errors
  Increasing shipment times
  Sudden increases in errors/invalids or no results
  One off increase in reporting errors
Conduct further investigation to determine root cause
Undertake corrective actions
Monitor effectiveness of actions

## Slide 19: EXAMPLE: CHANGE IN TESTING VOLUME
Chart: Xpert MTB/RIF tests performed
| Jan | Feb | Mar | Apr
Series 1 | 200 | 215 | 307 | 211
At the testing site, 200 and 215 Xpert MTB/RIF tests were performed in January and February respectively. In March the number increased to 307, and decreased to 211 in April
What could be the reason?
Increases in the number of Xpert MTB/RIF test performed could suggest that the the NTP conducted a special screening campaign

## Slide 20: EXAMPLE: CHANGE IN NUMBER OF RIFAMPICIN RESISTANT RESULTS
Chart: Xpert MTB/RIF tests RR
| Jan | Feb | Mar | Apr
 | 7 | 1 | 8 | 6
At the testing site, seven rifampicin resistant Xpert MTB/RIF test were detected in January, one in February, and eight and six in March and April respectively
What could be the reason for the decrease in February?
Changes in the number of Xpert MTB/RIF RR results could suggest errors in sample collection, instrument malfunction or errors in processing. The number of RR cases is often small , variation in monthly reported cases is frequent but less evident when reviewed on quarterly basis

## Slide 21: EXERCISE: REVIEWING QUALITY INDICATORS
Purpose
Aim: The objective of this exercise is to review quality indicator graph and to identify the problem(s), and list possible root causes of the problem(s)
Total Time
15 minutes: 10 minutes review, 5 minutes feedback
Process
Review data provided in the worksheet and answer the questions in groups of 4
Select one group to provide feedback and discuss their findings

## Slide 22: DISCUSSION QUESTIONS
  Why is collecting quality indicators important?
  What general quality indicators should be collected to evaluate the functioning of the TB laboratory?
  What quality indicators should be collected from the GeneXpert instrument to evaluate proper use?
  What quality indicators should be collected from Xpert MTB/RIF Ultra testing to evaluate proper use?

## Slide 23: KEY MESSAGES
Routine monitoring of quality indicators is a critical element of QA
Every month, each testing site should collect and analyse quality indicators to ensure quality of testing
Any unexplained change in quality indicators should be documented and investigated (root cause analysis)
Corrective actions should be put in place to to correct the problem, and their affect must be measured

## Slide 24
THANK YOU
© John Rae