## Slide 1
MODULE 6 (M6)
QUALITY ASSURANCE OF THE XPERT MTB/RIF (ULTRA)

## Slide 2: MODULE CONTENTS
Essential elements of a comprehensive quality assurance (QA) programme at Xpert MTB/RIF (Ultra)* testing sites
How to implement the following QA activities:
  Making the testing site safe and functional
  Conducting training
  Standardizing policies and documentation
  Maintaining and servicing equipment
  Ensuring adequate supplies and reagents
  Testing quality samples
  Monitoring performance indicators
  Reporting accurate results
  Implementing PT
* Refers to Xpert MTB/RIF and Xpert MTB/RIF Ultra

## Slide 3: LEARNING OBJECTIVES
At the end of this module, you will be able to:
Understand essential elements of a QA programme at testing site level
Describe the requirements for each QA component at the testing site level
Understand how to meet the requirements for each QA component at your testing site

## Slide 4: WHAT ARE QUALITY DIAGNOSTICS?
Quality diagnostics are accurate, reliable and reported timely in order to be useful in a clinical or public health setting
A QA programme ensures QA activities are implemented to ensure quality diagnostics
Implementing quality activities for diagnostic testing guarantees that the right test results are provided for the right patient, at the right time

## Slide 5: PROGRAMMATIC QA MANAGEMENT
  A QA programme is developed and implemented at central level by the Ministry of Health (MOH), National TB Programme (NTP) or National TB Reference Laboratory (NTRL)
  A QA programme provides general guidance and tools for standardized QA activities, as well as supervision of the QA activities and monitoring of the adherence to the procedures at the testing site level

## Slide 6: ESSENTIAL QA CENTRAL LEVEL ACTIVITIES
See Programme Module 4 (PM4): How to Plan and Implement a Quality Assurance Programme
Coordinate onsite supervisory visits
Standardize policies & documentation
Maintain & service equipment
Conduct
training
Implement proficiency testing
Strengthen the supply chain
Accurate, reliable and timely results
Monitor quality indicators

## Slide 7: ESSENTIAL QA TESTING SITE LEVEL ACTIVITIES

## Slide 8: MAKE THE TESTING SITE SAFE AND FUNCTIONAL
The Xpert MTB/RIF (Ultra) tests are low-risk procedures, with similar risk to performing direct AFM smear microscopy
  Perform testing in an adequately ventilated area (Module 1 (M1): TB Biosafety)
  Provide suitable Personal Protective Equipment (PPE) and train staff to use and how to use the PPE
  Ensure waste is segregated and disposed of according to local national guidelines or regulations
  Treat used cartridges and other used supplies as biohazardous
  Avoid incineration if possible

## Slide 9: MAKE THE TESTING SITE SAFE AND FUNCTIONAL
Ensure sufficient space for the instrument (10-15 cm clearance); away from wet procedures, direct sunlight or an open window
Secure premises for the equipment to prevent theft of the GeneXpert unit and the computer/laptop
Stable electricity supply or sufficient measures to ensure uninterrupted supply (generator, solar panels, battery/UPS backup, etc.)
Appropriate measures to prevent ambient temperature above 30°C where equipment is installed (e.g. ventilation, air conditioning) and storage of cartridges

## Slide 10: CONDUCT TRAINING
Train GeneXpert users in the operation of the instrument, performing of the Xpert MTB/RIF and Ultra tests and the associated QA activities
Sensitise clinicians who are responsible for managing TB patients on sample collection, recording and reporting and interpretation of Xpert MTB/RIF and Ultra test results

## Slide 11: CONDUCT TRAINING
National training curricula including the Xpert implementation plan developed by NTP/NTRL should be used for all trainings conducted
Testing site managers must select the appropriate staff to attend trainings:
  Select at least 1-2 staff per testing site for basic user training
  Candidates must have basic computer literacy and knowledge of laboratory registers
  Candidates must be those who will perform the GeneXpert testing on a regular basis
Adapt according to NTP guidelines in your country

## Slide 12: CONDUCT TRAINING
User Training – usually 3-5 days and must include practical and theoretical components
A pre-assessment should be used to measure efficacy of training
Competency assessments should be performed after training and on an annual basis to determine whether users are able to perform the test correctly
  Users are competent if they passed the training and competency assessment

## Slide 13: CONDUCT TRAINING
  Competency assessments must be performed by competent staff (e.g. testing site manager or advanced user)
All training and competency assessment must be documented and data sent to appropriate focal point
Conduct refresher training as needed (e.g. after failure of PT, high error rates, or other issues)
To ensure a competent workforce, create regular training opportunities for new staff, and re-training opportunities for ) existing staff
Experienced user / supervisor must monitor performance of new users (i.e establish a programme for “on-the job” training)

## Slide 14: CONDUCT TRAINING
Training is required for laboratory personnel & clinicians when transitioning from Xpert MTB/RIF to Xpert MTB/RIF Ultra*
User or Clinician Training- half or full-day
Training should include changes in diagnostic algorithms Including testing of eligible groups, extrapulmonary TB etc.
* http://www.stoptb.org/wg/gli/assets/documents/GLI_ultra.pdf

## Slide 15: STANDARDIZE POLICIES AND DOCUMENTATION
Standard Operating Procedures (SOPs) are usually developed at the national level to ensure consistency and standardization
Review the testing site SOPs and documents and modify to ensure they are consistent with national requirements
Modify SOPs to fit local context if required
Ensure SOPs are up to date and available onsite

## Slide 16: MAINTAIN & SERVICE EQUIPMENT
The GeneXpert is a precision instrument that requires regular maintenance to ensure that it provides accurate and precise results
  Perform daily, weekly, monthly, and annual maintenance tasks according to manufacturer’s instructions (refer to Cepheid training materials)
  Ensure all maintenance tasks are recorded on appropriate logs

## Slide 17: MAINTAIN & SERVICE EQUIPMENT
Troubleshoot testing or instrument failures:
Document corrective action(s) and ensure they have been effective
Ensure warranty or service contracts are in place and terms and conditions are adhered to
Maintain dated service records
Document all maintenance & servicing

## Slide 18: MAINTAIN & SERVICE EQUIPMENT
GeneXpert verification:
  Verification ensures that the GeneXpert is “fit for purpose” i.e. it can be used for testing clinical specimens
  Verification is performed using a verification panel (Xpert Check)
  Perform at least one verification test per module after:
    Instrument installation
    Post-calibration or swapping of instrument modules

## Slide 19: MAINTAIN & SERVICE EQUIPMENT
GeneXpert verification:
  The Xpert Check kit is used to check the optical system, verify the thermal system and perform a series of system-level tests to ensure full system functionality
  All modules must be verified every year with Xpert Check cartridges
  Follow the instructions provided with the verification panels

## Slide 20: MAINTAIN & SERVICE EQUIPMENT
GeneXpert verification:
  If an invalid/error/no result is obtained on any module, repeat the test in that module using the extra sample provided
  Report the instrument verification results to the appropriate person responsible for overseeing testing in the country
  Contact Cepheid immediately in order to assist with any issues encountered during verification process

## Slide 21: ENSURE ADEQUATE SUPPLIES AND REAGENTS
Accurately forecasting Xpert MTB/RIF (Ultra) test supply needs reduces the risk of an interruption in service due to shortage of reagents
Plan the procurement of reagents based on actual test consumption data
Regularly monitor cartridge consumption and stock in hand at site
Rotate stock to ensure that oldest reagents are used first
See Module 5 (M5): Procurement & inventory management

## Slide 22: ENSURE ADEQUATE SUPPLIES AND REAGENTS
New lot testing
  Designed to ensure consistent quality of different lots, ensure deterioration has not occurred during shipment and storage
  Incoming batches of Xpert cartridges are tested using a sample of cartridges with known positive and negative specimen to ensure expected performance
  To reduce costs, new lot testing may be done at a central level before distribution of a new batch of cartridges to testing sites
  Testing sites should monitor the testing results (i.e. rate of unsuccessful results) of new lots to ensure that no issues have arisen during distribution that might affect the central-level lot testing result

## Slide 23: ENSURE ADEQUATE SUPPLIES AND REAGENTS
New lot testing
  Continuous monitoring of performance indicators at testing sites is important for the early detection of any problems with cartridge batches due to local storage conditions or other factors
  If new lot reagents do not produce the desired results, these reagents should not be used
  Contact the manufacturer for further troubleshooting

## Slide 24: TEST QUALITY SAMPLES
To ensure a quality Xpert MTB/RIF (Ultra) test results, a good quality specimen is required
Inform healthcare workers on the sample requirements for Xpert MTB/RIF (Ultra) testing
Ensure patients are instructed in good sputum collection technique and quality of specimens is checked before sending for testing
Ensure specimens are labeled correctly and request forms are completed

## Slide 25: TEST QUALITY SAMPLES
Reject specimens that are incorrectly labeled, leaking, in broken containers or of insufficient volume
Record the date specimens arrive in the laboratory
Evaluate and record the specimen quality
Specimens can be refrigerated at 2–8°C for a maximum of 10 days prior to testing
If necessary, specimens can be stored at a maximum of 35°C for up to 3 days, and then refrigerated at 2–8°C for a combined maximum duration of 10 days
See Module 2 (M2): Sample Collection and Transport

## Slide 26: MONITOR PERFORMANCE INDICATORS
  A standard set of performance indicators should be collected by all testing sites conducting Xpert MTB/RIF (Ultra) testing*
  Additional performance indicators should be collected for Xpert MTB/RIF Ultra testing**
  Performance indicators should be analyzed, compared to pre-determined goals
The review of the performance indicators must always be linked to solutions (corrective actions) if unexpected results (non-conformities) or trends are observed
* http://www.stoptb.org/wg/gli/assets/documents/GLI_practical_guide.pdf
** http://www.stoptb.org/wg/gli/assets/documents/GLI_ultra.pdf

## Slide 27: MONITOR PERFORMANCE INDICATORS
A system should be in place for centralized reporting of monthly performance indicators from the testing site
Remote monitoring (data connectivity) can make collection of QA data simpler and more efficient, as well as providing data for other purposes, including procurement and reporting of patient results
See Module 5 (M5): Monitoring quality indicators
See Programme Module 6 (PM 6): How to Monitor & Evaluate your Programme

## Slide 28: REPORT ACCURATE RESULTS
The Xpert MTB/RIF (Ultra) test reports whether the sample contains M. tuberculosis complex, and whether it is rifampicin resistant
Ensure the Xpert MTB/RIF (Ultra) test results are reportedin accordance with country recommendations (e.g. within 24 hours of sample receipt)
Report results in a standard format, according to national guidelines:
  Xpert MTB/RIF (Ultra) requires the reporting of an additional ‘trace’ category of results (see Module 4 TM4: Recording & Reporting Results)
Sensitize clinicians on interpretation of results

## Slide 29: REPORT ACCURATE RESULTS
Put measures in place to prevent the unauthorized access to data:
  Access to testing site areas should be restricted to authorized staff only
  Protect the data by:
  Restricting network and system access (e.g. memory disks)
  Enforcing user authentication (i.e. using passwords)
  Maintaining antiviral software

## Slide 30: IMPLEMENT PT
PT is a means to determine the quality of the results generated at the testing site
PT compares testing site PT results with a reference result (and with other laboratory results for the same panel of samples) to determine inter-testing site comparability
The PT programme is an important tool for communicating with and motivating staff. Used correctly, the PT programme can be used to identify & resolve problems in testing
A variety of PT panels using different sample matrices are available. These include artificial sputum, dried tube specimens, dried culture spots, lyophilized samples and liquid samples

## Slide 31: IMPLEMENT PT
Receive the PT panels (these are sent to the testing sites several times per year)
Perform PT panel testing as specified by the provider, and submit the completed PT reporting form to the provider
Analyse PT result report
Store the PT reports and analyse for trends over time
Troubleshoot PT failures and identify corrective actions
Evaluate whether corrective actions were effective
See Module 7 (M7): External Quality Assurance

## Slide 32: EXERCISE: QA SCENARIOS
Purpose
To review Xpert MTB/RIF QA results and suggest actions to address the situations presented
Total time
50 minutes: about 15 minutes discussion in small groups, 20 minutes to report back to the large group, 10 minutes group discussions
Process
Work in four groups
On the worksheet provided (1 scenario per group) review the scenario and address the questions
Select a spokesperson and share your findings with the group

## Slide 33: EXERCISE SCENARIO 1
You notice that only one staff member performs Xpert MTB/RIF PT. In fact, the PT was not performed in the last round of testing because that staff member was on leave. When you questioned the other staff members they explained they are uncomfortable performing the PT. During the same discussion, you discover that in the previous round tested, the laboratory results were unsatisfactory (< 80%).
How should you handle Xpert MTB/RIF PT at your facility?
What should you do if your laboratory receives unsatisfactory Xpert MTB/RIF PT results?

## Slide 34: EXERCISE SCENARIO 2
You recently updated the GeneXpert testing SOP to include Xpert MTB/RIF Ultra testing. Later in the month, you note that results for the ‘trace’ category are not being recorded. When you question the staff member, you discovered he/she only had access to the old version of the SOP that did not reflect the updated changes. You realize there are several copies of this SOP, but you cannot recall where they are all located.
What actions can you take in managing your documents and records?

## Slide 35: EXERCISE SCENARIO 3
You walk into the store room and see 25 Xpert MTB/RIF cartridges ready to expire next week. You know from the last order and physical inventory that this should not be the case. When you check the cartridges in use in the TB laboratory, you see that their expiration date is six months from now. After reviewing the reagent log, you realize that the lot number with the longer expiration date has been used for the past several months.
How will you handle the current situation?
What steps will you take to prevent this situation from reoccurring?
How will you monitor future inventory cycling of stock to ensure the corrective action is effective?

## Slide 36: EXERCISE SCENARIO 4
Upon monthly review of the maintenance records for the GeneXpert instrument, no maintenance was recorded as having been done. How will you address:
The staff member who is responsible for performing and documenting the activities?
The staff member who says they forgot or did not know it was expected?

## Slide 37: DISCUSSION QUESTIONS
Which QA activities are recommended for Xpert MTB/RIF (Ultra) at testing site level?
List the criteria for choosing a staff to attend GeneXpert user training?
List three activities to ensure quality reagents are used for Xpert MTB/RIF (Ultra) testing?

## Slide 38: KEY MESSAGES
Recommended QA activities at the testing site include:
  Make the testing site safe and functional
  Conduct training
  Standardize policies and documentation
  Maintain and service equipment
  Ensure adequate supplies and reagents
  Test quality samples
  Monitor performance indicators
  Report accurate results
  Implement PT

## Slide 39
THANK YOU
© John Rae