## Slide 1
MODULE 7 (M7)
EXTERNAL QUALITY ASSURANCE

## Slide 2: MODULE CONTENTS
Introduction to External Quality Assurance (EQA) and Proficiency Testing (PT) for Xpert MTB/RIF (Ultra)*
  PT panel distribution & the PT process
  Common problems experienced with PT
  On-site supervision
Conducting on-site supervision (Advanced users)
  Supervision cycle
  Process of conducting on-site supervision
  Reporting on-site supervision visits
* Refers to Xpert MTB/RIF and Xpert MTB/RIF Ultra

## Slide 3: WHAT ARE EXTERNAL QUALITY ASSURANCE (EQA) AND PROFICIENCY TESTING (PT)?
EQA entails:
  Panel testing or Proficiency testing (PT)
  On-site supervision
  Slide re-checking
  Blinded re-checking (recommended for TB smear microscopy EQA) is not appropriate for Xpert MTB/RIF (Ultra( testing since the entire specimen is usually required to perform the test and no specimen can be saved for later re-checking
PT entails:
Sending a panel of biological materials with known results to testing sites for testing
The PT panel results from the testing site are compared to a reference result (and with other laboratory results for the same panel of samples) to determine inter-testing site comparability

## Slide 4: CLSA DEFINITIONS
  EQA
“Inter-laboratory comparisons and other performance evaluations that may extend throughout all phases of the testing cycle, including interpretation of results; determination of individual and collective laboratory performance characteristics of examination procedures by means of inter-laboratory comparison; NOTE: the primary objectives of EQA are educational and may be supported by additional elements.” [CLSI GP27-A2]
  PT
“A program in which multiple specimens are periodically sent to members of a group of laboratories for analysis and/or identification, in which each laboratory’s results are compared with those of other laboratories in the group and/or with an assigned value, and reported to the participating laboratory and others.” [CLSI GP27-A2]

## Slide 5: IMPORTANCE OF EQA
It can be used to improve performance across the laboratory network
It is an important tool for communicating with and motivating staff
It’s designed to identify and resolve serious problems with testing
It can be used to access the competency of test operators, and is a ISO requirement

## Slide 6: INTRODUCTION TO PT
PT programme checks pre-analytical, analytical, and post-analytical processes occurring at the testing site
Panels of samples are sent to testing site several times per year
PT samples are analyzed and the reported results are compared to reference values
Results indicate quality of personnel performance and test site operations
Results should be reviewed after each round and be monitored for trends over time
PT does not measure routine laboratory performance but may identify laboratories with major deficiencies

## Slide 7: PT ROLES & RESPONSIBILITIES
PT provider * | PT programme | Testing site | Advanced users
Manufacture PT panels | Collate PT results from all testing sites in the country | Receive panel | Conduct on-site supervision visits
Distribute PT panels to testing sites | Investigate PT problems from all testing sites to determine trends | Test panel | Troubleshoot PT problems
Collate reported PT panel results for all testing sites | Arrange supervisions to testing sites with PT problems | Report panel result | Determine effectiveness of corrective actions
Report PT outcomes to PT programme and testing sites | Re-train testing sites that repeatedly fail PT | Investigate & implement corrective actions for PT problems | 
 |  | Determine effectiveness of corrective actions | 
 |  | Monitor PT performance | 
Adapt according to NTP guidelines in your country
* NTRL, SRL or commercial PT provider

## Slide 8: PT CYCLE AT THE TESTING SITE
PT panels received
PT panels tested
PT panel results reported
PT panels sent
PT performance monitored
PT failures investigated
Corrective actions implemented
PT performance re-evaluated
PT panel results analysed

## Slide 9: RECEIVING PT PANELS
The testing site must:
  Check all contents when PT samples are received
  Document the date PT samples were received
  Confirm receipt and report any problems to the coordinator for the PT programme (e.g. NTRL, Quality Assurance Unit)

## Slide 10: TESTING & REPORTING PT RESULTS
As much as possible, samples should be treated in the same manner as patient samples:
Select who will process the PT panel
Carefully read the instructions received with the PT panel & fully understand what is required
Process and complete testing as directed by the panel provider
Record results
Complete PT reporting form
Return the PT results to the specified address using the specified format before the closure date

## Slide 11: RETURNING PT RESULTS
The testing site should:
  Review the results and undertake troubleshooting and corrective actions for any PT failures
  Troubleshooting should include a review of the pre-, analytical & post-analytical process to identify problems
  Troubleshooting may require the input from an advanced user to identify problems
  Determine whether the corrective actions have been effective

## Slide 12: EXAMPLE OF A XPERT MTB/RIF PT REPORT
Testing site information
Site results
Final score
Report on the performance of all the sites that took part in the testing

## Slide 13: COMMON PT PROBLEMS
Problem | Cause
PT results were not submitted to PT provider | Testing site did not receive the PT panel PT panels received damaged PT panel not tested Testing site sent PT report to wrong address PT report not returned
PT result report not available or reviewed | PT result report misplaced
PT results did not match expected results | PT result report not completed correctly or transcription error PT panel sample mix-up PT panel sample over-diluted PT panel sample contaminated

## Slide 14: MONITORING PT
The testing site should record the outcome of each proficiency testing round and look for trends over the year
Any negative trends should be reviewed and corrective actions implemented
Positive trends should be used to motivate the laboratory staff

## Slide 15: PT RESULT COMPILATION AND REVIEW
Testing site proficiency testing data should be recorded by the PT programme (NTP/NTRL) and reviewed after each panel and annually for patterns
All sites achieving less than 100% for the PT panel (even if scored a “Pass”) should investigate cause of errors and implement corrective actions
All testing sites that fail a PT panel should receive a onsite supervisory visit
Testing sites failing more than one panel should be retrained
This slide is not needed when training basic GeneXpert users

## Slide 16: EXERCISE: PT REPORT ANALYSIS
Purpose
The objective of this exercise is to review a PT report and to identify problems
Total Time
25 minutes: 10 minutes to complete the questions, 15 minutes for feedback
Instructions
In groups of four, review the provided PT report
Identify the PT the problem(s), and list possible causes of the problem(s)
List possible corrective actions to address the problem(s)
How could the these corrective actions be monitored?
Select someone to report your findings & suggestions

## Slide 17: ON-SITE SUPERVISION VISITS
On-site supervision visits should be conducted at regular intervals to assess laboratory/testing site practices
On-site supervision visits are usually conducted by NTP/NTRL and/or partners
The site will be contacted before the visit to ensure access and availability of staff
On-site supervision visits are conducted using standardized checklist for consistency and completeness of information

## Slide 18: DURING AND AFTER THE ON-SITE SUPERVISION VISIT
During
  Ensure assessor access to all testing areas and answer all their questions
  Receive feedback from the assessor
  After
  Review the assessors report
  Implement recommendations
  Address problems

## Slide 19: ADDRESSING PROBLEMS
Problem
Perform root cause analysis
Implement corrective action
Monitor effective-ness of corrective action
Record and document the process
Problems (or non-conformities) identified during onsite supervision visits must be investigated

## Slide 20: DISCUSSION QUESTIONS
What are the three components of EQA?
Why is EQA important?
What are the responsibilities of the testing site if they receive an incorrect PT result?

## Slide 21: KEY MESSAGES
EQA consists of PT, on-site supervision and slide re-checking (slide rechecking is not applicable to Xpert MTB/RIF & Xpert Ultra)
Multiple PT samples are periodically sent to testing sites, analyzed and the reported results compared to the reference values
PT is used to improve performance across the laboratory network, and is an important tool for communicating with and motivating staff
The testing site should review the PT results and undertake troubleshooting and corrective actions for any incorrect results

## Slide 22
THANK YOU
© John Rae

## Slide 23: CONDUCTING ONSITE SUPERVISION
(For Advanced GeneXpert Users)

## Slide 24: ONSITE SUPERVISION
On-site supervision is part of EQA
Site visits should be planned at regular intervals to assess laboratory/testing site practices
On-site supervision is usually conducted by NTP/NTRL and/or partners
  May be conducted by national level or regional/district level staff
  May be conducted by advanced users
  Should be integrated with other on-site supervision where possible (e.g. quarterly NTP site visits)
Adapt according to NTP guidelines in your country

## Slide 25: ONSITE SUPERVISION CYCLE
Supervisor schedules onsite supervision
Supervisor conducts onsite supervision
Supervisor provides verbal feedback
Supervisor prepares report
Supervisor sends report to testing site and central
SUPERVISOR
TESTING SITE

## Slide 26: SUPERVISION ROLES & RESPONSIBILITIES
Supervisor | Testing site | Central level (NTP/NTRL)
Schedules the onsite supervision with the testing site | Schedule the onsite supervision visit and ensure availability and access of the staff and facilities | Develop and circulate standardized checklists
Conducts the onsite supervision using a standardized checklist | After the visit, perform root cause analysis on the problem(s) identified | Identify which sites require onsite supervision visits and inform the supervisor(s)
Provides supportive supervision and mentoring as required | Develop action plans, with clear timeframes to address the cause of the problem(s) | Collate reports from onsite supervision visits and analyse data for trends
Provides verbal feedback to the testing site | Implement the corrective actions | 
Prepares the onsite supervision report and sends to the testing site and central level | Monitor the effectiveness of the corrective actions | 
 | Record & document what was done | 
 | Store the report from the supervisor for review at the next onsite supervision visit | 

## Slide 27: ONSITE SUPERVISION
On-site supervision should be conducted using standardized checklist for consistency and completeness of information
On-site supervision is used to determine how well a site is performing and provide the opportunity for in person mentoring and immediate remediation of urgent problems
The frequency of on-site supervision should be determined by the NTP/NTRL
Well performing testing sites can be visited less frequently but should still receive at least annual visit
Adapt according to NTP guidelines in your country

## Slide 28: ON-SITE SUPERVISION VISIT PROCESS
The supervisor conducting the on-site supervision will schedule the site visit in advance
The supervisor may need to discuss problems with laboratory management and the testing staff. The supervisor should ensure they are available on the day of the onsite supervision visit
The supervision visit can take several hours depending on the nature of the visit and level of the testing site

## Slide 29: ON-SITE SUPERVISION VISIT PROCESS
Supervisors may require access to sample reception, GeneXpert testing area & GeneXpert instrument, store and documentation
Supervisors may visit clinical staff and TB treatment areas. This information will be provided in advance to ensure availability of the clinical staff and facilities
Supervisors may wish to observe processing to determine the reason(s) for PT problems

## Slide 30: PROCESS OF ON-SITE SUPERVISION VISITS
If any problems are identified during the onsite supervision, the supervisor might collect additional data for troubleshooting or conduct some onsite troubleshooting
The supervisor may provide supportive supervision and mentoring to the GeneXpert staff
The supervisor will provide verbal feedback to the laboratory management, staff involved in testing and hospital management on the day of the onsite supervision visit
Adapt according to NTP guidelines in your country

## Slide 31: PROCESS OF ON-SITE SUPERVISION VISITS
The supervisor provides a completed checklist and report within two weeks of completing the onsite supervision visit
The report will include a list of problems identified (non-conformities)
The site will perform root-cause analysis, implement corrective actions to address the non-conformities and put measures in place to monitor the effectiveness of the corrective action
The measures will be evaluated at the next onsite supervision visit
Adapt according to NTP guidelines in your country

## Slide 32: ORGANISING A ONSITE SUPERVISION VISIT
Contact the testing site management and ensure they are aware of the time and date of visit
  Ensure accessibility to testing and clinical facilities
Contact laboratory management and ensure GeneXpert staff will be available during your visit
Adapt according to NTP guidelines in your country

## Slide 33: ORGANISING A ONSITE SUPERVISION VISIT
Before you go have the following items ready:
  Correct checklist
  Letter from programme granting access to the testing site
  Arrange transport if necessary
  Have a rewritable CD (1 per site) to collect data
  Have the previous checklist and nonconformities with you
Adapt according to NTP guidelines in your country

## Slide 34: ARRIVING AT THE TESTING SITE
Introduce yourself and explain the purpose of your visit to the site management
Have your letter of introduction ready to show the management
Introduce yourself to the laboratory management
Introduce yourself to the testing site staff
Adapt according to NTP guidelines in your country

## Slide 35: DURING YOUR ONSITE SUPERVISION VISIT
Be polite
Provide opportunity for staff to ask questions
Provide mentoring and support
Collect performance indicator data
Document evidence for non-conformities, do not include in the report personal opinions that are not backed up by evidence
Adapt according to NTP guidelines in your country

## Slide 36: COMPLETING THE CHECKLIST
Complete every question on the checklist
Write detailed notes, remember you are completing this so someone knows all about the testing site without having been to the testing site!
Do not be afraid to mark “NO” if a condition or requirement is not fulfilled. Always explain why you say so
Do not leave spaces empty, write “NA” (not applicable) if no information is available
Adapt according to NTP guidelines in your country

## Slide 37: FINISHING VISIT
When complete:
  Give detailed feedback, positive first then negative with constructive ways to improve
  Complete the site visit form to be left in the laboratory
Meet with hospital management for overall feedback, highlight areas in which hospital management can assist
Adapt according to NTP guidelines in your country

## Slide 38: AFTER THE VISIT
Complete your report
Send report to the GeneXpert Focal Person or other supervisor for review
Review comments and finalize the report
Send finalized report to the testing site, NTP/NTRL and regional coordinators
Adapt according to NTP guidelines in your country

## Slide 39: ON-SITE VISIT COMPILATION AND REVIEW
Site visit report should be recorded at a central level (NTP/NTRL) and reviewed after each quarter and annually for trends
Data should be used to:
  Determine sites which require re-training
  Plan for equipment needs
  Plan distribution of national documents
Adapt according to NTP guidelines in your country

## Slide 40: EXERCISE: ON-SITE SUPERVISION CHECKLIST
Purpose
The objective of this exercise is to develop a checklist of items to be used during an on-site supervision visit
Total Time
25 minutes: 15 minutes to complete the questions, 10 minutes for feedback
Process
In groups of four, develop a checklist of items that should be considered by a supervisor when planning, conducting and reporting an onsite supervision visit
Select a someone to report your findings and suggestions

## Slide 41: DISCUSSION QUESTIONS
Why is it important to contact the testing site before conducting an on-site supervision visit?
What are the responsibilities of the testing site if problems (non-conformities) are identified during the on-site supervision visit?
What is the purpose of analysing data from on-site supervision visits at central level (NTP/NTRL)?

## Slide 42: KEY MESSAGES
  On-site supervision is used to determine how well a site is performing and provide the opportunity for in person mentoring and immediate remediation of urgent problems
  On-site supervision should be conducted using standardized checklist for consistency and completeness of information
  The testing site should review the feedback from an on-site supervision visit and undertake troubleshooting and corrective actions for any deviation in results or problems

## Slide 43
THANK YOU
© John Rae