- NAME OF EQUIPMENT: Binocular microscope with LED fluorescence capability

Description of function and use:

This microscope is required for the detection of acid fast bacilli (AFB), mainly in sputum smears, but also in other clinical materials. It can be used in microscopy laboratories for the detection of AFB, and if desired for other types of microscopic examinations, either in brightfield or fluorescence microscopy (FM). Because of their high cost and lack of robustness, the classical fluorescence microscopy systems using mercury vapor lamps or other non-LED lamps are not recommended. They should only be used in laboratories performing a wide variety of tests in FM at widely differing excitation wavelengths, where LEDs would be impractical. In case only few FM applications with different excitation wavelengths are needed, modular LED FM systems with multiple or easily exchangeable LED lamps are preferred . LED-based microscopy is recommended by WHO to be phased in as an alternative for conventional ZN microscopy in both high and low-volume laboratories if the following specifications are met. Purchasers have at this moment three main options for replacement of existing microscopes with LED-FM. - dual-use microscopes designed for simple switching between reflected FM and transmitted bright-field - standard microscopes plus attachments consisting of special objectives and an LED light source that can adapt the microscopes for reflected FM - standard microscopes plus an attachment consisting of an LED light source that is fitted onto the illumination system of the microscope to adapt it for transmitted FM.

| - Tender specifications |  | ##### Bidder’s Specification |
| --- | --- | --- |
| ### Quantity | ##### MANUFACTURER: |  |
|  | ##### Type / Model |  |
|  | ##### Country of Origin |  |
| MAIN SPECIFICATIONS: |  |  |
| - Body: - Sturdy, stable, base body with focus adjustment devices positioned for a prolonged comfortable use and easy, precise movement. - The body shall be epoxy powder coated durable metal, heat treated and resistant to standard reagents used for staining, organic solvents used for cleaning of lenses and disinfectants. |  |  |
| Optical system: All optical parts including objectives, eye pieces, lenses, prisms, should have anti-reflective and anti-fungal coating with a warranty of 1 year at minimum. |  |  |
| Binocular eyepiece Binocular eyepiece, preferably with a tube flexible for use in an upper and lower position to ease the use by different lab workers, an ergonomic viewing angle of 30°, a tube rotatable through 360° at interpupillary distance ranging from at least 48–75 mm, maintaining parafocality. |  |  |
| Eyepieces Paired, high-quality, achromatic, wide field, 10x magnification. The eyepieces should have a minimum field number of 18, an eyepiece diopter adjustment on at least one eyepiece and no pointer. |  |  |

| Objectives For fluorescence microscopy, at least two infinity-corrected Plan Achromat objectives, corrected for use without cover glass (no-cover-slip slides). Objectives ≥50x should be of spring-loaded type. - 10x NA: 0.25 optional - 20x NA: 0.40 essential - 40–63x dry: essential for confirmation - 100x NA: 1.25, for oil immersion (optional) Marking and identification All objectives should be engraved with the following information: - name or code of the manufacturer - magnification and numerical aperture (NA). |  |
| --- | --- |
| Nosepiece Revolving nose piece to accommodate at least 4 objectives, any ports not covered by an objective should be closed with dust proof metallic or hard plastic screw caps. The nose piece should be provided with ribbed grip for easy rotation on a precision ball bearing mechanism for smooth and accurate alignment with precise click stops. In changing from one objective to another or reintroducing the same objective by rotation, the object in the centre of the field should not appear displaced by more than 0.04 in the object plane in any direction. |  |
| - Stage - Rectangular built in, uniformly horizontal, mechanical stage. - The stage should be provided with a spring-loaded slide holder for safe and exact positioning of the slide. The construction should allow a smooth travel in traverse directions. - Travel range of at least 75 x 30 mm (w x d), fatigue-proof position of knobs for movement, right or left-hand operation. |  |
| Condenser: Sub stage condenser of Abbe type, 0.9/1.25, focusable with rack and pinion arrangement incorporating an iris diaphragm. |  |
| Substage illuminator Substage illuminator for brightfield microscopy, with possibility to switch easily between FM and brightfield without tools with either a: - built in light source (e.g. 20–30 W, 6 V halogen lamp), including a constant power supply with automatic voltage recognition and surge protector. The system should provide a light intensity adjustment device, and an easily accessible on/off switch. The lamp should be provided with a lamp socket for easy replacement of the bulb. The housing for the light source should be designed to prevent dispersion of light and mounted not to heat up the body of the microscope or - 3 W, 6 V white LED could be used, built-in or modular. |  |
| Fluorescence illumination Fluorescence illumination, depending on microscope type may be: - a built-in LED blue light source with maximum wavelength close to 450 nm for reflected light examination - an LED blue light source attachment with maximum wavelength close to 450 nm that is fitted into a special objective for reflected light examination - an LED blue light source attachment with maximum wavelength close to 450 nm that is fitted onto the illumination system of a standard microscope and sliding barrier 510 nm long-pass filter for transmitted light examination. |  |
| Focusing knobs Co-axial coarse and fine focusing knobs capable of smooth fine focusing movement over the full range of coarse travel. The fine focusing movement should have a sensitivity of at least 500 μm per rotation over the entire coarse focusing range. Focusing knobs should be at both sides. A focusing stop safety |  |
| All metallic parts of the microscope to be corrosion proof, acid proof and stain proof. |  |

| Electricity requirements |  |
| --- | --- |
| Supply voltage: 100 -230 ± 10V, AC, 50/60 Hz Voltage and plugs shall be adapted to those used inside the country. |  |
| Power Consumption: Will depend on the illumination equipment (max. 30 W) |  |
| Conform to electrical safety IEC-60601-1, UL 61010-1, EN 61010-1 |  |
| Power supply, wide range input with 6 V converter. |  |
| Protection class (in accordance with EN 60529) |  |
| Designed not to interfere with circuit radio (in accordance with EN 55014). |  |
| Documentation |  |
| Manufacturer’s Certificate The manufacturer must have a management system certified to ISO 9001 or equivalent and a type-test certificate of relevant optical and mechanical tests |  |
| Quality and Safety standards met by the product offered have to be listed |  |
| Accessories: |  |
| Dust cover made of antistatic material |  |
| - Battery pack (Optional) - Battery pack (6 V DC) with charger working with 100-230 V, AC, 50 – 60 Hz supply as well as from the cigarette lighter of a car, 12 V DC |  |
| Eye protection (Optional) - If using an FM without a darkroom, a pair of well-fitting soft rubber pieces to protect the eyes and block stray light. |  |
| One antistatic cleaning brush for each microscope |  |
| At least one blue filter per microscope, (wrapped separately in a box). |  |
| A self-standing mirror unit (Optional) adapted to the space between base and substage condenser, providing bright illumination when used in brightfield. |  |
| If a special plug is needed for the country, a set of adapters should be provided. |  |
| Operation, maintenance and installation |  |
| Operation and Maintenance Manual: At least one set of operation, maintenance and service manuals for each microscope, written in United Nations languages (or at least in English) and preferably also in the official national language of the country requesting the microscope. The manuals to include instructions for: - setting up the microscope - routine cleaning and maintenance (including how to change the bulb) - changing the batteries inside the pack - installing and using the mirror (for when the electric lamp is not working) - planning periodic maintenance. |  |
| Installation and maintenance The bidder must arrange for the equipment to be installed by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail. The bidder to also provide user training (including how to use and maintain the equipment) and a comprehensive maintenance plan. The cost of the maintenance plan to be defined and guaranteed over the period of warranty. The supplier to provide an after-sale service that covers the whole country. The service to have competent staff, adequate infrastructure and sufficient spare parts to be able to respond to any complaints, and to repair or replace the microscope within 14 days. |  |
| Standard Maintenance Tools: All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication material, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |  |
| Spare Parts: Each assembled microscope should be accompanied by an authorized list of accessories and spare parts |  |
| Packing data: Packing data are not necessarily part of the bidding process, but are needed for shipment and customs declaration |  |
| Net weight |  |
| Gross weight |  |
| Dimensions (W x H x D) in cm |  |
| Appliances must be transported upright (Y/N) |  |
| Custom Tariff Number |  |
| Warranty: At least one year. |  |
| Remarks: The equipment offered, including its power supply, to be designed and constructed to operate properly and continuously in the conditions of the purchaser’s country; the equipment may need to tolerate high humidity, ambient temperatures of 5–40 °C, fungi, and spikes in the electricity supply. Bidder may propose additional products to the requirements listed above. |  |
| Optional : |  |
| Case A sturdy, lockable case for carrying or storing the microscope, with a handle and receptacles to hold accessories, objectives and eyepieces. The case to be padded and equipped with fixing points and straps to hold the microscope in place and to eliminate shocks during transportation.. |  |