## Slide 1: Lateral flow lipoarabinomannan testING and quality assurance training materials

## Slide 2: TRAINING agenda
Registration (as needed) Introductions, Training Objectives, and Outline Review Training Slides: Overview of TB Disease and TB/HIV Co-Infection Introduction to LF-LAM Testing The Alere Determine TB-LAM Ag Testing Procedure Continuous Quality Improvement of LF-LAM Testing De-brief and Closure
Note: This training is only conducted by trained and competent trainers.
The trainers will stay for the whole training period to ensure consistency, facilitate group interactions, demonstrate the testing procedure, and assess trainee competence to perform the test.

## Slide 3: Trainer & Participant Introductions

## Slide 4: Training GOAL AND objectives
Goal
    Build capacity to effectively provide quality-assured LF-LAM testing services in a safe environment
Objectives
    Equip participants with the knowledge, skills and understanding required to:
      Perform LF-LAM testing accurately and reliably
      Comply with quality and biosafety requirements
      Interpret test results accurately for patient care

## Slide 5: Participants
Participants may include stakeholders involved in LF-LAM test implementation across the TB cascade of care, including:
Testers and testing sites
Healthcare workers and TB or HIV clinicians
Technical Advisory/ Working Group members
Above site supervisors
Policy and planning personnel
Participants should be available to attend the entire didactic training. Testers should attend the full training, including testing procedure demonstration & independent completion.

## Slide 6: Overview of TB Disease and TB/HIV Co-Infection

## Slide 7: TUBERCULOSIS (TB)
What is TB?
An infectious disease caused by the bacterium Mycobacterium tuberculosis (Mtb)
Most commonly affects the lungs (pulmonary)
The most common cause of death from an infectious disease worldwide
How is TB transmitted?
TB can only be spread by those with active TB disease (not just TB infection)
TB is spread through the air
Activities that can lead to transmission include speaking, singing, and coughing
TB can remain in the air for hours
Global tuberculosis report 2023 (who.int)

## Slide 8: HIV co-infection impact on TB/HIV patients
TB continues to be the leading cause of death among HIV-infected people
People living with HIV (PLHIV) are 18 times more likely to fall ill with TB and experience more TB symptoms
Early TB/HIV diagnosis and prompt initiation of TB treatment:
  Reduces mortality
  Improves treatment outcomes and patient quality of life
  Reduces transmission of TB
*Fact sheet – World TB Day 2022 (unaids.org)

## Slide 9: Knowledge Check – Questions (Select the most accurate choice)
1. TB is a virus.
  True
  False
2. TB is transmitted when someone who is infected:
  Coughs
  Speaks
  Singing
  All of the above
3. HIV infection increases a TB patient’s risk for which of the following?
Becoming sick with TB
  Experiencing symptoms of TB
  Dying from TB
  All of the above
Notes: ANSWER KEY:
1. (a) true
2. (d) all of the above

## Slide 10: AN Introduction to LF-LAM testing

## Slide 11: What is LF-LAM?
LF-LAM stands for “Lateral Flow LipoArabinoMannan”
LAM is a component of the Mtb cell
LF-LAM tests are used to detect LAM in urine collected from eligible HIV-positive individuals
The Lateral Flow design of the LF-LAM test is simple, small, portable, and instrument-free

## Slide 12: What is lipoarabinomannan (LAM)?
A component of the outer cell wall of TB bacteria
An antigen, meaning it causes an immune response when it enters the human body
Sheds off from Mtb cells in the body during active TB disease. Many people living with HIV/AIDS have disseminated TB spread throughout the body, including TB in their kidneys (renal TB)
When TB bacteria in the kidneys shed off LAM, it is excreted into urine, which is how the urine LF-LAM test can detect LAM, and associated TB disease
Modified from the Treatment Action Group’s Activist Guide to the LAM Test

## Slide 13: Principle of the determine tb lam ag TEST
Urine is added to the sample pad
Urine flows across a membrane where LAM is bound by capture antibodies with gold reporter tags
LAM-antibody complexes then migrate across, and are bound by, another set antibodies fixed to the membrane along the Test line signaling detection of LAM
Remaining (unbound) capture antibodies migrate to a second set of antibodies along the Control line signaling test validity
2020 Global Laboratory Initiative LF-LAM Practical Guide

## Slide 14: ADVANTAGES AND BENEFITS OF LF-LAM TESTING FOR TB DETECTION AMONG PLHIV
Pulmonary TB is more difficult to diagnose among People Living with HIV (PLHIV)
    Most diagnostic tests for TB rely on sputum samples
    Sputum is difficult for PLHIV to produce and often contains low or variable amounts of mycobacteria
PLHIV experience higher rates of Extra-Pulmonary TB (EPTB) than those without HIV
    Majority of EPTB patients do not have detectable mycobacteria in sputum
LF-LAM is test that can bacteriologically confirm TB among PLHIV unable to produce sputum
Simple test; requires minimal infrastructure and supplies
Point-of-care assay that can be performed by trained lab- & non-lab testers
Fastest diagnostic test time for same-day result reporting for patient impact & clinical action
Mortality and treatment initiation benefits for PLHIV irrespective of setting
Urine is easier to collect from PLHIV and is relatively less biohazardous compared to sputum .
2021 Nathavitharana RR, et al. Coch. Rev; *Adult PLHIV at hospitals in Malawi, South Africa, Tanzania, Zambia and Zimbabwe; ^Adult PLHIV at primary healthcare clinics in South Africa

## Slide 15: PRACTICAL LIMITATIONS OF LF-LAM TESTING
Variable sensitivity (39-67% for the Determine TB LAM Ag test*)
Only PLHIV are eligible for LF-LAM testing (not HIV-negative persons)
Cannot distinguish tuberculous and non-tuberculous mycobacteria
Cannot detect drug resistance
Not recommended for TB treatment monitoring
* Determine TB LAM Ag test package insert

## Slide 16: WHEN to USE LF-LAM: [COUNTRY] LF-LAM testing algorithm
[INSERT NATIONAL DIAGNOSTIC ALGORITHM DIAGRAM HERE]
Always follow the sequence of the tests in the national algorithm

## Slide 17: when not to USE LF-LAM: 2021 WHO Global Guidelines
WHO recommends against using LF-LAM to assist in the diagnosis of active TB in HIV-positive adults, adolescents and children in outpatient settings:
without assessing TB symptoms (strong recommendation, very low certainty in the evidence about test accuracy)
without TB symptoms and unknown CD4 cell count or CD4 cell count greater than or equal to 200 cells/µl (strong recommendation, very low certainty in the evidence about test accuracy)
without TB symptoms and with a CD4 cell count of 100–200 cells /µl (conditional recommendation, very low certainty in the evidence about test accuracy)
Lateral flow urine lipoarabinomannan assay (‎LF-LAM)‎ for the diagnosis of active tuberculosis in people living with HIV: policy update 2019 (who.int)

## Slide 18: Knowledge Check – Question (Use the questions below for trainee response and group discussion)
What is LAM?
What are the benefits of LF-LAM testing?
What are the limitations of LF-LAM testing?
According to the national algorithm, which patients are eligible for LF-LAM testing?
According to the national algorithm, what action should be taken after a positive LF-LAM test result? A negative result?
Who should not be tested with LF-LAM?
Notes: ANSWER KEY:
(b) false – TB is a bacterium
coughs and (b) sneezes

## Slide 19: Applying the LF-LAM testing algorithm: Case scenario review & practice

## Slide 20: Notes on applying the algorithm: Clinic-tester coordination for success
Before LF-LAM Testing
Clinical team is aware of, and compliant with, the national testing algorithm when requesting an LF-LAM test
Clinical teams are aware of LF-LAM test availability
Clinical team completes relevant patient information (HIV status, other) when requesting an LF-LAM test
Sputum (or other WHO recommended sample type) must be collected in parallel to perform molecular WHO-recommended rapid diagnostic (mWRD) tests for confirmation of Tb diagnosis and detection of drug resistance.
After LF-LAM Testing
Tester communicates inability to test, invalid or indeterminate/equivocal test results (if repeated and confirmed), and valid test results promptly to the referring clinical team
Tester communicates with the clinical team regarding need for any additional testing

## Slide 21: Case scenario 1
A person living with HIV (PLHIV) with cough visits a tertiary clinic. LF-LAM tests and trained testers are available at the clinic, but there is no facility to perform an mWRD. What would you do?
Modified from WHO Operational Handbook on Tuberculosis. Module 3: Diagnosis
Brief Group Discussion
Prompts:
What if you could collect sputum?
What if the LF-LAM test was positive?

## Slide 22: Case scenario 1
A person living with HIV (PLHIV) with cough visits a tertiary clinic. LF-LAM tests and trained testers are available at the clinic, but there is no facility to perform an mWRD. What would you do?
PLHIV with cough
Perform LF-LAM
Send the sputum sample for Xpert MTB/RIF Ultra
LF-LAM test is positive
Initiate TB treatment with first line drugs
The test was positive for rifampicin susceptible Mtb
Continue to treat the patient with first line drugs
Modified from WHO Operational Handbook on Tuberculosis. Module 3: Diagnosis

## Slide 23: Case scenario 2
A person living with HIV (PLHIV) with cough visits a clinic. LF-LAM tests and trained testers are available at the clinic. The clinic is also has the facility to perform an mWRD. What would you do?
Modified from WHO Operational Handbook on Tuberculosis. Module 3: Diagnosis
Brief Group Discussion
Prompts:
What if you could collect sputum?
What if the LF-LAM test was negative?

## Slide 24: Case scenario 2
A person living with HIV (PLHIV) with cough visits a clinic. LF-LAM tests and trained testers are available at the clinic. The clinic is also has the facility to perform an mWRD. What would you do?
PLHIV with cough
Perform LF-LAM
Concurrently sent the sputum sample for Xpert MTB/RIF Ultra
LF-LAM test is negative
Wait for mWRD results
The test was positive for rifampicin susceptible Mtb
Initiate TB treatment with first line drugs
Modified from WHO Operational Handbook on Tuberculosis. Module 3: Diagnosis

## Slide 25: Case scenario 3
A person who is HIV negative with cough visits a tertiary clinic. LF-LAM tests and trained testers are available at the clinic. The clinic also has the facility to perform an mWRD. What would you do?
Modified from WHO Operational Handbook on Tuberculosis. Module 3: Diagnosis
Brief Group Discussion
Prompts:
Which samples would you collect?
Which tests would you request or do?

## Slide 26: Case scenario 3
A person who is HIV negative with cough visits a tertiary clinic. LF-LAM test and trained tester are available at the clinic. The clinic is also has the facility to perform an mWRD. What would you do?
HIV negative person with cough
Will NOT perform LF-LAM test
Send the sputum sample for Xpert MTB/RIF Ultra
The test was positive for rifampicin susceptible Mtb
Initiate TB treatment with first line drugs
Modified from WHO Operational Handbook on Tuberculosis. Module 3: Diagnosis

## Slide 27: Case scenario 4
A person living with HIV (PLHIV) with cough visits a tertiary clinic. LF-LAM testing is available at the clinic, but the trained LF-LAM tester is on leave for the day. There is another health care worker who has observed the LF-LAM test being performed. The clinic also has the facility to perform an mWRD, but is backed up for two days. What would you do?
Brief Group Discussion
Prompts:
Which samples would you collect?
Who would do the testing?

## Slide 28: Case scenario 4
A person living with HIV (PLHIV) with cough visits a tertiary clinic. LF-LAM test is available at the clinic, but the trained LF-LAM tester is on leave for the day, but there is another health care worker who has observed the LF-LAM test being performed. The clinic also has the facility to perform an mWRD, but is backed up for two days. What would you do?
PLHIV with cough
The untrained tester WILL NOT perform LF-LAM test
Send the sputum sample for Xpert MTB/RIF Ultra
LF-LAM test is positive
Initiate TB treatment with first line drugs
The test was positive for rifampicin susceptible Mtb
Continue to treat the patient with first line drugs
Store the urine sample according to IFU and wait for the trained tester

## Slide 29: Abbott/ Alere Determine TB LAM Ag testing Procedure

## Slide 30: Requesting an LF-LAM Test
Complete the Test Request Form
Label specimen collection containers with patient name, ID, and [insert any setting- or site-specific required labels]
[Insert any additional LF-LAM test request procedure or guidance]
[Insert image of the national/ site test request form here]

## Slide 31: Getting Ready
Wear appropriate PPE (minimum gloves and gown)
Prepare testing station and materials needed (test kit, precision pipette- if relevant, timer, biohazard discard bags)
Verify the condition of test kits (expiry date and storage) and record lot number.
Remove the test kit and the reference card to be used from the foil pouch and reseal the pouch containing the remaining tests.
Collect urine specimen from patient.
Note: Use 1 strip per test and be sure to preserve the lot number on the remaining packet of strips.

## Slide 32: Importance of biosafety
To reduce health-related risks associated with handling biological hazards in testing environments
Every specimen and test
should be treated as
though it is infectious

## Slide 33: LF-LAM biosafety considerations
Low-complexity, rapid, lateral flow test that is completed with minimal sample handling and manipulation
Urine is a non-respiratory sample type with low risk of Mtb aerosolization, regardless, to maximize tester safety: urine collection, handling, and testing should be done in a well-ventilated space.
PPE: Gloves and lab coat (always follow the site specific or basic universal biosafety protocol)
Infrastructure: No biosafety cabinet required
Reagents: All testing reagents provided in the kit
Waste Disposal: Follow the site specific/ or basic universal biosafety protocol (urine specimen should be treated as infectious and discarded appropriately)

## Slide 34: Ensure Proper Storage of LF-LAM test kits
Keep in a clean, organized and locked storeroom
Store according to manufacturer’s instruction's
Place in well ventilated room
Store away from direct sunlight
Place items on shelves
Organize existing and new shipments by expiration dates
Recommended storage temperature for LF-LAM test kits : 2°C–30°C
Shelf-life of test kits under recommended storage conditions: 2 years from date of manufacture
First expiry,
first out

## Slide 35: Collecting urine for LF-LAM testing
Urine Collection Area
      Identify and designate an area with adequate privacy (a toilet or a covered shed) for the patient to collect the urine sample privately.
      Ensure that there is facility for handwashing (with soap and clean water) within or close to the designated urine collection area for the patient and/or the caregiver and the healthcare worker to use before and after collecting and handling the urine sample.
Procedure for Urine Collection
      Label the urine collection container with patient identifiers at the time of urine collection.
      Prior to sample collection, instruct the patient or caregiver to:
          Clean the urogenital area with a sterile cleansing wipe and dispose of the wipe according to biowaste disposal procedures.
          After cleansing, to collect a minimum of 1 ml of midstream urine in the labelled urine collection container.
      If the sample is to be tested immediately, proceed according to the standard operating procedure (SOP) for Determine TB LAM Ag testing.
      If the sample is not used immediately, store the sample as described in table below and proceed according to the SOP for Determine TB LAM Ag testing when the sample taken out from storage and used for testing.
Temperature | Room temperature (18 to 30 °C) | Refrigerated temperature 2 to 8 °C | Frozen temperature ≤ -20 °C
Storage times | ≤ 8 hours | 3 days | 3 years#
# Samples may only undergo freeze thaw three times prior to testing.

## Slide 36: Using the Abbott Determine TB lam Ag Test
Tests packaged in kits of 25 test strips
After removal of each test strip, the kit/ bag should be completely resealed to prevent moisture build-up
Each test is separated for use via a perforated edge, labeled and its aluminum protective cover is removed from the surface
60ul of urine is pipetted onto the sample pad with timer set for 25 minutes
Resultant ‘Patient’ and ‘Control’ bands are read against a Reference Scale Card
Practical Considerations for LF-LAM TB Diagnostic Testing, Patricia Hall (TCMT/ILB/DGHT/CGH/CDC)
2020 Global Laboratory Initiative LF-LAM Practical Guide

## Slide 37: Result reading, interpretation and reporting
To assist in reading and interpretation, read test strips in normal indoor lighting. Note: do not read strips in direct sunlight
Check if for a visible band in the control and patient window of the test strip and determine results as indicated below:
Using the Reference Scale Card select and result log or register, read and document the test result and the line intensity (1,2,3,4) that is closest to the test line (refer to figure below)
Adapted from GLI Practical implementation of lateral flow urine lipoarabinomannan assay for detection of active tuberculosis in people living with HIV

## Slide 38: Troubleshooting Invalid Results
Problem
No control line or band present
Potential cause
Damaged test
Proper procedure not followed
Action
Repeat the test using new test and same sample (within time)
Follow each testing step according to the procedure
Re-check specimen volumes
Wait for the specified time before reading the test
Check expiration date of kit, if outdated, use test from in-date kit.

## Slide 39
What supplies do you need for urine specimen collection?
At what temperature / where should urine specimens be stored before or after being tested?
How much volume of urine is added to the sample pad?
What is the waiting time for reading Alere Determine TB LAM Ag test results?
Invalid results are a sign of positive test results
  True
  False
If the results are invalid, the test should be repeated.
  True
  False
Blood may be used for Alere Determine TB LAM Ag testing.
  True
  False
Knowledge Check – Question

## Slide 40: OPTIONAL Demonstration of LF-LAM testing
Trainer performs Alere Determine TB LAM Ag testing

## Slide 41: Continuous Quality Improvement of LF-LAM Testing

## Slide 42: Continuous Quality Improvement
Continual Quality Improvement (CQI) – A formal process/plan that is a cyclic series of steps designed to enhance quality assurance processes leading to improved outcomes.
Critical for improving LF-LAM testing reliability, accuracy, timeliness, effectiveness and impact
Planning
Implementation
Monitoring and Evaluation
Quality Improvement

## Slide 43: Considerations for LF-LAM Test Quality Assurance
Pipettors
Centrifuges
Refrigerators
Freezers
Initial
6 Months
Annually
Regular
Checklists
Above-Site and
Site-Level
quality
assurance testing
Randomized New Lot
QC Testing
EQA/PT
M&E Tool Use with Corrective Actions for Gaps
Algorithms
SOPs
Forms
Trainings
Material Lists
Forecasting
Monitoring
Storage
43

## Slide 44: Examples of quality assurance gaps and solutions for pre-testing, testing, and post-testing phases of the lf-lam procedure
 | Problem | Solution
Pre-Testing | Sample mislabeled or unlabeled Urine/ tests stored inappropriately Tests expired Work area not cleaned before/ between tests | Label urine & tests completely/ accurately Always store urine & tests appropriately Check inventory and expiration dates Disinfect work area before patient sample testing
Testing | Testing algorithm not followed (urine from HIV-negative persons tested) Incubation time not followed Invalid result missed and reported (+) or (-) Improper volume of specimen used Incompatible solution used as QC material | Only test samples collected from eligible patients Strictly follow urine volume & incubation times Take time to interpret valid and invalid results Never use water, PBS, or other solutions not compatible with LF-LAM tests as QC material
Post-Testing | Transcription error in reporting Negative results not reported Positive results not used to start TB treatment | Re-check patient/client identifier Always report valid results Use positive TB test results in clinical decisions

## Slide 45: Types of Quality Controls: internal and external
Internal Control is Included with the LF-LAM Test:
  Ensure the test is able to produce a valid result.
  Interpreted with every test run; if Control band is not present – the result is Invalid.
External Controls are NOT Included with LF-LAM Tests:
  External samples with known results are tested to verify that a tester at a testing site can use a standardized procedure to produce both valid and accurate results.
  External controls may be produced by a testing site (mWRD-positive and –negative urine), by a regional or national TB external quality assurance program provider, or by a commercial source
  Depending on availability, recommended for use: (1) at the beginning of each testing week, (2) once for each new lot of tests, (3) anytime the temperature exceeds the recommended range for LF-LAM testing (2 to 30°C), and (4) to investigate challenges with testing accuracy or reliability.
  When received, the testing site should: (1) store according to instructions, (2) record date of use (for urine), and (3) ensure use prior to expiry/ reporting dates

## Slide 46: MONITORING AND EVALUATION
Monitoring
Evaluation
Why? To provide indications of progress (or lack of progress)
Why? To improve the program using the data and information collected
How? Continuous collection of data and information to track progress to improve practices and have accountability
How? Determine the effectiveness, impact and sustainability of a program and understand why it is (or it is not) working
Adapted from Program Evaluation Home – CDC & Monitoring and Evaluation (aphl.org)

## Slide 47: Who conducts Monitoring and evaluation?
M & E is a collective responsibility

## Slide 48: M & E indicators for lf-lam testing
Testing Site M & E
  Data collected includes routine LF-LAM test and QC activities at the testing site
    Examples: Number of specimens and QC samples tested, days when LF-LAM testing could not be performed, status of enrolment in External Quality Assurance programs (if available), number of trained and competent testers
Supervisory Site M & E
  Consolidation of data from all the sites within the supervisory LF-LAM testing network
  Inclusive of in-patient, out-patient, and community/ outreach-based testing sites
LF-LAM Testing Cascade M & E
  Data collected on the use of LF-LAM test and its impact on timely bacteriological confirmation and diagnosis of TB in the context of the country algorithm

## Slide 49: Quality Improvement Tools
Onsite assessment and certification
Stepwise Process for Improving – Rapid Test (SPI-RT) for LF-LAM: Checklist for testing site quality assurance
Quality indicator monitoring
LF-LAM Testing M & E Tool with Automated Analyses
Improving Clinic lab interphase
CLICQ! | Clinic-Laboratory Interface – Continuous Quality Improvement & Diagnostic Cascade Evaluation (DiCE) Toolkit

## Slide 50: SUMMARY – KEY Lessons Learned
    Tuberculosis is caused by Mycobacterium tuberculosis, which is spread through the air
    LF-LAM tests are used to detect the lipoarabinomannan (LAM) cell wall component of Mtb in urine from HIV-positive persons
    LF-LAM is the only test that can bacteriologically confirm TB among PLHIV unable to produce sputum
    When testing is done according to the national algorithm and testing procedure by trained testers, LF-LAM may reduce TB-related morbidity and mortality among PLHIV
    The Alere Determine TB-LAM Ag Test is a point of care lateral flow assay that can be performed by trained lab- & non-lab testers in in-patient, out-patient, and community/ outreach settings
    The Determine TB LAM Ag Test is a low-complexity and rapid test that requires minimal sample handling and manipulation
    Common quality assurance challenges with LF-LAM testing can be addressed by ensuring testers receive appropriate initiation and refresher training, test according to the national algorithm, and routinely monitor and evaluate services
    There are a tools available for sites and supervisors for LF-LAM M&E to continuously improve the program & patient impact

## Slide 51: Acknowledgements
[ENTER COUNTRY/ TRAINING-SPECIFIC ACKNOWLEDGEMENTS HERE]